Oceanographic Data Collected in the Beaufort Sea

April 1991 - May 1991

# Equipment

Three Applied Microsystems STD-12 CTD's were used for data collection.  The serial numbers of the three units were 405, 517, and 510.  CTD 510 was brought along as a backup and only used at a few stations which were also measured using one of the other two CTD's.  CTD 510 was not calibrated and so no data for that unit are presented in this report.

## Stations

The stations visited during the field trip are show in Figures 1a, 1b and 1c.  All station locations were determined using GPS positioning systems.  When the fixed wing or rotary wing aircraft were used to reach a station, the aircraft GPS was used.  When a station was reached by skidoo a Magellan hand-held GPS model ??? was used.  All GPS units used C/A codes for positioning.  Over the duration of the field trip GPS "selective availability" was turned off and so the positions are expected to be within 50m of the true position.

The stations are divided into several main groups, each discussed below.

Off-Shore Transect Stations

-	Most of the stations were part of a series of transects that extended from the shore out into deeper water along lines approximately perpendicular to the isobaths. The transects done were as follows:

HI01 to HI02	-	Off Hershel Island

KP01 to KP03	-	Off Kay Point

GI01 to GI03	-	Off Garry Island

PI01 to PI05	-	Off Pelly Island

CM01 to CM05	-	Off Kugmallit Bay

and C006	-	Note: Station C006 was on the far side of a wide lead that was existed ??? miles off shore throughout the duration of the field trip.

AP01 to AP06	-	Off Atkinson Point.

Note: Station AP06 is not in line with the other AP stations.  The lead and ice conditions made shifting the position of the station necessary.

CD01 to CD05	-	Off Cape Dalhousie

REV Stations

-	CTD casts were performed at two stations labelled REV1 and REV2.  These stations were positioned along the CM transect and were also the sites for the deployment of two ARGOs REV units.

10 Meter Isobath Stations

-	Stations L001, L002, and L003 were positions approximately on the 10m isobath between off-shore transects in order to provide a more complete data set for that depth.

Mackenzie River/Delta Stations

-	CTD casts were done at several stations on the Mackenzie River and Mackenzie River Delta.

Cape Bathurst Stations

-	Stations CB01 to CB09 were done around Cape Bathurst and the mouth of Liverpool Bay

Liverpool Bay Stations.

-	A series of CTD stations, LB01 to LB09, were done in Liverpool Bay and into the Eskimo Lakes.

Mason Bay Station.

-	A single CTD cast, MB01, was performed at a station in Mason Bay.

Lead Stations

-	Three casts were done from a small Zodiak on the open lead north of the Tuktoyaktuk Peninsula.  Casts were done at stations BOAT and ISZ1.  Cast TOW1 consisted of towing the CTD behind the boat along the surface for some distance.

## Applied Microsystem Conductivity-Temperature-Depth System

Three AML model STD-12 CTD's were used.  Their serial numbers were 405, 517 and 510.  405 and 517 were the primary instruments used for data acquisition and 510 was used as a backup.  This report does not contain any data collected with unit 510. 405 and 517 were each configured with dual high/low range conductivity sensors for both fresh and salt water use.

For all but one station CTD casts were done with only one of the two primary CTD's.  At station SL01, however, casts with both 405 and 517 were taken.  At several stations a cast was done with 510 as well as one of the others.  Table 1 lists each station, its latitude and longitude, and the cast numbers associated with each of the CTD's at those stations.  An "F" next to the cast number indicates that the CTD was set to fresh water (low salinity) mode for that cast.

Data for each probe were logged in internal memory in each CTD during each cast.  When a cast was completed, the data were transferred to RAM-disk cards using an Atari Portfolio computer.  At the end of each day the cast files were copied from the Atari ram cards to a PC hard drive and backups were kept on floppy disk.  Data was collected at a nominal rate of 10 samples per meter.

Calibrations

Both 405 and 517 were calibrated at the factory just prior to going into the field.  510 had been calibrated sometime in the past but it became obvious once we started to use it that it's calibration coefficients were out of date.

Post trip calibrations were done on 405 and 517 but no calibration was done on 510.  Only the high conductivity range was checked in the post calibrations for both units. Post trip calibrations verified no significant change in the pressure or temperature responses of either unit.  Shifts in conductivity response, however, were measured for each unit.

Table 2 lists the sensor ranges, Pre-trip calibration standard deviations and post trip calibration differences for units 405 and 517.

From Table 2 we can see that unit 517 had a very consistent -0.0033 shift in conductivity ratio over the range of conductivities measured in the calibration(which was 0.56 to 0.98).  At -1 deg. C this would amount to a shift of approximately 0.2 PSU in salinity.  This shift is indicative of a problem with the conductivity cell in unit 517 which was also noted from the bottle data comparisons as discussed in section ???.

The shift of -0.00057 in unit 405's conductivity response would be equivalent to a shift approximately 0.02 PSU at -1 deg. C.  While this is a significant difference, it is probably more a reflection of the general accuracy limit of the instrument rather than due to a specific problem.

### Data Processing

Several steps were performed in the processing of each CTD cast.  These steps are outlined below:

1)	The raw values for pressure, temperature, and conductivity were converted to engineering units using the calibration coefficients determined in the pre-cruise laboratory calibrations.

At the start of each cast a reading was recorded with the CTD in the air and this value was subtracted from the pressures in the cast to remove any zero-offset.

2)	Salinity was calculated from temperature and conductivity using the Practical Salinity Scale-78.

3)	Erroneous values (spikes) were removed from the data using both automatic and manual techniques.  This mostly consisted of deleting the top 2 meters of data from casts in which the cast was done through a hole in the ice.

4)	The pressure readings for each cast were smoothed using an 11 point (approx 1 meter), un-weighted, running average.  This was done to eliminate noise in the data introduced by the pressure sensor prior to the next steps being performed.

5)	The pressure data for each cast was made monotonic by removing any data values from the casts in which the pressure was less than that for the previous data value accepted.

6)	The pressure, temperature, and salinity data for each cast was decimated into half-meter bins using a simple averaging process.

7)	An offset of -0.0617 PSU was applied to all salinity values taken with the CTD 517 (as outlined in section ???)

8)	Derived quantities were calculated from the pressure, temperature, and salinity data using the algorithms given in the UNESCO TECHNICAL PAPERS ON MARINE SCIENCE, No. 44 by Fofonoff and Millard.

The decimated data, and derived quantities, were used in producing the plots and tables found in Appendix 1.

### Data Validation

A direct intercomparison between CTD 405 and CTD 517 is available for one station: SL01.  This was the first station done during the field trip and three casts were done at it, two using 517 (casts #3 and #4) and one using 405 (cast #5).  Figures ??? and ??? show the temperature and salinity profiles for each of the three SL01 casts.

Comparing the two 517 casts we see that there is no perceptible shift in temperature but there is a shift of approximately 0.05 PSU between the two salinity readings.  This is five times the instrument specified accuracy.

Comparing the 517 casts with the 405 casts we see that there is a difference of less than 0.005 deg. C between the two.  This is well within the instruments accuracy.  In the salinities, however, there is up to a 0.12 PSU difference between the two units.   This intercomparison is strong evidence suggesting a problem with the 517 conductivity cell.

Further validation of the CTD readings was done by comparing the CTD data at stations in which bottle data were also collected. The bottle data were collected and analyzed for salinity as part of the chemical program (see subsequent sections and the chemical data tables).  In the intercomparison, it should be noted that the CTD casts and bottle samples were not collected simultaneously (although they were always collected within 1 hour of each other), and that depth for the CTD was determined from the pressure sensor, while depth for the bottles was estimated by "wire out".  Because of the relatively shallow depths of the casts, and the stable conditions of working from the ice, the depth measurements using the two systems are expected to be within +/- ??? meters of each other (1 meter?).

To quantify the bottle - CTD intercomparison, the property differences D = Xbottle - Xctd were calculated for each salinity in which the salinity gradient at the depth of the bottle was less than 0.01 PSU/meter.  For the most part, only the stations farthest away from shore had uniform enough profiles to be used in the bottle - CTD intercomparisons.

Figures ??? and ??? show the salinity differences (bottle-ctd), as a function of cast number, for CTD's 405 and 517 respectively.  Figure ??? shows the salinity differences as a function of depth for CTD 405 casts #9, #10, #11 and #43. Casts #9, #10, and #11 were done consecutively over the period of an afternoon.  Cast #43 is unique in that it was taken after the CTD had already been in the water for 1 hour being towed behind the Zodiak on the lead.  Figure ??? shows the salinity differences as a function of depth for CTD 517 casts #21 and #22.  Casts #21 and #22 were taken 3 1/2 hours apart one afternoon.

The figures for unit 405 show a fairly random scattering of differences in the +/- 0.04 PSU range.  For any given cast the spread of differences was within a range of about +/-0.01 PSU which is comparable to the instrument's specified accuracy.  The shift in mean difference from cast to cast could be do to any number of things but is probably due to temperature effects on the instrument's electronics.

The figures for unit 517 show differences between bottle and CTD salinities up to -0.1 PSU with a range of up to +/- 0.05 PSU in a single cast.  Except for one point, the CTD salinities are higher than the bottle salinities and there is a clear trend towards higher CTD salinity errors at greater depths.  There is also an indication that the offset in CTD salinity is somewhat long lasting in that cast #22 starts off at a large negative offset probably induced by cast #21.  The behavior of the CTD 517 salinity readings is consistent with what would happen if there was a leak in its conductivity cell.

Table 3 give the mean and standard deviations of the salinity differences calculated for both CTD 405 and 510.  From these statistics, the decision was made to apply a

-0.0617 PSU offset to all CTD 517 salinities to bring them to better agreement with the bottle data.  The post-trip calibration of CTD 517 and the comparison between CTD's 405 and 517 suggest that the actual errors in the 517 salinity readings could be as large as 0.2 PSU for some casts and so the actual uncertainty for that unit is most certainly higher than 0.033 PSU, probably closer to 0.1 PSU.

Description of Data Plots

Each page represents one cast with the cast number being the last three digits in the reference number.  There are two plots on each page: a profile of temperature and salinity, and a T-S plot.  In addition to the data, the T-S plots show the freezing point line and two or three sigma-t lines (with the sigma-t value shown next to the line).  A table of measured and computed properties is also provided for each cast. The units used for each property are as follows:

Pressure			-	decibars

Depth			-	meters

Temperature		-	degrees Celcius

Theta (Pot. Temp)	-	degrees Celcius

Salinity			-	PSU

Sigma-t			-	grams/meter^3

Svan (Specific Vol.	-	???

(Anomally)

Dynamic Height		-	meters

Potential Energy	-	???

Sound Speed		-	meters/second

Only one cast for each station is reported when multiple casts from CTD 405 and CTD 517 are available.  Specifically: at station LB04 cast #62 is reported, at station REV2 cast #46 is reported, and at station SL01 cast #5 is reported.

Each cast was plotted using a temperature scale of -2 to 0 degrees C and a salinity scale of 0 to 35 PSU.  Where warrented, a second page may also be provided showing all or part of the cast on a smaller scale to reveal more detail.

CTD Cast Numbers

Station

Name

Latitude

Longitude

S/N 405

S/N 517

S/N 510

AP01

69o 59.7' N

131o 28.0' W

60

AP02

70o 04.2' N

131o 25.2' W

59

AP03

70o 08.5' N

131o 25.1' W

24

AP04

70o 15.5' N

131o 28.5' W

22

23

AP05

70o 18.0' N

131o 21.3' W

11

AP06

70o 18.4' N

131o 54.2' W

36

BOAT

70o 11.1' N

133o 38.8' W

41

C006

70o 28.3' N

133o 51.5' W

21

CAN2

69o 24.5' N

133o 49.0' W

34F

CB01

70o 31.3' N

128o 18.1' W

48

CB02

70o 30.3' N

128o 14.9' W

50

CB03

70o 29.7' N

128o 13.3' W

51

CB04

70o 29.8' N

128o 28.4' W

52

CB05

70o 29.7' N

128o 25.4' W

53

CB06

70o 30.0' N

128o 21.1' W

54

CB07

70o 38.6' N

128o 18.0' W

55

CB08

70o 38.6' N

128o 15.9' W

56

CB09

70o 38.4' N

128o 15.9' W

57

CD01

70o 20.0' N

129o 40.0' W

44

CD02

70o 25.0' N

129o 40.0' W

45

CD03

70o 31.5' N

129o 39.5' W

40

CD04

70o 40.0' N

129o 40.8' W

39

CD05

70o 48.3' N

129o 40.4' W

37

38

CM01

69o 45.0' N

133o 17.7' W

32

CM02

69o 50.0' N

133o 20.0' W

17

CM03

69o 56.0' N

133o 21.6' W

16

CM04

70o 01.0' N

133o 23.9' W

15

CM05

70o 08.1' N

133o 25.3' W

10

FLGP

69o 27.5' N

133o 04.4' W

13F

GI01

69o 30.2' N

136o 09.5' W

71

GI02

69o 38.1' N

136o 22.0' W

77

GI03

69o 43.9' N

136o 31.9' W

78

HI01

69o 28.0' N

138o 50.6' W

72

HI02

69o 36.8' N

138o 47.8' W

73

ISZ1

70o 11.6' N

133o 40.8' W

43

Table 1 - Station Locations and Cast Numbers

CTD Cast Numbers

Station

Name

Latitude

Longitude

S/N 405

S/N 517

S/N 510

KITI

69o 22.6' N

133o 41.7' W

12F

KP01

69o 18.8' N

138o 10.5' W

74

KP02

69o 23.9' N

138o 07.5' W

75

KP03

69o 30.0' N

138o 08.4' W

76

L001

69o 53.0' N

133o 60.0' W

31

L002

69o 54.0' N

132o 13.8' W

35

L003

70o 16.3' N

130o 29.7' W

58

LB01

70o 15.0' N

128o 52.4' W

47

LB02

69o 59.8' N

129o 10.4' W

64

LB03

69o 42.3' N

130o 25.8' W

63

LB04

69o 30.2' N

131o 29.8' W

62,66

LB05

69o 29.0' N

131o 45.0' W

65

LB06

69o 20.4' N

131o 57.6' W

67

LB07

69o 13.4' N

132o 26.8' W

68

LB08

69o 11.2' N

132o 41.5' W

69

LB09

68o 59.4' N

132o 57.5' W

70

MB01

69o 33.2' N

134o 04.5' W

61

PI01

69o 43.8' N

134o 31.2' W

30

PI02

69o 48.7' N

134o 38.2' W

29

PI03

69o 55.0' N

134o 48.0' W

20

PI04

69o 57.9' N

134o 53.7' W

19

18

PI05

70o 02.0' N

135o 00.0' W

9

REV1

69o 42.8' N

133o 11.6' W

1

REV2

69o 56.0' N

133o 20.7' W

2,8,46

SL01

69o 53.6' N

135o 39.2' W

5

3,4

SL02

69o 47.6' N

135o 37.8' W

6

SL03

70o 00.0' N

131o 28.6' W

7

TKAK

69o 31.5' N

133o 08.0' W

14

TPT1

69o 39.3' N

132o 56.1' W

25

TPT2

69o 39.8' N

132o 59.3' W

27

TPT3

69o 40.2' N

133o 01.1' W

28

WSTC

69o 21.4' N

133o 57.7' W

33

Table 1(cont) - Station Locations and Cast Numbers

Parameter

S/N 517

S/N 405

Range

0 to 1378 dBars

0 to 690 dBars

Pressure

Calib. Std. Dev.

+/- 0.31 dBars

+/- 0.07 dBars

Post-Trip Calib. Difference

0.69 +/- 0.08 dBars

0.36 +/- 0.17 dBars

Range

-2 to +38

deg. C

-2 to +38

deg. C

Temperature

Calib. Std. Dev.

+/- 0.0067

deg. C

+/- 0.0089

deg. C

Post-Trip Calib. Difference

-0.00224 +/- 0.0116

deg. C

-0.00171 +/- 0.0111

deg. C

Range

0 to 40

(RATIO)

0 to 40

(RATIO)

Conductivity

(High Range)

Calib. Std. Dev.

+/- 0.00047

(RATIO)

+/- 0.00017

(RATIO)

Post-Trip Calib. Difference

-0.0033 +/- 0.00017

(RATIO)

-0.00057 +/- 0.00021

(RATIO)

Range

0 to 2

(RATIO)

0 to 2

(RATIO)

Conductivity

(Low Range)

Calib. Std. Dev.

+/- 0.00036

(RATIO)

+/- 0.000067

(RATIO)

Post-Trip Calib. Difference

N/A

N/A

Table 2. - CTD Calibration Information

Salinity Diff. (PSU)

Bottle-CTD

S/N 405

S/N 517

Mean

0.000084

-0.0617

Std. Dev.

0.02331

0.0334

Table 3 - Bottle - CTD Salinity Difference Statistics