File: 1_PROCESSING.DOC

Date: September 24, 1991

by:   Edmand Fok

Processing summary:

- run FIELD_DATA_TO_RAW program

- all temperature moved up 0.004 degree C by adding 0.004 to offset.

0.0022 + 0.004 => 0.0062

- after initial calculation, deduct conductivity ratio offset by

0.000173977 to reduce salinity value by 0.01.

(-0.00108 - 0.000173977) => -1.253977E-03

- equation used to calibrate:   value = A + B*N

offset (A)         coefficient (B)

PRESSURE               0.63000             1.000290

TEMPERATURE            0.00620             1.000758

CONDUCTIVITY RATIO    -1.253977E-03        1.002050

THERMISTOR 1          -0.0020              0.997950

THERMISTOR 2           0.0006              0.998960

- edit HEADER_RAW files as follows:

- CALIBRATION SECTION is updated as follows:

- number of channels in raw data is set to 8

PRESSURE                    dbars    10  2  0.63000        1.000290

TEMPERATURE                 celsius  10  2  0.00620        1.000758

TEMPERATURE                 celsius  10  2  0.00620        1.000758

CONDUCTIVITY RATIO                   10  2  -1.253977E-03  1.002050

CONDUCTIVITY RATIO                   10  2  -1.253977E-03  1.002050

TEMPERATURE : THERMISTOR 1  celsius  10  2  -0.0020        0.997950

TEMPERATURE : THERMISTOR 2  celsius  10  2   0.0006        0.998960

FLAG                                 10  2  0.0000000E+00  1.000000

- number of calculated parameters is set to 2

SALINITY          30   3   2   4   1

DENSITY  kg/m**3  40   2   2  -1

- START DATE is set according to information from DRIFTCHK.ANS

- INCREMENT TIME is set to 0.04 sec = 40 msec

- NUMBER OF RAW DATA RECORDS is set to the number of records in the

raw data file.

- use VIEW program to truncate beginning of profiles:

Profile                Record

6840                  1 to 137

6841                  1 to 139

6842                  1 to 131

6843                  1 to 135

6844                  1 to 134

6845                  1 to 159

6846                  1 to 138

6847                  1 to 135

6848                  1 to 135

6849                  1 to 137

6851                  1 to 135

6852                  1 to 149

6853                  1 to 142

6854                  1 to 133

6855                  1 to 137

6856                  1 to 140

- run auto_edit program

File CHANNEL.DAT:

PRESSURE

TEMPERATURE

CONDUCTIVITY RATIO

TEMPERATURE : THERMISTOR 1

TEMPERATURE : THERMISTOR 2

FLAG

File RUN_AUTO_EDIT.COM:

$ AUTO_EDIT$

3         ! 1 (save com and run), 2 (com only), (3 run only)

INPUT.DAT  ! Input control file

CHANNEL.DAT  ! Channel description file

T,T       ! perform SPIKE_EDIT (T/F), perform FALL_SPEED_EDIT (T/F)

PRESSURE

TEMPERATURE

CONDUCTIVITY RATIO

TEMPERATURE : THERMISTOR 1

TEMPERATURE : THERMISTOR 2

$EOD   ! End of channel selection

T      ! if .true., all input channels are correct

T      ! Edit against record number (T) or data values of a selected channel (F)

T      ! Use variable bounds (T/F)

T      ! Change of epsilon value for each channel

5.000E-02  ! epsilon value for channel: PRESSURE

2.000E-03  ! epsilon value for channel: TEMPERATURE

3.000E-05  ! epsilon value for channel: CONDUCTIVITY RATIO

0.002  ! epsilon value for channel: TEMPERATURE : THERMISTOR 1

0.002  ! epsilon value for channel: TEMPERATURE : THERMISTOR 2

T      ! Change of Abs. Upper bound & Abs. Lower bound for each channel

0.001, 0.0 ! Abs. Upper & Abs. Lower bound for channel: PRESSURE

1.0, -2.0  ! Abs. Upper & Abs. Lower bound for channel: TEMPERATURE

8.0,  0.5  ! Abs. Upper & Abs. Lower bound for channel: CONDUCTIVITY RATIO

1.0, -2.0  ! Abs. Upper & Abs. Lower bound for channel: TEMPERATURE : THERMISTOR 1

1.0, -2.0  ! Abs. Upper & Abs. Lower bound for channel: TEMPERATURE : THERMISTOR 2

T         ! Change of bound limits for each channel

3.000000E+00  ! Bound limits for channel: PRESSURE

3.000000E-01  ! Bound limits for channel: TEMPERATURE

7.500000E-03  ! Bound limits for channel: CONDUCTIVITY RATIO

3.000000E-01  ! Bound limits for channel: TEMPERATURE : THERMISTOR 1

3.000000E-01  ! Bound limits for channel: TEMPERATURE : THERMISTOR 2

21        ! Window length, number of data within window

2.7       ! Standard deviation limit : multiple of standard deviation

T         ! if T, perform SPIKE_EDIT for the whole file.

13, 0.3   ! # scans over which to average Fall Speed, Min. acceptable Fall speed

T         ! if T, perform FALL_SPEED_EDIT for the whole file.

_EDIT     ! Part of output file extension

- run sensor_adjustment program

!  Command file:       RUN_SENSOR_ADJUSTMENTS.COM

$ SENSOR_ADJUSTMENTS$

3         ! 1 (save com and run), 2 (com only), (3 run only)

FILES_TO_ADJUST.DAT  ! Input control file

0.0   ! Separation of pressure diaphragm & centre of conductivity cell

0.07  ! Separation of temperature sensor & centre of conductivity cell

1.00000E-07  ! Pressure coefficient for conductivity on PYREX glass

-3.20000E-06  ! Temperature coefficient for conductivity on PYREX glass

0.04  ! Sample interval between records in seconds

32    ! Separation of scans to compute fall speed.

2     ! # of temperature and conductivity per sample.

0.05  ! Delay in temperature response in seconds.

0.05  ! Temperature response time.

0.077 ! C_Cell response distance in meter.

_ADJUST  ! Part of output file extension

- run calirbation program

$ RUN CALIBRATION$

INPUT.DAT

F,T    ! PRINT FILE, SAVE FILE

T      ! FILE FORMATTED

F      ! CREATE TIME CHANNEL

/      ! FIRST RECORD, LAST RECORD

F      ! DO NOT CALIBRATE ALL CHANNELS

5      ! total number of channels to calibrate

17

22

1

20

3

- run state_parameters program

File OUT_CHANNEL.DAT:

PRESSURE

TEMPERATURE

SALINITY

TEMPERATURE : (T - TF) : SURFACE

GAMMA

PHI

File RUN_STATE_PARAMETERS.COM:

$ STATE_PARAMETERS$

3         ! 1 (save com and run), 2 (com only), (3 run only)

files_to_final.dat  ! Input control file

OUT_CHANNEL.DAT  ! Channel description file

T    ! if [F], use following entries instead of entries from channel file.

dbars     ! Unit for  PRESSURE

F7.2      ! Format for  PRESSURE

T         ! Expecting this channel from input data file

-99999    ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

celsius   ! Unit for  TEMPERATURE

F7.3      ! Format for  TEMPERATURE

T         ! Expecting this channel from input data file

-99999    ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

! Unit for  SALINITY

F7.3      ! Format for  SALINITY

T         ! Expecting this channel from input data file

-99999    ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

celsius   ! Unit for  TEMPERATURE : (T - TF) : SURFACE

F7.3      ! Format for  TEMPERATURE : (T - TF) : SURFACE

F         ! Expecting this channel from input data file

-99999    ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

kg/m**3   ! Unit for  GAMMA

F7.3      ! Format for  GAMMA

F         ! Expecting this channel from input data file

0         ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

J/kg      ! Unit for  PHI

F9.4      ! Format for  PHI

F         ! Expecting this channel from input data file

0         ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

C         ! Confirmation of record format

_FINAL    ! Part of output file extension

10        ! Window size for pressure running average.

1.00000E+00  ! Interpolate pressure at every db level.

20       ! Size of Interpolation window.

- for profile 6840 to 6846, because of the timing problem with the

conductivity ratio and the temperature, many spikes are found in

pressure-salinity plot.

- the following adjustments are made on profie 6840 to 6846:

- low pass filter is applied to profile 6840 to 6846 after auto_edit

program.

File RUN_DLAN.COM:

$ run dlan

4

3500

490

L

6840.header_edit

F T F T F F F F

6841.header_edit

F T F T F F F F

6842.header_edit

F T F T F F F F

6843.header_edit

F T F T F F F F

6844.header_edit

F T F T F F F F

6845.header_edit

F T F T F F F F

6846.header_edit

F T F T F F F F

- after salinity is calculated, multiplication factor (1.000185)

is applied on conductivity ratio to increase salinity by 0.007.

- temperature and conductivity ratio channels are changed as

follows:

SMOOTHED TEMPERATURE  celsius  10  2  6.2000002E-03  1.000758

TEMPERATURE           celsius  10  2  6.2000002E-03  1.000758

CONDUCTIVITY RATIO             60  3 -1.2539770E-03  1.002050  1.000185

COLD                           60  3 -1.2539770E-03  1.002050  1.000185

- salinity is calculated by smoothed temperatute and smoothed

conductivity.

- density is calculated by smoothed salinity and smoothed temperature.

- calibration program is run by:

File CAL_CHANNEL.DAT:

PRESSURE

TEMPERATURE

TOLD

CONDUCTIVITY RATIO

COLD

TEMPERATURE : THERMISTOR 1

TEMPERATURE : THERMISTOR 2

SALINITY

DENSITY

File RUN_CALIBRATION.COM:

$ run ccs$physics:[edf.debug.exe]calibration

input_to_calib.dat

F,T    ! Print and/or Save each channel value (T/F) ?

T      ! Output file (T) Formatted or (F) Unformatted ?

F      ! Create Time channel in output file from start time

1,999999    ! Start record value , End record value

F      ! Calibrate each channel within each file

cal_channel.dat

T      ! select channel

PRESSURE

T

TEMPERATURE

T

TOLD

T

CONDUCTIVITY RATIO

T

COLD

T

SALINITY

T

DENSITY

F

T

- run state_parameters program:

File OUT_CHANNEL.DAT:

PRESSURE

TEMPERATURE

TOLD

SALINITY

DENSITY

PHI

File RUN_STATE_PARAMETERS.COM:

$ STATE_PARAMETERS$

3   ! 1 (save com and run), 2 (com only), (3 run only)

input_to_state.dat  ! Input control file

OUT_CHANNEL.DAT  ! Channel description file

T   ! if [F], use following entries instead of entries from channel file.

dbars     ! Unit for  PRESSURE

F7.2      ! Format for  PRESSURE

T         ! Expecting this channel from input data file

-99999    ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

celsius   ! Unit for  TEMPERATURE

F7.3      ! Format for  TEMPERATURE

T         ! Expecting this channel from input data file

-99999    ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

celsius   ! Unit for  TOLD

F7.3      ! Format for  TOLD

T         ! Expecting this channel from input data file

-99999    ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

! Unit for  SALINITY

F7.3      ! Format for  SALINITY

T         ! Expecting this channel from input data file

-99999    ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

kg/m**3   ! Unit for  DENSITY

F7.3      ! Format for  DENSITY

T         ! Expecting this channel from input data file

0         ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

J/kg      ! Unit for  PHI

F9.4      ! Format for  PHI

F         ! Expecting this channel from input data file

0         ! Reference pressure (dummy value if -ve)

C         ! Confirmation of setting for the current channel

C         ! Confirmation of record format

_FINAL    ! Part of output file extension

9         ! Window size for pressure running average.

1.00000E+00  ! Interpolate pressure at every db level.

21        ! Size of Interpolation window.

- edit all header files to add remaks and other information.