## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 8-Feb-2012 | - Added three productivity casts, containing Primary Productivity, Chlorophyll, Nutrients, POC and PON, from Frank Whitney’s Productivity spreadsheet. |

Processing notes           Jun 11, 1993

Cruise 92-06

Mar 27 - Apr 12, 1992

73 original files, 67 processed files.

casts 1 - 74

There were 2 CTD probes used on this cruise.

Probe 53,501 was used for casts  1 - 43 and 46-74.

Probe 58,483 was used for casts 44 & 45.

1.  Calibrations  (Program Calib)

- For probe 53,501

offset        slope

------        -----

pressure        -0.41          2015.1  Ron Cooke's lab calibrations

temperature      0.012       0.999773  Average of pre and post cruise calibrations

conductivity    -0.00103     1.000984  Bottle salinities

- For probe 58,483

offset       slope

------       -----

pressure          0.0         3000.0   Oct  7/91 lab calibration

temperature       0.0         1.0      Nov 20/91 lab calibration

conductivity     -0.00016     1.0      Nov 20/91 lab calibration

2.  Despiking

The following DESPIKE parameters were used.

DESPIKE TABLE

C  FIT    OVER-  MIN     MAX      MIN      MAX     SPIKE

H  WIDTH  LAP    VALUE   VALUE    STDEDV   STDDEV  TOL.

-  -----  -----  ------  -------  -------  ------  -----

P     25      5    0.00  2200.00   0.1000  30.000   2.50

T     25      5    0.00    16.00   0.0010   3.000   2.50

C     25      5    0.50     1.00   0.0005   0.005   2.50

3.  Time compensation

The following input parameters were used for both probes:

Temp. Probe Dist from Cond. Cell Mouth  0.07 meters

Sample Period                           0.04 sec

4.  Swell removal  (Program DELETE)

The following delete parameters were used.

Pressure NOT filtered

Swells deleted

Low Drop Rates NOT deleted

5.  Editing

Editing was done by Joe Linguanti and Ron Perkin using program CTDEDIT.

6.  Particulars

cast 14 - deleted, upcast for 13.

cast 20 - deleted, cast aborted.

cast 25 - deleted, cast aborted.

cast 27 - deleted, bad pressure data.

cast 49 - deleted, cast aborted.

cast 59 - missing

Joe Linguanti

# Moving files to new archive. July 25, 2002 – Joe Linguanti

- Nine hydro files containing temperature, salinity, oxygen and nutrient data were obtained from the OSAP archives. Bernard had merged the nutrient data from Frank Whitney’s data set in 1999.  The nutrient data in Bernard’s files were in umol/kg so I removed them from the files and obtained the nutrient data in umol/L from Frank’s spreadsheet file.

- Program HEADER EDIT was used to change non-standard channel names and units and Data Description, Scientists name, Project, Platform and Instrument fields.

- Program Clean was used to;

- reset number of records, minimum and maximum values

- delete empty channels where applicable

- set channel name initial letters to uppercase

- Set last 4 characters of file name to event – set cast numbers to be the same as event numbers in file.

- Files .CLN were copied to new archive and renamed to the new file convention “YYYY-CC-NNNN.* “.

May 21, 2003 – Joe Linguanti

- The loop file was obtained from Frank Whitney. The spreadsheet file was converted to IOS Header format with meta-data added to it. The original .XLS file is in the DOC directory.