| 9207 Events |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Date | CTD Casts | Stations | 1030 | 1031 | Comments | Servicing |
| April 14 | 1 2 | E600 E800 | ü ü |  | Doug & Tom Bad Casts |  |
| April 15 |  |  |  |  |  |  |
| April 16 |  |  |  |  |  |  |
| April 17 | 3 4 5 6 7 | E800 E800 R292 R292 R192 | ü ü ü ü | ü | Doug & Tom & Rick - At station R292 CTD took a jolt which may have introduced air into inlet tube |  |
| April 18 |  |  |  |  |  | -Topped up oil in PT inlet tube - Replace batteries in both CTD's |
| April 19 | 8 9 10 11 12 | E100 E100 E200 E200 E300 | ü ü ü ü ü |  | Rick & Dave - Took sample bottles with casts 9 & 11 |  |
| April 20 | 13 14 15 16 17 18 19 20 | E400 E400 E500 E600 E630 E900 E800 E700 | ü ü ü ü ü ü ü ü |  | Doug & Dave - Took sample bottles with casts 14 & 20 |  |
| April 21 |  |  |  |  |  | -Replaced DO membrane (bubbles had become trapped under membrane) |
| April 22 |  |  |  |  |  | - Modified inlet tube on both CTD's by stretching latex under cap |
| April 23 | 21 22 23 | T100 T100 T100 | ü ü | ü | Skidoo Test - Pressure readings are garbage. | Did calibration of O2 probe, SOC = 0.1134 BOC = -0.0495 |
| April 24 | 24 | T200 | ü |  | Skidoo Test - Pressure readings good again | Restored pt inlet tubes to normal configuration |
| April 25 | 25 26 | T100 T100 | ü | ü | Skidoo Test - Good profiles | Modified PT ports with AML bladder |
| April 26 | 27 28 | R0B0 E300 | ü ü |  | Skidoo to Huskey Lakes - ROB0 bad pumping - Chemistry at E300 |  |
| April 27 | 29 30 31 32 | E900 E900 E230 E220 | ü ü | ü ü | Twin Otter to E900 with full chemistry Skidoo to E230 and E220 |  |
| April 28 | 33 34 35 36 | E110 E200 E200 E210 |  | ü ü ü ü | Skidoo in Huskey Lakes - Batt failed during #33 and #34 - Full Chem at E200 | Replaced batteries at E200 and toggled reset switch. |
| April 29 |  |  |  |  |  |  |
| April 30 |  |  |  |  |  | Did test of pump flow rates: 1030: 5l in 2:37 1031: 5l in 5:06 |
| May 1 | 37 38 39 | E660 E650 E640 |  | ü ü ü | Skidoo in Huskey Lakes |  |
| May 2 | 40 41 42 | E630 E620 E600 |  | ü ü ü | Skidoo in Huskey Lakes - Full Chem at 630 - E600 done free flushing because of ice in pump |  |
| May 3 | 43 | E500 | ü |  | Twin Otter - Full Chem |  |
| May 4 | 44 45 46 47 48 49 | E100 E100 E400 E400 E800 E800 | ü ü ü | ü ü ü | Twin Otter - Full Chem and CTD intercomparissons at each station | Changed DO sensor membrane |
| May 5 | 50 51 52 53 | E670 E200 E665 E663 | ü ü | ü ü | Twin Otter - Full Chemistry at 670 - Freon at 200 Helicopter - CTD with no pump at 665 and 663 |  |
| May 6 |  |  |  |  |  |  |
| May 7 | 54 55 56 57 58 59 | R192 R292 E400 E370 E360 E350 | ü ü | ü ü ü ü | Twin Otter - Mooring Deployments at R192 and R292 Skidoos to Huskey Lakes at E370 to E350 - LVF at E400 - CTD's use no pump |  |
| May 8 | 60 61 62 63 64 65 66 | E341 E330 E225 E300 E661 E662 E670 | ü ü ü ü ü | ü ü | Skidoos at Huskey Lakes E341 and E330 - CTD's use no pump Twin Otter remaining stations |  |
| May 9 | 67 68 69 | E300 SITL SITL | ü ü ü |  | Twin Otter |  |
| May 10 |  |  |  |  |  |  |
| May 11 | 70 71 72 73 | MB92 MB92 E300 E300 | ü ü | ü ü | Twin Otter - Full Chemistry at MB92 - CTD calibration bottles at E300 | DO sensor plenum inlet arm broke off. Glued back together before going to E300 |
| May 12 | 74 | E800 | ü |  |  | Did I replace the DO membrane before going out today? I have no notes of it but I expect that I would have. |