Table of Contents

µ1.	Stations	1

1.1.	Station Nomenclature	1

1.2.	Station Locations	1

2.	Sea-Bird Conductivity-Temperature-Depth System	1

2.1.	Data Collection Methods	2

2.1.1.	Working Environment	2

2.1.2.	CTD Handling Procedures	2

2.2.	Instrument Problems	4

2.2.1.	Freezing	4

2.2.2.	Battery Failure	4

2.3.	Data Processing	5

2.4.	Data Validation	6

2.4.1.	Inspection of the Multiple ProfilesTaken for Each 
	Cast.	6

2.4.2.	Comparison of the CTD Data with the Bottle data for 
	Salinity and Dissolved Oxygen	6

2.4.3.	Intercomparision of the Two CTD's for Temperature 
	and Salinity.	10

2.4.4.	Comparision of the Pre-trip and Post-trip Laboratory 
	Calibrations.	12

3.	List of Stations	14

4.	Appendix A. - CTD Handling Guidlines Used During 
	Field Trip 92-07	18

# 1.	Stations

## 1.1.	Station Nomenclature

All stations have been given a 4 character designation.   The majority of the stations reside in the Husky Lakes and Liverpool Bay areas and these station names all consist of an "E" followed by three digits.  The "E" Numbers run from E100 at the start of the Husky Lakes to E900 at the mouth of Liverpool Bay.

There are six stations in addition to those in the Husky Lakes and Liverpool Bay area: SITL (Sitigi Lake sp?), MB92 (Mason Bay), T100 and T200 (Tuktoyaktuk Harbour), and R192 and R292 (at the mouth of Kugmallit Bay).

## 1.2.	Station Locations

Station locations, shown in Figures 1 and 2, were determined using a number of GPS systems including a hand held Trimble Transpak II (used during snow mobile operations) and the GPS systems aboard the fixed and rotary-wing aircraft.  The positions are expected to be within 100m of the true position.  Refer to table 7 for a complete list of the stations and the casts taken for each CTD.

Many of the stations were visited more than once over the period of the field trip.  In most cases the same location was used for each visit (as marked by the hole drilled by the first visit).  In some cases, however, slightly different locations were used in an attempt to find a deeper basin.

# 2.	Sea-Bird Conductivity-Temperature-Depth System

Two Sea-Bird model SBE-19 probes were used over the duration of the cruise.  They each had pumped conductivity cells and one had a pumped O2 sensor.   Table 1 lists the configuration of each CTD.

| S/N | Max Pressure (dBars) | Pumped Cond Cell | O2 Probe |
| --- | --- | --- | --- |
| 1030 | 200 | Yes | Yes |
| 1031 | 600 | Yes |  |

Table 1 - Sea-Bird CTD Configurations

The Sea-Bird CTD's were self-recording and all data were logged internally during the casts and then transferred to a desk-top PC computer by serial cable at the end of each day.   Also, data were often transferred to a portable PC in the field to check data integrity.  Data were collected at a rate of 2 samples per second.

## 2.1.	Data Collection Methods

### 2.1.1.	Working Environment

All stations were located on ice-covered waters.  The stations were reached by fixed-wing, and rotary-wing aircraft and by snow mobile.  In all cases, holes had to be augured or melted through the ice in order to be able to take the CTD casts and water samples.

The working temperatures varied from about 0o C to -20o C.   These temperatures posed a significant problem when working with the Sea-Bird SBE-19 CTD's because of the susceptibility of the pump and plumbing systems to freezing.   This problem was complicated by the fact that we could not keep the conductivity cell filled with alcohol (and hence prevent freezing) because the manufacturer warned that the platinized electrodes on the conductivity cell would deteriorate if soaked in alcohol for extended periods.

When working from an aircraft we were able to keep the CTD's inside and warm until just before doing the cast and so avoided most freezing problems.  When we worked from the snow mobiles we kept the CTD warm by hanging it in a box mounted to a sled that was towed by the snow mobile.  The box contained a catalytic heater which was able to keep the temperature of the CTD above freezing most of the time.  We did find, however, that when we did a number of casts in one day the CTD became progressively more cold-soaked and freezing became a problem.

### 2.1.2.	CTD Handling Procedures

The actual procedure used for each cast varied over the course of the field trip because our methods adapted to our on-going experiences.   By cast #8, however, we had developed a set of guidelines that each team was to follow when taking casts and for the most part these were the steps that were performed for each cast.   A copy of the finalized guidelines we used in the field are reproduced in Appendix A.

The main points to note regarding the handling procedures are:

·	Prior to each cast we plugged the end of the conductivity cell and filled the conductivity cell (and oxygen probe if present) with warm 40 ppt water.  We did not pull the plug from the cell until the CTD was submerged.    We did this with the hopes of preventing the cell from freezing on the way from the plane to the hole.

·	The CTD was lowered by hand using a kevlar line marked out in 1 meter and 10 centimeter increments.

·	We let the CTD soak for 5 minutes before performing the cast.  The CTD was programmed to turn on the pump after 4 minutes.

·	The CTD was lowered at 4 seconds/meter so that we could get 8 samples per meter.

·	For each cast we did two to three profiles of the water column.  This was done so that we could verify the repeatability of the data.  This procedure proved very useful when it came time to process the data.

·	When we had the oxygen probe, the last profile was done in steps to allow the oxygen probe to stabilize for 1 minute at a number of depths.  This was done because of the long time constant associated with the oxygen probe,  The data from the soaking periods was used as a comparison to the continuous profile.

·	We often took a salinity sample for calibration in conjunction with a CTD cast. This was always done as a subsequent cast to the main multi-profile cast for the station.  We hung a 1.7 litre Nisken bottle on the line 2 meters above the CTD and then lowered the CTD so that it profiled the water being sampled just before the bottle was tripped.

For some casts we found it necessary to operate without the pump flushing the conductivity cell.  In those cases we lowered the CTD at 1 meter per second in order to provide adequate flushing of the conductivity cell.

In order to be able to compare the pumped verses free-flushing conductivity data we did a number of casts in which we did three profiles with the pump attached followed by one or two more with it detached.  Detaching the pump was a simple matter of pulling the hose off of the top end of the conductivity cell while the CTD was momentarily pulled out of the water.

The resulting data showed no changes in the conductivity readings except where the CTD passed through an abrupt halocline.  In these cases the free flushing conductivity data lagged the pumped data in a manner consistent with the differences in lowering rate and the time constant of the cell.

## 2.2.	Instrument Problems

### 2.2.1.	Freezing

For many casts, the data were corrupted by the instrument freezing in one way or another.  The three main freezing problems were:

1.	Pressure Inlet Port Freezing

The instrument came supplied from the manufacturer with an oil-filled, open-ended tube connected to the pressure inlet port.  This configuration caused problems in  a couple of casts because the oil would get knocked out of the end of the tube allowing water to enter and freeze.  This had the effect of "damping" the pressure signal being measured so that true depths were not recorded.

This problem was fixed in later casts by replacing the inlet tube with an oil-filled bladder covering the pressure inlet port.

2.	"Y" Fitting Air Hole Freezing

Because the pump on the SBE-19 is not self-priming, it is necessary to allow the air to bleed from the plumbing prior to the pump being turned on.  This is done by allowing the air to bleed  out through a small hole in the "Y" fitting at the top of the instrument.  This hole is very small, however, and quite prone to freezing.  It was necessary to poke the ice out of this hole just before the instrument went in the water or the pump would not prime.  Several casts have bad pumping because of ice reforming in this hole before the CTD went in the water.

3.	Pump Impeller Freezing

If the CTD became cold-soaked to any degree, the pump impeller had a tendency to freeze as soon as it was put in the water.  In many cases it was not possible to warm the CTD sufficiently to avoid this problem.  In some cases, if we were aware of the pump freezing we would disconnect it from the conductivity cell and allow the cell to free-flush during the cast.

### 2.2.2.	Battery Failure

A few of the casts were corrupted because the CTD battery voltage dropped too low and the CTD automatically turned itself of in the middle of the cast.  We were normally able to catch this problem because we noticed the pump was not running when the CTD came out of the water and we would replace the batteries and redo the cast.  Cast #33 at station E110 was not caught until after we left the station, however, and so the data from that cast, and station, was lost.

## 2.3.	Data Processing

Several steps were performed in the processing of each CTD cast.  These steps are outlined below:

1)	The raw values for pressure, temperature and conductivity were converted to engineering units using calibration coefficients determined in pre-trip laboratory calibrations.

2)	Each cast was processed to compensate for the differences in the time responses of the temperature sensor and the conductivity cell.  This was done by advancing the temperature data by 2 records (1 second) and the conductivity by .5 records (.25 seconds).  These values were determined experimentally by comparing up and down casts and by minimizing spikes in the computed salinity.

3)	Salinity was calculated from temperature and conductivity using the UNESCO algorithms [Fofonoff and Millard; 1983].

4)	Each cast was examined to determine if any freezing or any other problems affected the integrity of the cast.   All subsequent processing was done using only casts showing no problems with the data.

5)	Based on comparisons done between the CTD and bottle salinities, an offset of 0.0145 was added to all 1030 salinities and an offset of 0.0124 was added to all 1031 salinities.  See the section on data validation for more information.

6)	For all casts using the oxygen probe, dissolved oxygen was calculated using the pre-cruise calibration coefficients and the formulas provided by the manufacturer.  Part of the calculations involved advancing the oxygen channels by 12.5 seconds (25 records) to compensate for the sensors time response.  This value was determined by comparing up and down cast data for a number of stations.

7)	All dissolved oxygen values were adjusted based on bottle oxygen values determined by Winkler titration.  Table 3 lists the offsets added as a function of cast number.  See the section on data validation for more information.

8)	The data from all casts was then processed to remove all up-casts.  The pressures were also adjusted at this time so that the pressures at the surface would be zero.

9)	The top two meters were removed from each cast based on the assumption that this data is in the ice hole and not representative of the water column.

10)	The best profile was selected from those taken during each casts.  In most cases this was the first profile but in some cases a subsequent profile was chosen.

11)	A few casts were hand-edited to remove spikes.

12)	0.007 deg C. was added to the temperature data from all 1030 casts.  See the section on data validation for more information.

13)	The pressure, temperature, salinity, and oxygen data for each cast were decimated into one-meter bins using a simple averaging process.

14)	Derived oceanographic quantities were calculated from the pressure, temperature, and salinity data using the algorithms given in the UNESCO TECHNICAL PAPERS ON MARINE SCIENCE, No. 44 by Fofonoff and Millard.

The decimated data, and derived quantities, were used in producing the plots and tables found in Appendix 1.

## 2.4.	Data Validation

Validation of the CTD readings was done in four steps as described below

### 2.4.1.	Inspection of the Multiple Profiles Taken for Each Cast.

Each cast was plotted with all profiles overlapping and also with each channel as a function of time.  From these plots we were able to determine which casts had freezing or battery problems.   Any casts which showed problems are not included in this report.  Table 7 lists each of the casts, what problems, if any, were associated with that cast, and whether or not the cast is used in this report.

### 2.4.2.	Comparison of the CTD Data with the Bottle data for Salinity and Dissolved Oxygen

For each station in which bottle samples were taken for salinity and/or oxygen, comparisons were done between the CTD and bottle data.  Some bottle samples were collected from bottles hung on the line above the CTD and some were taken from separate bottle casts.

Figures 3 and 4 shows the differences between bottle and CTD salinities as a function of depth for CTD's 1030 and 1031.  From this data we decided to adjust the salinities from each CTD by a constant so as to give the residuals a zero mean.   Table 2 gives the offsets added to the salinities and the standard deviation of the residuals.

| CTD S/N | Constant Added to all Salinities | Standard Deviation of Residuals | Count |
| --- | --- | --- | --- |
| 1030 | 0.0145 | .011 | 20 |
| 1031 | 0.0124 | .007 | 22 |

Table 2 - Salinity Corrections and Standard Deviations

Figure 3 - CTD 1030 Bottle-CTD Salinity Residuals

Figure 4 - CTD 1030 Bottle-CTD Salinity Residuals

Figure 5 shows the differences between the CTD and bottle dissolved oxygen values as a function of date for CTD 1030.  The oxygen probe membrane was changed three times during the field trip.  Those three events show clearly on the figure.  We made the assumption that the dissolved oxygen values will drift linearly as a function of time for a given membrane, and that there is no way to interpolate from one membrane to the next.  Based on these assumptions, three least squares fits were done to the data to calculate an offset, as a function of date.  These offsets, and the casts that they affect are listed in Table 3.   Table 4 gives the standard deviation of the residuals of the corrected dissolved oxygen values.

Note: the assumptions made in correcting the oxygen data are not necessarily valid.  There is no explanation for the fact that for two membranes there was an increase in reading from one cast to the next and for a third membrane there was a decrease.

| Date | Cast #'s | Oxygen Offset (ml/l) |
| --- | --- | --- |
| April 14 | 1,2 | .985 |
| April 17 | 3,4,6,7 | -.681 |
| April 19 | 8,9,10,11,12 | -1.791 |
| April 20 | 13,14,15,16,17,18,19,20 | -2.347 |
| April 23 | 21,22 | 1.768 |
| April 24 | 24 | 1.690 |
| April 25 | 25 | 1.612 |
| April 26 | 27,28 | 1.533 |
| April 27 | 29,30 | 1.455 |
| May 3 | 43 | .984 |
| May 4 | 45,46,49 | -1.556 |
| May 5 | 50,51 | -1.373 |
| May 7 | 54,55 | -1.006 |
| May 8 | 62,63,64,65,66 | -0.823 |
| May 9 | 67,68,69 | -0.640 |
| May 11 | 70,72 | -0.273 |
| May 12 | 74 | -2.900 |

Table 3 - Offset Corrections Made to Dissolved Oxygen Data

| Standard Deviation of Residuals (ml/l) | Count |
| --- | --- |
| 0.35 | 120 |

Table 4 - Dissolved Oxygen Standard Deviation

Figure 5 - Bottle-CTD Dissolved Oxygen Residuals

### 2.4.3.	Intercomparision of the Two CTD's for Temperature and Salinity.

Comparisons were made between the temperature and salinity data from CTD's 1030 and 1031.   The comparisons were made for a select group of stations in which the two CTD's were used profile the water column within 1 hour of each other.  The stations, and casts numbers used are listed in Table 5.

| Station | 1030 Cast | 1031 Cast |
| --- | --- | --- |
| T100 | 25 | 26 |
| E100 | 45 | 44 |
| E400 | 46 | 47 |
| E800 | 49 | 48 |
| MB92 | 70 | 71 |
| E300 | 72 | 73 |

Table 5 - Casts Used for CTD Intercomparisons

The salinities were compared after being adjusted based on the bottle salinities. One half meter averaged data was calculated for both temperature and salinity and these values were compared.  The temperature difference (1030-1031) and salinity difference (1030-1031) were calculated at each depth.  The mean and standard deviations were then calculated for each temperature difference where DT/DP (for 1030) did not exceed 0.01 deg C/m and for each salinity difference where DS/DP (for 1030) did not exceed 0.01 PSU/m.  Figures 6 and 7 show plots of these temperature and salinity differences as a function of pressure.  Table 6 lists the mean and standard deviations for the Temperature and Salinity Differences

Figure 6 - Temperature Differences (1030-1031) vs. Pressure

Figure 7 - Temperature Differences (1030-1031) vs. Pressure

| Quantity | Mean | Std. Dev. | Count |
| --- | --- | --- | --- |
| Temp Diff (1030-1031) deg C | -.0071 | .0025 | 132 |
| Sal. Diff (1030-1031) PSU | -.0024 | .0054 | 128 |

Table 6 - Mean and Standard Deviations of the Temperature and Salinity Differences (1030-1031)

From the intercomparison of the two CTD's we can say that the two sets of CTD salinity data (adjusted to the bottle salinities) have no measurable difference between them, within a tolerance of 0.005 PSU.  The two sets of CTD temperature data, however, have a measurable difference of  0.007 +/- .002 deg C

### 2.4.4.	Comparison of the Pre-trip and Post-trip Laboratory Calibrations.

In order to explain the differences in the temperatures measured between the two CTD's.  It was necessary to look at temperature calibrations done on the instruments before and after the trip.

Figure 8 shows temperature residuals (CTD-Bath) for CTD 1030 based on a February 20,1992 calibration.  Residuals are shown for both the February 20,1992 calibration (pre-trip)  and a February 19,1993 calibration check (post-trip).  The two calibrations show an approximate 0.007 degree decrease in the temperature readings between the pre-trip and post-trip calibrations.

Figure 9 shows temperature residuals (CTD-Bath) for CTD 1031 based on a March 4,1992 calibration.  Residuals are shown for both the March 4,1992 (pre-trip) and November 5,1992 (post-trip) calibrations.   The two calibrations show no measurable offset in the temperature readings between the two calibrations.

The laboratory calibrations, in conjunction with the CTD intercomparisons, suggest that an event occurred between when 1030 was calibrated and the start of the field trip (perhaps a jolt in transporting the CTD) that caused a 0.007 deg C decrease in temperature readings that persisted from the start of the trip to the post-trip calibration.  They also suggest that the temperature readings from 1031 remained stable from the pre-trip calibration through the duration of the field trip.

Based on these measurements,  a value of 0.007 deg C. was added to all temperature data for CTD 1030.

Figure 8 - 1030 Bath Calibration Temperature Residuals (CTD-Bath) for Pre-Trip and Post-Trip Calibrations

µ §

Figure 9 - 1031 Bath Calibration Temperature Residuals (CTD-Bath) for Pre-Trip and Post-Trip Calibrations

# 3.	List of Stations

Table 7 lists all of the CTD casts taken during the field trip and the stations at which they were taken.   The table also lists which casts are valid and which ones are used in this report.

| Cast # | Station | CTD S/N | Trans-port * | Bottle on line with CTD? | Chem-istry Done At Stat-ion? | Valid? | Use? | NOTES |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | E600 | 1030 | T |  |  |  |  | Bad Pump |
| 2 | E800 | 1030 | T |  |  |  |  | Bad Pump |
| 3 | E800 | 1030 | T |  |  |  |  | Bad Pump |
| 4 | E800 | 1030 | T |  |  | ü | ü | Only one profile |
| 5 | R292 | 1031 | T |  |  |  |  | Bad Pump |
| 6 | R292 | 1030 | T |  |  |  |  | Bad Pressure |
| 7 | R192 | 1030 | T |  |  |  |  | Bad Pressure |
| 8 | E100 | 1030 | H |  |  | ü | ü |  |
| 9 | E100 | 1030 | H | ü |  | ü |  | Not a complete profile, for bottle comparison only |
| 10 | E200 | 1030 | H |  |  | ü | ü |  |
| 11 | E200 | 1030 | H | ü |  | ü |  | Not a complete profile, for bottle comparison only |
| 12 | E300 | 1030 | H |  |  | ü | ü | Did not find 18m basin. Cast done in only 5m of water |
| 13 | E400 | 1030 | H |  |  | ü | ü |  |
| 14 | E400 | 1030 | H | ü |  | ü |  | Not a complete profile, for bottle comparison only |
| 15 | E500 | 1030 | H |  |  | ü | ü |  |
| 16 | E600 | 1030 | H |  |  |  |  | Poor salinity intercomparison between profiles |
| 17 | E630 | 1030 | H |  |  | ü | ü |  |
| 18 | E900 | 1030 | H |  |  |  |  | Bad Pumping |
| 19 | E800 | 1030 | H |  |  | ü | ü |  |
| 20 | E700 | 1030 | H | ü |  | ü | ü |  |
| 21 | T100 | 1030 | S |  |  |  |  | Bad Pressure |

* T=Twin Otter, H=Helicopter, S=Snow Mobile

Table 7 - CTD Cast Information

| Cast # | Station | CTD S/N | Trans-port * | Bottle on line with CTD? | Chem-istry Done At Stat-ion? | Valid? | Use? | NOTES |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 22 | T100 | 1030 | S | ü |  |  |  | Bad Pressure |
| 23 | T100 | 1031 | S |  |  |  |  | Bad Pressure |
| 24 | T200 | 1030 | S |  |  | ü | ü |  |
| 25 | T100 | 1030 | S |  |  | ü | ü |  |
| 26 | T100 | 1031 | S |  |  | ü | ü |  |
| 27 | ROB0 | 1030 | S |  |  |  |  | Bad Pumping |
| 28 | E300 | 1030 | S |  | ü |  |  | Bad Pumping |
| 29 | E900 | 1030 | T |  | ü |  |  | Bad Pumping |
| 30 | E900 | 1030 | T |  | ü | ü | ü |  |
| 31 | E230 | 1031 | S |  |  | ü | ü | Evidence of a lot of turbulence |
| 32 | E220 | 1031 | S |  |  | ü | ü |  |
| 33 | E110 | 1031 | S |  |  |  |  | Batt failed during profile, did not measure entire water column |
| 34 | E200 | 1031 | S |  | ü |  |  | Batt Fail |
| 35 | E200 | 1031 | S |  | ü | ü | ü |  |
| 36 | E210 | 1031 | S |  |  | ü | ü |  |
| 37 | E660 | 1031 | S |  |  | ü | ü |  |
| 38 | E650 | 1031 | S |  |  | ü | ü |  |
| 39 | E640 | 1031 | S |  |  | ü | ü |  |
| 40 | E630 | 1031 | S |  | ü | ü | ü |  |
| 41 | E620 | 1031 | S |  |  | ü | ü |  |
| 42 | E600 | 1031 | S |  |  | ü | ü |  |
| 43 | E500 | 1030 | T |  | ü | ü | ü |  |
| 44 | E100 | 1031 | T |  | ü | ü | ü |  |
| 45 | E100 | 1030 | T |  | ü | ü | ü |  |
| 46 | E400 | 1030 | T | ü | ü | ü | ü |  |
| 47 | E400 | 1031 | T |  | ü | ü | ü |  |
| 48 | E800 | 1031 | T | ü | ü | ü | ü |  |
| 49 | E800 | 1030 | T |  | ü | ü | ü |  |
| 50 | E670 | 1030 | T |  | ü | ü | ü |  |
| 51 | E200 | 1030 | T |  |  | ü | ü |  |
| 52 | E665 | 1031 | H |  |  | ü | ü | No Pump |
| 53 | E663 | 1031 | H |  |  | ü | ü | No Pump |

* T=Twin Otter, H=Helicopter, S=Snow Mobile

Table 7 (con't) - CTD Cast Information

| Cast # | Station | CTD S/N | Trans-port * | Bottle on line with CTD? | Chem-istry Done At Stat-ion? | Valid? | Use? | NOTES |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 54 | R192 | 1030 | T |  |  | ü | ü |  |
| 55 | R292 | 1030 | T |  |  | ü | ü |  |
| 56 | E400 | 1031 | S |  |  | ü | ü | No Pump |
| 57 | E370 | 1031 | S |  |  | ü | ü | No Pump |
| 58 | E360 | 1031 | S |  |  | ü | ü | No Pump |
| 59 | E350 | 1031 | S |  |  | ü | ü | No Pump |
| 60 | E341 | 1031 | S |  |  | ü | ü | No Pump |
| 61 | E330 | 1031 | S |  |  | ü | ü | No Pump |
| 62 | E225 | 1030 | T |  |  | ü | ü | Second profile used |
| 63 | E300 | 1030 | T |  |  | ü | ü |  |
| 64 | E661 | 1030 | T |  |  | ü | ü |  |
| 65 | E662 | 1030 | T |  |  | ü | ü |  |
| 66 | E670 | 1030 | T |  |  | ü | ü |  |
| 67 | E300 | 1030 | T |  |  | ü | ü |  |
| 68 | SITL | 1030 | T |  |  | ü | ü |  |
| 69 | SITL | 1030 | T |  |  | ü | ü |  |
| 70 | MB92 | 1030 | T |  | ü | ü | ü |  |
| 71 | MB92 | 1031 | T |  | ü | ü | ü |  |
| 72 | E300 | 1030 | T |  | ü | ü | ü |  |
| 73 | E300 | 1031 | T |  | ü | ü | ü |  |
| 74 | E800 | 1030 | T |  |  | ü | ü |  |

* T=Twin Otter, H=Helicopter, S=Snow Mobile

Table 7 (con't) - CTD Cast Information

From the list of valid casts, a representative set has been chosen which includes one cast from each station.  These casts are listed in Table 8.  In the authors opinion, these are the casts that should be used for further analysis such as contour plots and volumetric analysis.

| Station | Cast# | Station | Cast# | Station | Cast# |
| --- | --- | --- | --- | --- | --- |
| E100 | 45 | E400 | 46 | E700 | 20 |
| E110 | 33 | E500 | 15 | E800 | 49 |
| E200 | 51 | E600 | 42 | E900 | 30 |
| E210 | 36 | E620 | 41 | MB92 | 70 |
| E220 | 32 | E630 | 40 | R192 | 54 |
| E225 | 62 | E640 | 39 | R292 | 55 |
| E230 | 31 | E650 | 38 | SITL | 69 |
| E300 | 72 | E660 | 37 | T100 | 25 |
| E330 | 61 | E661 | 64 | T200 | 24 |
| E341 | 60 | E662 | 65 |  |  |
| E350 | 59 | E663 | 53 |  |  |
| E360 | 58 | E665 | 52 |  |  |
| E370 | 57 | E670 | 66 |  |  |

Table 8 - Representative CTD Casts for the Survey Area

# 4.	Appendix A. - CTD Handling Guidelines Used During Field Trip 92-07

NOGAP Sea-Bird CTD Handling Procedures

Taking a Cast

1)	Keep the CTD in the aircraft and warm until just before the cast

2)	If the pump has been allowed to freeze then flush the pump with warm 40ppt water using the large syringe attached to the pump outlet tube.

3)	Fill the conductivity cell (and DO sensor) side of the plumbing with warm 40ppt water as follows:

-	Disconnect the left tube from the bottom of the "Y" at the top of the CTD plumbing.

-	Fill the large syringe kept in the CTD Field Kit with 40ppt water.  You will need at least 35 ml of water for CTD 1031 and 50 for CTD 1030.

-	Attach the syringe tube to the mouth of the conductivity cell and inject the water, insuring that the water level reaches the top of the tube you disconnected.

-	Put your thumb over the tube and remove the syringe.

-	Plug the cell mouth using the small stopper attached to the string then clip the other side of the string to the top of the CTD so that it will be handy.

-	Reconnect the tube to the bottom of the "Y"

4)	Make sure that the pin hole at the top of the "Y" is clear using one of the small wires supplied in the field kit..

5)	Clear all ice from the hole

6)	Take a sounding using the lead line.  With the lead weight on the bottom and the line pulled taught, attach a clip to the line just at water level.  This clip will serve as your bottom mark when doing the cast.  Bring up the lead weight and unclip it from the line.

7)	Take the CTD to the hole, clip it to the line, TURN IT ON, and lower it until the top of the "Y" is just submerged.  You should then see bubbles rising from the small pinhole at the top of the "Y".  This indicates that the plumbing system is filling with water properly.

8)	NOTE:	If you do not see bubbles at this time, quickly remove the CTD from the hole, TURN IT OFF, and clear the pinhole using one of the small wires available in the field kit.  If the weather is really cold you may have warm the CTD in order to do this.  Then start step 7 again.

9)	Gently pull on the string to remove the plug from the end of the conductivity cell (several short, light tugs should dislodge the plug).

10)	Lower the CTD to a depth of 15 meters, or just above the bottom (if depth < 15m) and let it soak for 5 minutes.  The pump will come on 4 minutes after the CTD was turned on.

11)	Bring the CTD back to the surface so that the top of the "Y" is just below the water (NEVER remove the CTD from the water at this point) and then start the profile, lowering the CTD at 4 seconds per meter.  Lower the CTD until the clip attached to the line in step 5 reaches the water level.  This will stop the CTD 1/2 meter above the bottom.

12)	Repeat step 10 (do a second complete profile)

13)	Do a third complete profile as in step 10.  If you are using CTD 1030 (the CTD with the DO sensor), however, modify the profile by stopping the CTD for 30 seconds four times during the cast at the 1/4, 1/2,  3/4 and bottom depths.  This will allow the DO sensor to catch up and give us four good points to pull the rest of the data to.

14)	Remove the CTD from the water and let the plumbing completely drain.  TURN IT OFF, and take it back to the plane.

Doing a bottle calibration with the CTD

1)	Select the depth, D (from the TOP of the ice), that you want to take the bottle at.  For this example we will use 40 meters.

2)	Perform  steps 1-6 of "Taking a Cast"  (page 2) to get the CTD ready.

3)	Attach the Nisken bottle at the 2 meter mark on the line (this is the first green line above the shackle).  Put the CTD in the water as per step 7 of  "Taking a Cast".

4)	Lower the CTD until the sounding line reaches the D -1 meter mark (39 meters in our case) and let the CTD soak for 5 minutes.

5)	Lower the CTD at 4 seconds per meter to the D+2 meter mark (or 42 meters).

6)	Let the bottle flush at that depth then release the messenger.

7)	Bring the bottle and CTD to the surface.  Remove the CTD from the water and let the plumbing completely drain.  TURN IT OFF, and take it back to the plane.  Then take your samples from the Nisken bottle.