Centre for Ocean Climate Chemistry

Cruise 9203

Freon Data Report

Wendy Richardson

Centre for Ocean Climate Chemistry

Institute of Ocean Sciences

Sidney, B.C., Canada

July 1993

# µ

Introduction	3

Methods	3

Sampling	3

Water	3

Air	4

Analysis	4

Calibration	4

Standards	4

Daily Calibration	4

Results	5

Positions of Hydrochemical Stations with Freon Profiles	5

Station Data (Analyzed in ship laboratory)	6

Comparison: Station HS14 vs. HS23	7

Station HS01- August 15, 1992	8

Station HS12- August 18, 1992	10

Station HS13- August 19, 1992	12

Station HS14- August 20, 1992	14

Station HS15- August 21, 1992	16

Station Data (Analyzed in container on aft deck)	18

Station HS17- August 24, 1992	18

Station HS18- August 25, 1992	20

Station HS19- August 26,1992	22

Station HS20 - August 27, 1992	24

Station HS21- August 29, 1992	26

Station HS22- August 30, 1992	28

Station HS23- September  1, 1992	30

Tests	32

Bottle Blanks	32

Precision Test	33

Precision by Pooled Standard deviation of Duplicates	34

Restrips of water samples	35

System Temperature Fluctuations in Container	35

Baseline Zero Values	36

Air	37

Freon Levels of Outside Air	37

Freon Levels in Container Air	37

Discussion	38

Recommendation from Bullister (1991)	38

Problems	38

Conclusions	39

Recommendations	39

References	39

# Introduction

The August September 1992 COCC cruise to the North west Pacific Ocean and Bering Sea was the first time the Freon Analysis System had been shipped across the world and set up on a Foreign vessel.  Freon sampling and analyses were attempted at all hydrochemical stations.

Due to phenomenally high Freon levels in the lab air (estimate of  over a hundred parts per million as opposed to the normal hundreds of parts per trillion) it was deemed necessary to move the Freon analysis system out to the metal shipping container on the aft deck of the ship.  Analyses of hydrocasts taken from stations HS17, HS18, HS19, HS20, HS21, HS22, and HS23 were carried out in the container.

# Methods

## Sampling

### Water

Freon samples were taken from 10 liter Niskin bottles only since the large bottles have a better O-ring surface area/water volume ratio than smaller Niskins.  These bottles were fairly well aged, but no specific precautions had been taken to lessen contamination of Freon from O-rings or elastic tubing in these Niskin bottles.  They were however, stored out on deck where the air is not so heavily contaminated with CFC's.  The first samples drawn from the Niskin were the Freon samples since they are the most easily contaminated by air entering the head space of the bottle.  The samples were drawn directly into 100 cm precision ground glass syringes.

Once the Niskin bottle was on deck, the integrity of the bottle was verified by pushing in the drain valve while the vent remained closed.  If there was no leakage, the syringe valve was inserted into the drain valve and the vent opened.  The syringe was filled with about 10 cc. of seawater  The hydrostatic pressure of the water in the Niskin bottle forced the syringe plunger outward thus reducing the chance of drawing an air bubble into the syringe which is likely if the plunger was pulled out.  The syringe was then removed while the water was still flowing and emptied into the outflow (but not into the Niskin).  This 10 cc rinse was repeated three times and then the syringe was filled with approximately 100cc, inverted and plunger spun to send any bubbles to the top. while pushing out bubbles and sample.  The three 10 cc rinse technique was repeated and the syringe then filled to about 100cc; stopcocks closed.  An elastic was put on the syringe to maintain a positive pressure on the water sample.  The syringe was placed in the sample bucket which had at least one inch of recently drawn seawater above the syringe barrel.  If an air bubble was drawn into the syringe while sampling, the rinses were started all over again.

Initially, the bucket of samples was kept in a sink in the lab, with a continual overflow of seawater from the seawater loop manifold. However, after discovering the extreme contamination of the lab air and how quickly the seawater in the bucket picked up the contamination, the storage procedure of samples was changed.  Water for the bucket was taken from three Niskins from the bottom of the first cast.  The first cast at station was the shallow cast so the water was collected at a depth below 600 to 1000 meters. The buckets were stored out on the back deck with a cover of plastic in an effort to keep some of the smoke stack soot out of the bucket.

Bottle blanks of fourteen Niskins, used to sample 600 meters and below, were evaluated with 2000 meter water at station HS23.  Unfortunately numbered Niskins were not designated to specific depths, therefore requiring many bottle blanks and hampering any evaluation of possible Niskin contamination of samples.

Precision of analysis was checked at stations HS22 and HS23.  At station HS22 three Niskins were tripped at 95, 100, and 105 meter depths and at station HS23 four Niskins were tripped at 85, 90, 95, and 100 meters.  Three subsamples from each Niskin were drawn and analyzed.

Carrier blanks, stripper blanks, restriped samples, lab air, container air, sampling deck air, and Foredeck ("clean") air were also analyzed throughout the cruise.

### Air

Syringe air samples were taken from the Foredeck, sampling deck, inside the lab, and inside the container. The designated air syringes were flushed at least ten times then a sample drawn, stopcock closed, and tension elastic put in place.

## Analysis

Analysis was carried out using the  Chloroflorocarbon Analytical System built by Glenn Smith,  Centre for Ocean Climate Chemistry, Institute of Ocean Sciences.  The COCC analytical system is a modified version of the one described by Bullister and Weiss (1988).

## Calibration

### Standards

Working standard, tank number 63098 was used throughout this cruise. (F11,  250.300 ppt, standard deviation 0.87; and F12,  143.50 ppt, standard deviation 0.28) This standard was made up of outside air and nitrogen by Glenn Smith using the COCC facilities.  The tank was calibrated against COCC's  lab standard tank number 63088 (F11,  445.41 ppt, standard deviation 6.95; and F12,  257.3 ppt, standard deviation 0.88).  This lab standard in turn had been calibrated against Ocean Chemistry's Freon standard tank number 70016 (F11,  271.0 ppt, standard deviation 0.42; and F12, 511.8 ppt, standard deviation 0.69).  The Ocean Chemistry Freon standard was obtained and certified by Weiss's lab.

### Daily Calibration

Calibration curves used to determine sample concentrations were generated by multiple injections of predetermined volumes of standard gas.  Verifying the carrier blank would be the first procedure of the day.  This was followed by a suite of S(small loop)*1, S*2, S*3, S*4, and L(large loop)*1, L*2, L*3.  If these areas looked reasonable based on the previous days results, a stripper blank was analyzed and then water sample analyses begun.  A L*1 standard was run every four to six samples, approximately hourly, throughout the day to monitor drift in detector sensitivity.

# Results

Over five hundred and fifty Freon analysis runs were carried out during this cruise.  A total of two hundred and fifty water samples were run.

Data reduction was carried out using the Ocean Chemistry/NOGAP PC adapted computer program based on the Scripps computer program (Weiss).  This program requires salinity and temperature for calculations.  The former was taken from CTD cast data for depths down to 2000m, below this depth Salinity was taken from the Salinometer data.  Temperature was read from the sample bucket when the syringe was removed and attached to the extraction system.  The Instrument detection limit while set at threshold 2, was found to be F11 x<0.013 (Pmol/Kg) and F12 x<0.038 (Pmol/Kg).

There are two weak points  in the calculation procedure.  The first is that the caliper readings may not be as consistent as desired, due to the variation of the curve of the glass syringe.  Secondly, a slight variation of second aliquot results could be due to the temperature change as the sample warms up in the air; since the temperature used for calculation is that of the sample storage bucket.

This is an informal data report only. Because of the nature of the Freon data reduction program, confirming recalculation of the results would be advisable before considering results for publication.  As well, there has been no decision about sample handling blanks and bottle blanks.  Therefore, the results presented in this report are in raw form with no corrections. Until a standard protocol has been decided upon, no corrections shall be applied to data report results.

## Positions of Hydrochemical Stations with Freon Profiles

| Station | Date 1992 | Latitude North | Longitude East/West |
| --- | --- | --- | --- |
| HS01 | August 15 | 45 59 | 146 56 E |
| HS12 | August 18 | 45 00 | 155 00 E |
| HS13 | August 19 | 45 00 | 160 00 E |
| HS14 | August 20 | 44 59 | 165 02 E |
| HS15 | August 21 | 44 59 | 170 02 E |
| HS17 | August 24 | 45 00 | 176 49 W |
| HS18 | August 25 | 50 26 | 176 34 W |
| HS19 | August 26 | 53 43 | 178 47 E |
| HS20 | August 27 | 55 59 | 179 59 W |
| HS21 | August 29 | 54 30 | 172 30 E |
| HS22 | August 30 | 51 24 | 165 16 E |
| HS23 | September 1 | 44 58 | 165 02 E |

## Station Data (Analyzed in ship laboratory)

As stated previously, the contaminated ship had levels of Freon 12 estimated at over a hundred parts per million as opposed to the normal hundred of parts per trillion.  The Freon analysis could not be conducted properly in such an atmosphere.

The quality of data analyzed within the contaminated ship was poor.  It is included in the report only to indicate the possibly interesting levels of Freon at the first five stations.  The following comparison plot of a repeated station indicates that the F11 data may not be too inaccurate.  However this plot also demonstrates the F12 contamination of samples that were expected to be low level.

It should be noted that the carrier blanks for these stations were not zero, and that the carrier blank rules the standard curve in the Freon computer program.  There was also an incredible drift of the 1 large loop standard during these analysis.  The response of the GC varied greatly during this time as well.  The baseline drift was very inconsistent while the system was in the ship.

Some of these problems may have been due to a required long period for stabilization of the Freon system after reassembly and a two month transit period in a hot shipping container.  Weiss (1991) discusses the severe contamination problems encountered on their Moana Wave cruise which probably arose from the packing foams used for other equipment in their shipping container.  Their CFC measurement system was badly contaminated, as were their Niskin sampling bottles.  Weiss used an exponential decay function depending on bottle type to determine sample blank histories.

The Freon computer program allows the operator to exclude suspicious standard values while trying to generate the best fitting standard curve.  While computing the data from samples analyzed within the ship it was necessary to omit many of these "fliers", whereas during normal cruises it is very rare to omit any of the standards.

### Comparison: Station HS14 vs. HS23

Station HS14 was sampled August 20 and Freons were analyzed in the ship laboratory.  Station HS23 was sampled September 1 with Freons analyzed in the container.  The location for both was 45 North and 165 East.

µ §

### Station HS01- August 15, 1992

Calibration Curve, Percent Standard Error: F11 = 5.577; F12 = 5.861

Definitions:	n/c:	not computed	n/intg:	not integrated

s/c:	suspected contamination	p/intg:	poor integration

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 4.674 | 5.036 |
| 51 | 7.216 | 5.838 |
| 101 | 5.903 | 4.489 |
| 202 | 6.315 | 3.836 |
| 405 | 5.142 | 3.043 |
| 506 | 4.118 | 2.341 |
| 607 | 3.254 | 1.630 |
| 808 | 1.720 | 0.780 |
| 1013 | 0.760 | 0.403 |
| 1417 | 0.289 | 0.633 |
| 1618 | 0.180 | s/c |
| 1823 | 0.190 | 0.459 |
| 2279 | 0.154 | 0.654 |
| 2785 | 0.068 | 0.659 |
| 3038 | 0.177 | 1.480 |

µ §

### Station HS12- August 18, 1992

Calibration Curve, Percent Standard Error: F11 = n/c; F12 = 1.891

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 3.488 | 1.810 |
| 50 | 6.319 | 3.022 |
| 100 | 6.852 | 3.914 |
| 150 | 3.977 | 3.468 |
| 201 | 2.190 | 2.309 |
| 301 | 0.881 | s/c |
| 500 | 0.426 | 0.421 |
| 601 | 0.419 | 0.800 |

µ §

### Station HS13- August 19, 1992

Calibration Curve, Percent Standard Error: F11 =4.340; F12 = 21.01

The F12 results are not presented because of the absurdly high Percent Standard Error of 21.

| DEPTH (m) | F11 Pmol/Kg |
| --- | --- |
| 0 | 3.014 |
| 51 | 5.837 |
| 100 | 6.175 |
| 151 | 4.940 |
| 201 | 2.832 |
| 302 | 2.163 |
| 408 | 1.038 |
| 510 | 0.537 |
| 612 | 0.733 |
| 1220 | 0.114 |
| 1427 | 0.030 |
| 1640 | 0.100 |

µ §

### Station HS14- August 20, 1992

Calibration Curve, Percent Standard Error: F11 = 0.925; F12 = 5.491

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 3.330 | 1.952 |
| 50 | 5.671 | 2.887 |
| 101 | 5.349 | 2.626 |
| 152 | 4.891 | 2.283 |
| 204 | 3.618 | 1.667 |
| 304 | 2.687 | 1.470 |
| 304 | 2.595 | 1.445 |
| 405 | 1.713 | s/c |
| 509 | 1.350 | 1.116 |
| 610 | 0.547 | 0.833 |
| 812 | 0.311 | 0.899 |
| 1014 | 0.257 | s/c |
| 1621 | 0.085 | 0.902 |
| 1824 | 0.168 | s/c |

µ §

### Station HS15- August 21, 1992

Calibration Curve, Percent Standard Error: F11 = 0.580; F12 = 3.635

The sample results of the 150 meter Niskin demonstrate why the Freon samples must be drawn first.  The expected F12 result should have been between 1.630 Pmol/Kg (200m) and 2.915 Pmol/Kg (100m).  The Freon sample was not drawn first, and had an F12 value of 6.027 Pmol/Kg!

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 4.540 | 2.875 |
| 50 | 5.837 | 2.940 |
| 101 | 5.844 | 2.915 |
| 201 | 2.620 | 1.630 |
| 301 | 1.793 | 1.205 |
| 301 | 1.621 | 1.219 |
| 402 | 1.070 | 0.969 |
| 502 | 0.678 | 0.540 |
| 602 | 0.429 | p/intg |
| 804 | 0.257 | 0.296 |
| 1005 | 0.115 | 0.348 |
| 1206 | 0.102 | 0.352 |
| 1407 | 0.091 | s/c |
| 1609 | 0.052 | 0.277 |
| 1810 | 0.060 | 0.272 |

µ §

## Station Data (Analyzed in container on aft deck)

### Station HS17- August 24, 1992

Calibration Curve, Percent Standard Error: F11 = n/c; F12 = 0.39

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 3.515 | 1.727 |
| 50 | 4.914 | 2.182 |
| 101 | 4.935 | 2.304 |
| 152 | 4.604 | 2.305 |
| 203 | 3.248 | 1.574 |
| 305 | 1.844 | 0.906 |
| 305 | 1.879 | 0.979 |
| 406 | 1.534 | s/c |
| 508 | 0.811 | 0.445 |
| 611 | 0.383 | 0.272 |
| 814 | 0.200 | 0.262 |
| 1017 | 0.091 | 0.121 |
| 1221 | 0.056 | 0.088 |
| 1424 | 0.047 | 0.156 |
| 1628 | 0.037 | 0.105 |
| 2010 | n/intg | 0.068 |
| 2513 | 0.034 | 0.077 |
| 3010 | 0.028 | 0.114 |

µ §

### Station HS18- August 25, 1992

Calibration Curve, Percent Standard Error: F11 = 0.23; F12 = 0.67

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 4.813 | 2.574 |
| 49 | 4.899 | 2.295 |
| 99 | 5.330 | 2.761 |
| 149 | 1.941 | 1.165 |
| 199 | 1.265 | 0.643 |
| 299 | 0.746 | 0.437 |
| 299 | 0.617 | 0.346 |
| 400 | 0.450 | 0.403 |
| 500 | 0.292 | 0.306 |
| 831 | 0.159 | 0.334 |
| 1029 | 0.156 | 0.456 |
| 1221 | 0.086 | 0.209 |

µ §

### Station HS19- August 26,1992

Calibration Curve, Percent Standard Error: F11 = 1.07; F12 = 1.12

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 4.455 | 2.343 |
| 50 | 4.946 | 2.365 |
| 100 | 5.071 | 2.393 |
| 151 | 4.765 | 2.405 |
| 300 | 1.394 | 0.746 |
| 300 | 1.284 | 0.627 |
| 401 | 0.750 | 0.428 |
| 501 | 0.568 | 0.503 |
| 810 | 0.196 | 0.274 |
| 1010 | 0.121 | 0.282 |
| 1210 | 0.084 | 0.173 |
| 1611 | 0.065 | 0.167 |

µ §

### Station HS20 - August 27, 1992

Calibration Curve, Percent Standard Error: F11 = 1.42; F12 = 1.44

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 4.573 | 2.472 |
| 50 | 5.362 | 2.588 |
| 100 | 5.487 | 2.639 |
| 148 | 5.597 | 2.937 |
| 200 | 3.592 | 2.006 |
| 300 | 2.139 | 1.188 |
| 300 | 1.902 | 1.077 |
| 400 | 1.228 | 0.700 |
| 501 | 0.723 | 0.647 |
| 600 | 0.418 | 0.284 |
| 789 | 0.254 | 0.320 |
| 1000 | 0.142 | 0.277 |
| 1202 | 0.096 | 0.207 |
| 2506 | 0.032 | 0.068 |
| 3006 | 0.047 | 0.084 |

µ §

### Station HS21- August 29, 1992

Calibration Curve, Percent Standard Error: F11 = 1.07; F12 = 0.69

| DEPTH m | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 4.445 | 2.219 |
| 25 | 4.525 | 2.142 |
| 76 | 4.480 | 2.157 |
| 127 | 4.048 | 1.940 |
| 203 | 2.778 | 1.426 |
| 203 | 2.668 | 1.293 |
| 304 | 1.311 | 0.665 |
| 405 | 0.806 | 0.518 |
| 507 | 0.535 | 0.541 |
| 608 | 0.309 | 0.238 |
| 1000 | 0.143 | 0.346 |
| 1200 | 0.105 | 0.178 |
| 1600 | 0.054 | 0.186 |
| 2002 | 0.030 | 0.090 |
| 2002 | 0.044 | 0.107 |
| 2504 | 0.079 | 0.171 |
| 3005 | 0.040 | 0.176 |

µ §

### Station HS22- August 30, 1992

Calibration Curve, Percent Standard Error: F11 = 2.03; F12 = 2.55

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 4.397 | 2.468 |
| 50 | 6.155 | 2.894 |
| 150 | 4.364 | 2.153 |
| 200 | 2.495 | 1.340 |
| 300 | 1.545 | 0.859 |
| 300 | 1.393 | 0.764 |
| 400 | 0.874 | 0.534 |
| 492 | 0.569 | 0.497 |
| 599 | 0.302 | 0.238 |
| 795 | 0.189 | 0.265 |
| 997 | 0.118 | 0.259 |
| 1199 | 0.107 | 0.244 |
| 1392 | 0.075 | 0.167 |
| 1591 | 0.062 | 0.166 |
| 1989 | 0.045 | 0.177 |
| 2237 | 0.062 | 0.116 |
| 2734 | 0.048 | 0.086 |
| 2982 | 0.051 | 0.133 |

µ §

### Station HS23- September  1, 1992

Calibration Curve, Percent Standard Error: F11 = 1.10; F12 = 0.82

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 0 | 3.711 | 1.903 |
| 50 | 5.536 | 2.592 |
| 100 | 4.944 | 2.318 |
| 150 | 4.641 | 2.264 |
| 200 | 3.958 | 1.934 |
| 300 | 2.447 | 1.215 |
| 300 | 2.461 | 1.180 |
| 400 | 1.560 | 0.759 |
| 500 | 1.060 | 0.713 |
| 601 | 0.589 | 0.366 |
| 800 | 0.304 | 0.337 |
| 1002 | 0.204 | 0.226 |
| 1205 | 0.120 | 0.179 |
| 1407 | 0.103 | 0.193 |
| 1610 | 0.048 | 0.075 |
| 1812 | 0.061 | 0.115 |
| 2015 | 0.037 | 0.068 |
| 2015 | 0.033 | 0.075 |
| 2268 | 0.042 | 0.072 |
| 2522 | 0.050 | 0.096 |
| 2774 | 0.048 | 0.083 |
| 3027 | 0.035 | 0.053 |
| 3280 | 0.043 | 0.124 |
| 3534 | 0.055 | 0.217 |

µ §

## Tests

### Bottle Blanks

Only the deeper water Niskins were evaluated for bottle blanks since there is a greater percent bottle blank error for low value samples.  Thus the Niskins used for 600 meter sampling  and below were blank tested.  Unfortunately, on this cruise numbered Niskins were not designated to specific depths, therefore requiring many bottle blanks and hampering evaluation of possible Niskin contamination of samples. Niskins 3 and 13 were blank sampled, however the syringes had leaked thus precluding analysis.  Bottle blanks were done at Station HS23 on September 1, 1992 at a depth of 2000m which had a salinity of 34.581 parts per thousand.  The results demonstrate the problems of delayed analysis on low level samples.  Since the bottle blanks were done at the end of a station there was a delay while the profile samples were run.  Both Niskins number 5 and 12 had given lower values during deep water profiles.  Niskin 12 had values of F11 0.050, and F12 0.096, while Niskin 5 had an F12 of 0.075.

Calibration Curve, Percent Standard Error: F11 = 0.43; F12 = 0.91

| NISKIN | SYRINGE | RUN # | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- | --- | --- |
| 4 | 9839 | 512 | 0.103 | 0.271 |
| 5 | 5584 | 513 | 0.086 | 0.237 |
| 6 | 9838 | 514 | 0.058 | 0.150 |
| 7 | 5 | 515 | 0.045 | 0.107 |
| 8 | 3 | 517 | 0.050 | 0.106 |
| 9 | 24 | 518 | 0.054 | 0.109 |
| 10 | 16 | 519 | 0.041 | 0.060 |
| 11 | 9 | 520 | 0.045 | 0.065 |
| 12 | 29 | 522 | 0.152 | 0.433 |
| 14 | 15 | 523 | 0.048 | 0.103 |
| 15 | 7626 | 524 | 0.049 | 0.168 |
| 16 | 5547 | 525 | 0.078 | 0.252 |
|  |  | Mean | 0.067 | 0.172 |
|  |  | STD | 0.031 | 0.104 |

### Precision Test

Bottle depths for precision tests were chosen based on plots of CTD data in an attempt to find three depths of least variation.  However the "mixed" layer was not as homogeneous as desired as evident in the first test at station  HS22.  Again there was a problem with delayed analysis time, since the station profile had to be run first.  For station HS23, the Freon system was analyzing profile, precision and bottle blank samples from 2247 hours on August 31 until 2100 hours on September 1, 1993.  Long storage time before analysis is detrimental to the quality of the data.  It would be preferable not to conduct precision and blank tests at a station where a profile is also required.

| DATE | STN | NISKIN | SYRINGE | DEPTH M | SALINITY | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Aug 30 1992 | HS22 | 14 | 7626 | 95 | 33.232 | 5.940 | 2.871 |
| Aug 30 1992 | HS22 | 14 | 7633 | 95 | 32.232 | 5.992 | 2.967 |
| Aug 30 1992 | HS22 | 3 | 7677 | 100 | 32.240 | 5.832 | 2.953 |
| Aug 30 1992 | HS22 | 3 | 9828 | 100 | 32.240 | 5.963 | 2.972 |
| Aug 30 1992 | HS22 | 3 | 7634 | 100 | 32.240 | 5.911 | 2.834 |
|  |  |  |  |  | mean | 5.928 | 2.919 |
|  |  |  |  |  | STD | 0.055 | 0.056 |
|  |  |  |  |  | %STD | 0.923 | 1.925 |
| Aug 30 1992 | HS22 | 4 | 7631 | 105 | 32.240 | 5.683 | 2.777 |
| Aug 30 1992 | HS22 | 4 | 9823 | 105 | 32.240 | 5.672 | 2.754 |

| DATE | STN | NISKIN | SYRINGE | DEPTH M | SALINITY | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Sept 1 1992 | HS23 | 5 | 7745 | 85 | 33.454 | 5.297 | 2.534 |
| Sept 1 1992 | HS23 | 5 | 28 | 85 | 33.454 | 5.268 | 2.379 |
| Sept 1 1992 | HS23 | 5 | 0 | 85 | 33.454 | 4.978 | 2.274 |
| Sept 1 1992 | HS23 | 4 | 9942 | 90 | 33.454 | 4.849 | 2.282 |
| Sept 1 1992 | HS23 | 4 | 25 | 90 | 33.454 | 5.071 | 2.352 |
| Sept 1 1992 | HS23 | 4 | 5597 | 90 | 33.454 | 4.905 | 2.313 |
| Sept 1 1992 | HS23 | 6 | 7636 | 95 | 33.454 | 5.402 | 2.522 |
| Sept 1 1992 | HS23 | 6 | 13 | 95 | 33.454 | 5.118 | 2.364 |
| Sept 1 1992 | HS23 | 6 | 9829 | 95 | 33.454 | 5.237 | 2.627 |
| Sept 1 1992 | HS23 | 7 | 7677 | 100 | 33.454 | 5.366 | 2.752 |
| Sept 1 1992 | HS23 | 7 | 9828 | 100 | 33.454 | 5.338 | 2.787 |
|  |  |  |  |  | Mean: | 5.166 | 2.471 |
|  |  |  |  |  | STD; | 0.184 | 0.177 |
|  |  |  |  |  | %STD: | 3.565 | 7.148 |

### Precision by Pooled Standard deviation of Duplicates

| F11 X Pmol/Kg | F11 Y Pmol/Kg | F11 X-Y Pmol/Kg | F12 X Pmol/Kg | F12 Y Pmol/Kg | F12 X-Y Pmol/Kg |
| --- | --- | --- | --- | --- | --- |
| 2.687 | 2.595 | 0.092 | 1.470 | 1.445 | 0.025 |
| 1.793 | 1.621 | 0.172 | 1.205 | 1.219 | -0.014 |
| 1.844 | 1.879 | -0.035 | 0.906 | 0.979 | -0.073 |
| 0.746 | 0.617 | 0.129 | 0.437 | 0.346 | 0.091 |
| 1.394 | 1.284 | 0.110 | 0.746 | 0.627 | 0.119 |
| 2.139 | 1.902 | 0.237 | 1.188 | 1.077 | 0.111 |
| 2.778 | 2.668 | 0.110 | 1.426 | 1.293 | 0.133 |
| 1.545 | 1.393 | 0.152 | 0.859 | 0.764 | 0.095 |
| 2.447 | 2.461 | -0.014 | 1.215 | 1.180 | 0.035 |

µ §

Using the above equation to calculate the precision by the pooled standard deviation of duplicates the results were as follows.

F11  For the range 0.617 - 2.778 Pmol/Kg, the pooled standard deviation for duplicates was 0.599 Pmol/Kg.  During the previous Station Papa cruise five sets of duplicates with a range of 0.070 - 4.571 Pmol/Kg had a pooled standard deviation for duplicates of 0.075 Pmol/kg.

F12  For the range 0.346 - 1.470 Pmol/Kg, the pooled standard deviation for duplicates was 0.394 Pmol/Kg.  During the previous Station Papa cruise four sets of duplicates with a range of 0.074 - 2.090 Pmol/Kg had a pooled standard deviation for duplicates of 0.072 Pmol/Kg.

### Restrips of water samples

| STN | DEPTH | SAMPLE Run # | RESTRIP Run # | SAMPLE F11 Pmol/Kg | RESTRIP F11 Pmol/Kg | CARRY-OVER F11 % | SAMPLE F12 Pmol/Kg | RESTRIP F12 Pmol/Kg | CARRY-OVER F12 % |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HS18 | 800 | 293 | 294 | .159 | 0 | 0 | .334 | 0 | 0 |
| HS19 | 800 | 326 | 327 | .196 | 0 | 0 | .274 | 0 | 0 |
| HS21 | 0 | 392 | 393 | 4.445 | .014 | 0.3% | 2.219 | .0 | 0 |
| HS21 | 800 | 403 | 404 | .234 | 0 | 0 | .568 | 0 | 0 |
| HS22 | 200 | 428 | 429 | 2.495 | .013 | 0.5% | 1.340 | 0 | 0 |
| HS22 | 0 | 443 | 444 | 4.397 | .011 | 0.25% | 2.468 | 0 | 0 |
| HS23 | 2000 | 541 | 543 | .047 | 0 | 0 | .086 | 0 | 0 |

This table demonstrates the high stripper efficiency of the Freon analysis system on this cruise.

### System Temperature Fluctuations in Container

| Station | Date 1993 | Freon Loop Degree C | Sample Bucket Degree C |
| --- | --- | --- | --- |
| HS17 | Aug 24 | 14.0 to 17.4 | 9.9 to 10.1 |
| HS18 | Aug 25 | 15.0 to 16.6 | 5.7 to 8.8 |
| HS19 | Aug 26 | 12.5 to 15.0 | 6.0 to 8.4 |
| HS20 | Aug 27 | 12.1 to 14.1 | 5.1 to 8.4 |
| HS21 | Aug 29 | 12.2 to 21.3 | 6.0 to 8.5 |
| HS22 | Aug 30 | 15.5 to 24.6 | 8.8 to 11.8 |
| HS23 | Aug 31 | 16.5 to 25.8 | 8.0 to 15.3 |

This table shows the lower operating temperatures and wide temperature fluctuations that the Freon analysis system had to function under in the container

### Baseline Zero Values

| Location | Date | Run | Zero |
| --- | --- | --- | --- |
| IOS Lab | April 10 | 288 | 459.30 |
| IOS Lab | April 10 | 289 | 458.92 |
| IOS Lab | April 10 | 295 | 458.22 |
| IOS Lab | April 10 | 299 | 458.22 |
| Ship Lab | Aug 13 | 1 | 717.67 |
| Ship Lab | Aug 13 |  | 710.76 |
| Ship Lab | Aug 13 |  | 689.72 |
| Ship Lab | Aug 13 | 3 | 823.11 |
| Ship Lab | Aug 13 |  | 807.59 |
| Ship Lab | Aug 14 | 1 | 707.98 |
| Ship Lab | Aug 14 | 7 | 699.67 |
| Ship Lab | Aug 15 | 1 | 664.87 |
| Ship Lab | Aug 15 | 16 | 638.43 |
| Ship Lab | Aug 15 | 17 | 618.93 |
| Container | Aug 24 | 246 | 495.77 |
| Container | Aug 24 | 249 | 496.12 |
| Container | Aug 25 | 266 | 503.50 |
| Container | Aug 26 | 296 | 500.56 |
| Container | Aug 31 | 458 | 497.72 |
| Container | Aug 31 | 463 | 501.16 |

The above table shows the difficulty with the baseline zero while in the ship laboratory.

## Air

### Freon Levels of Outside Air

| DATE | STN | LATITUDE | LONGITUDE | RUN # | SAMPLE | F11 ppt | F12 ppt |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Aug 16 1992 | HS01 | 45'55.9N | 146'57.4E | 17 | Fore deck | 267 | 538 |
| Aug 20 1992 | HS14 | 44'59.3N | 165'03.5E | 140 | Fore deck | 312 | 629 |
| Aug 20 1992 | HS14 | 44'59.3N | 165'03.5E | 141 | Sampling deck | 298 | 458 |
| Aug 23 1992 | HS16 | 44'59.8N | 175'01.1E | 222 | Fore deck | 267 | 547 |
| Aug 24 1992 | HS17 | 45'02.5N | 176'49.4W | 245 | Fore deck | 253 | 417 |
| Aug 25 1992 | HS18 | 50'26.8N | 176'34.5W | 273 | Fore deck | 285 | 547 |
| Aug 25 1992 | HS18 | 50'26.8N | 176'34.5W | 286 | Fore deck | 255 | 431 |
| Aug 26 1992 | HS19 | 53'43.5N | 178'47.2E | 321 | Fore deck | 270 | 633 |
| Aug 27 1992 | HS20 | 55'58.9N | 179'56.4W | 352 | Fore deck | 292 | 511 |
| Aug 29 1992 | HS21 | 54'29.8N | 172'30.1E | 388 | Fore deck | 273 | 516 |
| Aug 30 1992 | HS22 | 51'24.0N | 165'16.5E | 445 | Fore deck | 257 | 495 |
| Sept 1 1992 | HS23 | 44'58.6N | 165'02.2E | 468 | Fore deck | 258 | 494 |

### Freon Levels in Container Air

| DATE | STN | LATITUDE | LONGITUDE | RUN # | SAMPLE | F11 | F12 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Aug 23 1992 | HS16 | 44'59.8N | 175'01.1E | 223 | Container | 370 | 3039 |
| Aug 24 1992 | HS17 | 45'02.5N | 176'49.4W | 260 | Container | 305 | 2189 |
| Aug 25 1992 | HS18 | 50'26.8N | 176'34.5W | 295 | Container | 272 | 2733 |
| Aug 26 1992 | HS19 | 53'43.5N | 178'47.2E | 306 | Container | 302 | 5782 |
| Aug 27 1992 | HS20 | 55'58.9N | 179'56.4W | 340 | Container | 283 | 4888 |
| Aug 27 1992 | HS21 | 54'29.8N | 172'30.1E | 406 | Container | 279 | 2722 |

# Discussion

## Recommendation from Bullister (1991)

"The trapping times required for multiple injections of standard gas can be longer than that required for a single injection of standard and for the extraction of CFC's from a seawater sample. For multiple injections of standards, the effect of losses of F11 and F12 by migration through the trap, broadening the chromatographic peaks due to long trapping times, and the accumulation of analytical blanks should be evaluated."(Bullister, 1991).  A new standard had been made up by Glenn Smith just before the cruise.  This new standard required multiple injections and the trapping times were not modified or evaluated before the system was shipped.

## Problems

One of the problems that arose due to the contamination in the ship lab, required a change in procedure for analysis of samples.  There appeared to be a resilient contamination problem with the tubing connecting the syringe and buret.  This meant that the first sample of the day was a higher value than it should have been even though the stripper blank had been zero.  The way around this problem was to load a dummy sample (an extra draw or one from the previous cast), analyze, restrip and then load the actual sample.

Attempting to train other scientists how to run the Freon analysis system while conducting an intense sampling schedule and combating the extreme contamination problems was absurd.  Training should involve enough time for specific tests of precision and handling by each analyst before sample analysis.  Training people with actual samples is unreasonable and detrimental to the quality of data.  Also it would have been much better to have only one new analyst dedicated to the Freon system instead of two who were also running another completely different analysis system that took much of their time and attention.  The unavailability of both new analysts to assist the primary operator while sampling also greatly hampered the desired speed of sampling and immediate analysis to reduce the storage time of samples.

Most of the analysis was performed in a shipping container on the back deck of the vessel in order to avoid the Freon contamination within the ship. This is not a suitable work environment for electronic instruments or the operator. This was only attempted in order to verify the possibility of obtaining reasonable results if an external lab module was provided on a ship know to be contaminated.

# Conclusions

## Recommendations

As stated previously no particular precautions, other than storing the Niskin bottles on deck, were employed to reduce the Freon levels of the Niskin bottles. If it is considered important to lower our deep water results and blanks, an attempt to treat the Niskins as described by Bullister and Weiss (1987) should be made. This entails replacing the elastic tubing with epoxy coated stainless steel springs, baking the O-rings for several days and then storing them in a gas tight container until installed on the bottles at sea.

As demonstrated at station HS15 Freon samples must be drawn first.

Samples cannot be analyzed in ppm level contaminated environments.

A shipping container on deck is not a suitable laboratory because of the humidity, corrosive atmosphere, temperature variations, and possible contamination from packing foams used for other equipment shipped in the container (Weiss 1991).

Precision and blank tests should not be added to a station requiring a profile.  For quality results each test should be sampled and quickly analyzed under optimal conditions.

# References

Bullister J.L., Weiss R.F. (1987)  Determination of CCl3F and CCl2F2 in Seawater and Air. Deep-Sea Research Vol. 35 No. 5, 839-853.

Bullister J.L., (1991) Chlorofluorocarbons, 3He-tritium and Small Volume Radiocarbon. WHP Operations and Methods, July 1991, 1-9.

Weiss R, Van Woy, R. (1991)  CFC-11 and CFC-12 measurements made during Moana Wave cruise 893, WHP Data Reporting Requirement, July 1991, 59-61.