## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 22-Feb-2012 | - Added one productivity cast, containing Primary Productivity, Chlorophyll, Nutrients, POC and PON, from Frank Whitney’s Productivity spreadsheet. - Added CFC profile data from John Page’s spreadsheet file that he had produced for Steve Diggs at CDIAC (linep_1993-03_cast_surf_bottle_data_02.xls). - Merged DIC data to bottle cast numbers 19, 23, 27 and 35 from Jim Christian’s DIC file located in the \DOC directory. For more information, including quality control analysis see John Page’s spreadsheet files in directory \\OSD_Data_Archive\Cruise_Data\Documents\Station Papa Data Archiving\Alkalinity and DIC. |

Processing notes

Cruise 93-03

March 7 - 15, 1993

33 original files, 28 processed files.

Casts 4 - 37, 99

There were 3 CTD probes used on this cruise.  The physics Guildline probes

58,483 and 53501, and the WOCE Guildline probe 59901, (model 8737).

The physics Guildline probes 58,483 and 53501 were used only for casts

#5 to #9.  The data were bad, John Love thought the problem was the probe,

so he tried the different CTDs.  He finally found that the problem was in

the deck unit.  The WOCE probe 59901 was used for the casts #4 and #10

to #37, plus #99.  The casts #1 to #3 were tests in Saanich Inlet.

1. Casts 4, 10-37, and 99

The .CNV files that Ron Perkin had produced where converted to .RAW files

using program WOCE_CNV.

Casts 6-9

These were already in .RAW format.

2. Calibrations (Program Calib).

The Calib program was used a first time on the *.RAW files. A "raw"

calibration type was done.  The pressure offset was 0.

The Calib program was then used a second time on the *.CAL files.

The pressure offset was again 0.  Only the conductivity ratio has been

modified by the following values :

Casts  1 to  3 : 0.999630

Cast   4       : 0.99993

Casts  5 to 18 : 0.999630

Cast  19       : 0.999507

Casts 20 to 22 : 0.999630

Cast  23       : 0.999542

Casts 24 to 26 : 0.999630

Cast  27       : 0.999585

Casts 28 to 34 : 0.999630

Cast  35       : 0.999672

Cast  36       : 0.999588

Cast  37       : 0.999709

Casts 38 to 98 : 0.999630

Cast  99       : 0.999807

3. Despiking (Program Despike)

The following parameters were used for all casts.

C     Fit        Over-      Min        Max       Min        Max       Spike

H     Width      lap        Value      Value     Stddev     Stddev    Tol

-     -----      -----      -----      -----     ------     ------    -----

P       25         5         0.00      3500.0    0.1000     30.000     2.50

T       25         5         0.00        20.0    0.0010      3.000     2.50

C       25         5         0.00         1.0    0.0005      0.005     2.50

4. Time compensation

The following input parameters were used for all probes :

Temp. Probe Dist above Cond. Cell Month        0.07 m

Sample Period                                  0.02 s

5. Swell removal (Program Delete)

The following delete parameters were used.

Pressure NOT filtered

Swells deleted

Low drop rates NOT deleted

6. Editing

Editing was done by Marie Robert using program CTDEDIT.  Howard Freeland

previously marked the plots for editing.  The plots were quite good.

Some of them had very large spikes easy to remove.  Surface data were

sometimes really bad.  A "smoothing operator" had to be applied for some

casts, always between 120 m and 300 m.

7. Bottle calibrations

Hydro comparisons were done by Ron Perkin and a Bi-linear correction

was applied to salinity for all casts except for cast 37.

IF pressure < 1000.0 dbars THEN

new_sal = old_sal + (0.005 - 0.5E-05*pressure)

ELSE IF pressure >= 1000.0 THEN

no correction

ENDIF

8. Particulars.

Casts 1-3 - deleted, test casts in Saanich Inlet.

Cast    5 - missing.

Cast    6 - deleted, aborted.

Cast    7 - deleted, aborted.

Cast   16 - deleted, upcast.

Cast   17 - deleted, upcast.

Cast   18 - deleted, aborted.

Cast   30 - missing.

Cast   99 - records from 0 to 22 dbars deleted since salinity was bad.

Marie Robert

# Moving files to new archive. May 20, 2003 – Joe Linguanti

- The CHE files in \\DATA_LIB\CTD contained nutrient data in umol/kg which is not the units we want in the archive, therefore the nutrient data in umol/L and chlorophyll data from Frank Whitney’s 9512-HYD.XLS spreadsheet was merged with the bottle data. Frank’s original file is in the DOC directory.

- Program HEADER EDIT was used to change non-standard channel names and units and add additional metadata to the CTD and Rosette files.

- Program Clean was used to;

- reset number of records, minimum and maximum values

- delete empty channels where applicable

- set channel name initial letters to uppercase

- Files .CLN were copied to new archive and renamed to the new file convention “YYYY-CC-NNNN.* “.

- The loop file was obtained from Frank Whitney. The spreadsheet file was then converted to IOS Header format with meta-data added to it. The original .XLS file is in the DOC directory.

- The SAIL files were copied from the TRACKOBS directory in the OSAP Data Archive.

- Channels not needed were removed.

- Program HEADER EDIT was used to change non-standard channel names and units and add additional metadata to the files.

- IOS Header Date and Time channels were added.

- Program Clean was used to set the start and end times and the geographical area.

October 19, 2004 – J. Linguanti

File 1993-03-0035.che – temperature below 150 m were incorrect (displaced by one sampling interval). Frank Whitney’s revised spreadsheet 9903-hyd.xls provided the corrected data as confirmed by CTD measurements.