Centre for Ocean Climate Chemistry

Cruise 9303

Freon Data Report

Wendy Richardson

Centre for Ocean Climate Chemistry

Institute of Ocean Sciences

Sidney, B.C., Canada

August 1993

# µ

Introduction	3

Methods	3

Sampling	3

Water	3

Air	4

Analysis	4

Calibration	4

Standards	4

Daily Calibration	4

Results	5

Quality of station data	5

Station Data	5

Station P01 - March 7, 1993	5

6

Station P04 - March 7, 1993	7

Station P12 - March 10, 1993	9

10

Station P16 - March 11, 1993	11

Station P20 - March 12, 1993	13

Station P26 - March 13, 1993	15

Station P26 -March 15, 1993	17

Tests	19

Bottle Blanks	19

Precision Test	20

Precision by Pooled Standard deviation of Duplicates	21

Restrips of water samples	22

Air	22

Freon Levels of Outside Air	22

Freon Levels of Lab Air	22

Discussion	22

Comparison of earlier data	22

Comparison of F11 at Station Papa over time	23

Comparison of F12 at Station Papa over time	24

Conclusions	25

Recommendations	25

References	25

# Introduction

The March 1993 COCC cruise to Station Papa was the third test of the Freon Analysis System in the field.  The level of contamination during the previous cruise on a Russian research vessel was a concern, and it was necessary to ensure the recovery of the freon analysis system.  This second trip to Station Papa was approximately one year since the first Preliminary testing of the Freon system.

# Methods

## Sampling

### Water

Freon samples were taken from 10 liter Niskin bottles only since the large bottles have a better O-ring surface area/water volume ratio than smaller Niskins.  These bottles were fairly well aged, but no specific precautions had been taken to lessen contamination of Freon from O-rings or elastic tubing in these Niskin bottles.  They were however, stored closed with the remnants of the previous sample out on deck where the air is not so heavily contaminated with CFC's.  The Freon samples were the first samples drawn from the Niskin since they are the most easily contaminated by air entering the head space of the bottle.  The samples were drawn directly into 100 cm precision ground glass syringes.

Once the Niskin bottle was on deck, the integrity of the bottle was verified by pushing in the drain valve while the vent remained closed.  If there was no leakage, the syringe valve was inserted into the drain valve and the vent opened.  The syringe was filled with about 10 cc. of seawater  The hydrostatic pressure of the water in the Niskin bottle forced the syringe plunger outward thus reducing the chance of drawing an air bubble into the syringe which is likely if the plunger was pulled out.  The syringe was then removed while the water was still flowing and emptied into the outflow (but not into the Niskin).  This 10 cc rinse was repeated three times and then the syringe was filled with approximately 100cc, inverted and plunger spun to send any bubbles to the top. while pushing out bubbles and sample.  The three 10 cc rinse technique was repeated and the syringe then filled to about 100cc; stopcocks closed.  An elastic was put on the syringe to maintain a positive pressure on the water sample.  The syringe was placed in a sample bucket filled with fresh seawater at least one inch above the syringe barrel.  If an air bubble was drawn into the syringe while sampling, the rinses were started all over again.  The bucket of samples was kept in a sink in the lab, with a continual overflow of seawater from the seawater loop manifold.

Hydrocast Freon profiles were carried out at stations P01, P04, P12, P16, P20, and P26.  Precision of analysis was checked with four Niskin bottles sampled at a depth of 30 meters at station P26.  Three subsamples from each Niskin were drawn and analyzed.  Bottle blanks were evaluated with 1200 meter water at 49.15.4 degrees North: 135.40 degrees West.

Carrier blanks, stripper blanks, and restripped samples were also analyzed throughout the cruise.

### Air

Syringe air samples were taken from the Helicopter deck, sampling deck and inside the lab.  The designated air syringe was flushed at least ten times, then a sample drawn, the stopcock closed, and tension elastic put in place.

## Analysis

Analysis are carried out using the Chloroflorocarbon Analytical System built by Glenn Smith, Centre for Ocean Climate Chemistry, Institute of Ocean Sciences.  The COCC analytical system is a modified version of the one described by Bullister and Weiss (1988).

## Calibration

### Standards

Working standard tank number 63100 was used throughout this cruise.  (F11, 428.55 ppt, standard deviation 3.72, and F12, 297.44 ppt, standard deviation 0.68)  This standard was made up of outside air and medical air by Glenn Smith using the COCC facilities.  The tank was calibrated against COCC's  lab standard tank number 63088 (F11,  445.41 ppt, standard deviation 6.95; and F12,  257.3 ppt, standard deviation 0.88).  This lab standard in turn had been calibrated against Ocean Chemistry's Freon standard tank number 70016 (F11,  271.0 ppt, standard deviation 0.42; and F12, 511.8 ppt, corrected standard deviation 0.69).  The Ocean Chemistry Freon standard was obtained and certified by Weiss's lab.

### Daily Calibration

Calibration curves used to determine sample concentrations were generated by multiple injections of predetermined volumes of standard gas.  Verifying the carrier blank would be the first procedure of the day.  This was followed by a suite of S(small loop)*1, S*2, S*3, S*4, and L(large loop)*1, L*2, L*3.  If these areas looked reasonable based on the previous days results, a stripper blank was analyzed and then water sample analyses begun.  A L*1 standard was run every four to six samples, approximately hourly, throughout the day to monitor drift in detector sensitivity.

# Results

Over three hundred and forty Freon analysis runs were carried out during this cruise.  A total of a hundred and forty water samples were analyzed.

Data reduction was carried out using the Ocean Chemisty/NOGAP PC adapted computer program based on the Scripps computer program (Weiss).  This program requires salinity and temperature for calculations; the former was taken from CTD cast data; and the latter was read from the sample bucket when the syringe was removed and attached to the extraction system.  The Instrument detection limit while set at threshold 2, was found to be F11 x<0.038 (Pmol/Kg) and F12 x<0.020 (Pmol/Kg).

There are two weak points  in the calculation procedure.  The first is that the caliper readings may not be as consistent as desired, due to the variation of the curve of the glass syringe.  Secondly, a slight variation of second aliquot results could be due to the temperature change as the sample warms up in the air, since the temperature used for calculation is that of the sample storage bucket.

## Quality of station data

This is an informal data report only.  Because of the nature of the Freon data reduction program, confirming recalculation of the results would be advisable before considering results for publication.  The results presented in this report are in raw form with no corrections.  Until a standard protocol has been decided upon, no corrections shall be applied to data report results.

## Station Data

### Station P01 - March 7, 1993

Calibration Curve, Percent Standard Error: F11 = 0.76; F12 = 2.31

Definitions:  s/c: suspected contamination

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 10.6 | 4.396 | 2.162 |
| 51.3 | 4.983 | 3.541 |
| 75.2 | 4.104 | 2.029 |
| 117 | 3.786 | 1.883 |

### µ §

### Station P04 - March 7, 1993

Calibration Curve, Percent Standard Error: F11 = 0.14; F12 = 1.35

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 11.3 | 4.425 | 2.185 |
| 49.3 | 4.091 | 1.977 |
| 98.7 | 3.378 | 1.749 |
| 148.2 | 2.704 | 1.351 |
| 298.0 | 1.702 | 0.844 |
| 396.0 | 1.163 | 0.481 |
| 595.9 | 0.209 | 0.148 |
| 796.5 | 0.071 | 0.105 |
| 1186.5 | 0.131 | 0.114 |

µ §

### Station P12 - March 10, 1993

Calibration Curve, Percent Standard Error: F11 = 0.75; F12 = 0.25

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 10.3 | 5.225 | 2.307 |
| 49.4 | 5.195 | 2.239 |
| 99.7 | 4.547 | 2.072 |
| 147.6 | 3.389 | 1.581 |
| 198.1 | 2.593 | 1.243 |
| 297.4 | 1.582 | .787 |
| 394.1 | .832 | .484 |
| 591.3 | .391 | .156 |
| 791.6 | .185 | .150 |
| 1191.6 | .081 | .030 |
| 1379.9 | .091 | <.020 |
| 1571.9 | .094 | .041 |
| 1776.2 | .097 | <.020 |
| 1970.8 | .084 | <.020 |

### µ §

### Station P16 - March 11, 1993

Calibration Curve, Percent Standard Error: F11 = 0.56; F12 = 0.26

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 8.7 | 4.929 | 2.202 |
| 47.6 | 4.931 | 2.226 |
| 98.5 | 4.819 | 2.230 |
| 147.5 | 4.148 | 1.987 |
| 147.1 | 4.063 | 2.000 |
| 195.9 | 3.553 | 1.771 |
| 295.4 | 2.344 | 1.160 |
| 391.7 | 1.314 | .638 |
| 591.7 | .333 | .181 |
| 790.9 | .125 | .051 |
| 790.1 | s/c | .075 |
| 986.9 | .087 | .124 |
| 1184.7 | .047 | .046 |
| 1377.4 | .038 | <.020 |
| 1579.4 | s/c | <.020 |
| 1777.5 | .069 | <.020 |
| 1973.1 | .076 | <.020 |

µ §

### Station P20 - March 12, 1993

Calibration Curve, Percent Standard Error: F11 = 0.68; F12 = 0.84

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 9.2 | 5.243 | 2.356 |
| 49.9 | 5.227 | 2.351 |
| 98.6 | 5.116 | 2.294 |
| 148.0 | 3.567 | 1.677 |
| 196.1 | 3.051 | 1.430 |
| 296.6 | 1.581 | .749 |
| 398.1 | 1.005 | .468 |
| 398.1 | 1.014 | .492 |
| 594.2 | .419 | .150 |
| 790.4 | .153 | .078 |
| 989.8 | s/c | <.020 |
| 1185.6 | .059 | <.020 |
| 1185.6 | .060 | .028 |
| 1383.6 | .047 | <.020 |
| 1581.5 | s/c | <.020 |
| 1777.6 | .070 | <.020 |
| 1974.4 | .081 | <.020 |
| 2955.7 | s/c | <.020 |

µ §

### Station P26 - March 13, 1993

Calibration Curve, Percent Standard Error: F11 = 2.19; F12 = 0.90

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 10.6 | 4.364 | 3.027 |
| 50.4 | 5.427 | 2.457 |
| 99.6 | 5.172 | 2.400 |
| 147.6 | 3.832 | 1.864 |
| 196.1 | 3.107 | 1.502 |
| 296.6 | 2.007 | .958 |
| 296.6 | 1.964 | .921 |
| 398.0 | 1.098 | .493 |
| 595.5 | .320 | .194 |
| 792.0 | .141 | .091 |
| 987.9 | .231 | .081 |
| 1185.7 | .062 | <.020 |
| 1383.6 | .095 | <.020 |
| 1580.3 | .045 | <.020 |
| 1777.2 | .054 | <.020 |
| 2169.4 | .053 | <.020 |
| 2365.4 | .043 | <.020 |
| 2561.7 | .069 | <.020 |
| 2955.3 | .184 | <.020 |
| 4200 | .064 | .050 |

µ §

### Station P26 -March 15, 1993

Calibration Curve, Percent Standard Error: F11 = 1.28; F12 = 0.63

| DEPTH (m) | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- |
| 48.2 | 5.557 | 2.550 |
| 100.7 | 5.214 | 2.424 |
| 146.3 | 3.786 | 1.920 |
| 198.8 | 2.936 | 1.452 |
| 295.1 | 2.065 | 1.054 |
| 295.1 | 2.029 | 1.049 |
| 295.1 | 2.039 | 1.029 |
| 593.2 | .325 | .272 |
| 791.6 | .174 | .201 |
| 986.1 | .050 | .161 |
| 1381.2 | .045 | .182 |
| 1581.9 | .072 | .108 |

µ §

## Tests

### Bottle Blanks

Sampling Location:  49.15.4N   135.40W

March 16, 1993  Depth 1200 M

| NISKIN ID | NISKIN Serial | SYRINGE | SALINITY | RUN # | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- | --- | --- | --- | --- |
| 2 | 10503 | 9823 | 34.419 | 1295 | .116 | .041 |
| 3 | 10210 | 16 | 34.419 | 1301 | .075 | <.020 |
| 4 | 10501 | 7625 | 34.419 | 1302 | .986 | .035 |
| 5 | 10186 | 24 | 34.420 | 1304 | <.038 | <.020 |
| 6 | 10496 | 5547 | 34.420 | 1305 | .055 | .040 |
| 7 | 8575 | 9828 | 34.420 | 1307 | .066 | .044 |
| 8 | 7684 | 3 | 34.421 | 1309 | .212 | .038 |
| 9 | 7685 | 5 | 34.420 | 1310 | .078 | .020 |
| 10 |  | 9838 | 34.421 | 1311 | .068 | .028 |
| 11 | 7691 | 100 | 34.421 | 1312 | .060 | .020 |
| 12 | 10494 | 5584 | 34.421 | 1313 | .259 | .079 |
| 13 | 10185 | 9839 | 34.421 | 1314 | .055 | .064 |
| 14 | 7692 | 7631 | 34.421 | 1316 | .107 | .070 |
| 15 | 10199 | 15 | 34.421 | 1317 | .047 | <.020 |
| 16 | 10505 | 5597 | 34.421 | 1318 | .053 | .033 |
| 17 | 10212 | 13 | 34.422 | 1319 | .114 | .050 |
| 18 | 8586 | 7633 | 34.422 | 1325 | .078 | .067 |
| 19 | 10216 | 9829 | 34.422 | 1326 | .054 | .053 |
| 20 | 10504 | 7677 | 34.422 | 1327 | .114 | .098 |
| 21 | 10491 | 7626 | 34.423 | 1328 | .084 | .095 |
| 22 | 10499 | 7746 | 34.422 | 1329 | .081 | .115 |
| 23 | 7693 | 9942 | 34.424 | 1330 | .056 | .072 |

Bottle blanks for Niskins 4, 8 and 12 were high, resulting in contamination of low level samples.  Therefore such samples were omitted from station tables and graphs.  Niskin 1 was broken before the bottle blank cast, but was obviously contaminated and those results have not been reported.  It would be adviseable to do bottle blanks at the beginning of a cruise, so that contaminated Niskins could be replaced,  preventing the loss of data.

### Precision Test

Calibration Curve, Percent Standard Error: F11 = 1.667; F12 = 0.387

March 14, 1993,  Station P26,  30 Meter Depth,  Rossette Cast

| SAMPLE | NISKIN | SYRINGE | SALINITY | RUN # | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- | --- | --- | --- | --- |
| 112a | 14 | 5584 | 32.563 | 1243 | 5.885 | 2.678 |
| 112b | 14 | 9839 | 32.563 | 1244 | 5.645 | 2.431 |
| 112c | 14 | 7631 | 32.563 | 1245 | 5.709 | 2.448 |
| 113a | 13 | 9829 | 32.563 | 1246 | 5.658 | 2.492 |
| 113b | 13 | 100 | 32.563 | 1247 | 5.482 | 2.439 |
| 113c | 13 | 13 | 32.563 | 1248 | 5.465 | 2.372 |
| 114a | 12 | 5597 | 32.562 | 1251 | 5.564 | 2..336 |
| 114b | 12 | 7633 | 32.562 | 1252 | 5.479 | 2.435 |
| 114c | 12 | 9942 | 32.562 | 1253 | 5.657 | 2.448 |
| 115a | 11 | 7677 | 32.562 | 1254 | 5.468 | 2.406 |
| 115b | 11 | 7626 | 32.652 | 1255 | 5.494 | 2.442 |
| 115c | 11 | 7746 | 32.652 | 1256 | 5.406 | 2.471 |
|  |  |  |  | Mean | 5.576 | 2.450 |
|  |  |  |  | STD | 0.132 | 0.080 |
|  |  |  |  | %STD | 2.368 | 3.248 |

### Precision by Pooled Standard deviation of Duplicates

Duplicates on this cruise were taken by one syringe for each of two Niskin Bottles tripped at the same depth.

| F11 X Pmol/Kg | F11 Y Pmol/Kg | F11 X-Y Pmol/Kg | F12 X Pmol/Kg | F12 Y Pmol/Kg | F12 X-Y Pmol/Kg |
| --- | --- | --- | --- | --- | --- |
| 1.005 | 1.014 | .009 | .468 | .492 | .024 |
| .059 | .060 | .001 | 0 | .028 | .028 |
| 2.007 | 1.964 | .043 | .958 | .921 | .037 |
| 2.065 | 2.029 | .036 | 1.054 | 1.049 | .005 |

µ §

Using the above equation to calculate the precision by the pooled standard deviation of duplicates the results were as follows.

F11  For the range  0.059 - 2.65 Pmol/Kg, the pooled standard deviation for duplicates was 0.056 Pmol/kg.

F12  For the range 0 - 1.054 Pmol/Kg; the pooled standard deviation for duplicates was 0.052 Pmol/Kg.

Comparison of Cruises

|  | Freon | 11 | Freon | 12 |
| --- | --- | --- | --- | --- |
| CRUISE | Range | Pooled STD | Range | Pooled STD |
| COCC9201 Stn PAPA | .070 - 4.571 | .075 | .074 - 2.090 | .072 |
| COCC9203 Russian N.Pac | .617 - 2.778 | .599 | .346 - 1.470 | .394 |
| COCC9303 Stn PAPA | .059 - 2.65 | .059 | 0 - 1.054 | .052 |

### Restrips of water samples

| STN | SAMPLE RUN # | RESTRIP RUN # | SAMPLE F11 Pmol/Kg | RESTRIP F11 Pmol/Kg | CARRYOVER F11 % | SAMPLE F12 Pmol/Kg | RESTRIP F12 Pmol/Kg | CARRYOVER F12 % |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| P12 | 1098 | 1099 | .181 | 0 | 0 | 0 | 0 | 0 |
| P16 | 1128 | 1130 | .306 | 0 | 0 | .164 | 0 | 0 |
| P20 | 1185 | 1188 | 5.243 | .009 | 0.17 | 2.356 | 0 | 0 |
| P26 | 1203 | 1204 | .165 | 0 | 0 | .071 | 0 | 0 |
| BB | 1302 | 1303 | .986 | .058 | 5.8 | .035 | 0 | 0 |
| BB | 1319 | 1324 | .114 | 0 | 0 | .050 | 0 | 0 |

The above table demonstrates the high stripper efficiency of the Freon analysis system on this cruise.

## Air

### Freon Levels of Outside Air

| DATE | STN | RUN TIME | RUN # | SAMPLE | F11 ppt | F12 ppt |
| --- | --- | --- | --- | --- | --- | --- |
| March 12 1993 | P20 | 334 | 1166 | SAMPLING DECK AIR | 261.66 | 499.78 |
| March 15 1993 | P26 | 102 | 1262 | HELI DECK AIR | 265.54 | 531.88 |
| March 15 1993 | P26 | 1029 | 1273 | SAMPLING DECK AIR | 268.37 | 614.53 |
| March 17 1993 | BB | 223 | 1322 | HELI DECK AIR | 265.43 | 622.51 |

### Freon Levels of Lab Air

| DATE | STN | RUN TIME | RUN # | SAMPLE | F11 ppt | F12 ppt |
| --- | --- | --- | --- | --- | --- | --- |
| March 7 1993 | P04 | 2107 | 1057 | LAB AIR | 304.40 | 1266.4 |
| March 15 1993 | P26 | 1017 | 1272 | LAB AIR | 386.89 | 2668.7 |
| March 17 1993 | BB | 0234 | 1323 | LAB AIR | 387.42 | 4384.3 |

# Discussion

## Comparison of earlier data

Comparison plots of F11 and F12 results from Station Papa over several years are shown on the following pages.  These data are from Gammon (1992), Schwarz (1988) and the data presented in the COCC 9201 Freon Data Cruise report (Richardson 1993), and this COCC 9303 Freon Data Cruise report.

### Comparison of F11 at Station Papa over time

### Comparison of F12 at Station Papa over time

# Conclusions

# Recommendations

As stated previously no particular precautions, other than storing the Niskin bottles on deck, were employed to reduce the Freon levels of the Niskin bottles.  If it is considered important to lower our deep water results and blanks, an attempt to treat the Niskins as described by Bullister and Weiss (1987) should be made.  This entails replacing the elastic tubing with epoxy coated stainless steel springs, baking the O-rings for several days and then storing them in a gas tight container until installed on the bottles at sea.

Considering the space restrictions of the John P. Tully, which result in many people and analyses being carried out in one lab, it is mandatory that all people respect the signs requesting that no spray cans be used in the lab.  Even without Freon propellants, di-chloro and tri-chloro ethane cause major contamination problems resulting in the loss of samples and triple the analysis time required for each sample.  This in turn results in a much longer storage time for samples overall, reducing their integrity.  These problems were also encountered and described by U.Schwarz (1989)

Bottle blanks should be checked before sampling stations begin, in order to prevent any contamination of samples by the Niskins.

# References

Bullister J.L., Weiss R.F. (1987)  Determination of CCl3F and CCl2F2 in Seawater and Air. Deep-Sea Research Vol. 35 No. 5, 839-853.

Gammon R.H., Cline J.D., Wisegarver (1982)  Chlorofluoromethanes in the Northeast Pacific Ocean: Measured Vertical Distributions and Application as Transient Tracers of Upper Ocean Mixing. J. Geophys. Res. 87(C12), 9441-9454.

Richardson W. (1993)  Centre for Ocean Climate Chemistry Cruise 9201 Freon Data Report, Unpublished manuscript, Institute of Ocean Sciences. Sidney, B.C. Canada.

Schwarz U. (1988)  Freon data for OC-85-OS-02 and OC-85-OS-03, Unpublished manuscript, Institute of Ocean Sciences. Sidney, B.C. Canada.

Schwarz U. (1989)  Freon Manual, Unpublished manuscript, Institute of Ocean Sciences, Sidney, B.C. Canada.

Weiss R.F.  Freon Lab Manual, Unpublished manuscript, Scripps Institute of Oceanography, San Diego, California, U.S.A.