Physical and Chemical Data Collected in the Beaufort, Chukchi and East Siberian Seas,

August - September 1993

R.W. Macdonald, M. O'Brien, E.C. Carmack,  R. Pearson, F.A. McLaughlin, D. Sieberg, J. Barwell-Clarke , D.W. Paton and D. Tuele

Institute of Ocean Sciences

Department of Fisheries and Oceans

Sidney, B.C.  V8L 4B2

1995

Canadian Data Report of Hydrography and Ocean Sciences 139

Canadian Data Report of Hydrography and Ocean Sciences No. 139

1995

Physical and Chemical Data Collected in the Beaufort, Chukchi and East Siberian Seas,

August - September 1993

by

R.W. Macdonald, M. O'Brien, E.C. Carmack, R. Pearson, F.A. McLaughlin,

D. Sieberg, J. Barwell-Clarke, D.W. Paton and D. Tuele

Institute of Ocean Sciences

Department of Fisheries and Oceans

Sidney, B.C.  V8L 4B2

Copyright Minister of Supply and Services Canada - 1995

Cat. No. Fs  97-16/139            ISSN  0711-6721

Correct citation for this publication :

Macdonald, R.W., M. O'Brien, E.C. Carmack, R. Pearson, F.A. McLaughlin, D. Sieberg, J. Barwell-Clarke, D.W. Paton and D. Tuele 1995, Physical and Chemical Data Collected in the Beaufort, Chukchi and East Siberian Seas, August - September 1993. Can. Data Rep. Hydrogr. Ocean. Sci.: 139, 288pp.

Table of Contents

1. INTRODUCTION

1.1 Overview of the Field Work

1.2 Station Locations and Information

2. METHODS

2.1 Field Sampling - Rosette Casts

2.2 CTD (Conductivity-Temperature-Depth) Systems

2.2.1 Data Collection Methods

2.2.2 Instrument Problems

2.2.2.1 FSI Current Setting Too Low

2.2.2.2 Guildline RF Interference

2.2.2.3 Pressure/Time Drift with Guildline Temperature & Conductivity Data

2.2.2.4 Air Bubbles Trapped in Guildline Conductivity Cell

2.2.2.5 Sea-Bird Electronic Problems

2.2.3 Data Processing

2.2.3.1 FSI CTD Data

2.2.3.2 Guildline CTD Data

2.2.3.3 Sea-Bird CTD Data

2.2.4 CTD Data Validation

2.2.4.1 Summary of Accepted Uncertainties

2.2.4.2 FSI CTD Data

2.2.4.3 Transmissometer, Fluorometer, and PAR Data:

2.2.4.4 Guildline CTD S/N CTD5 Data

2.2.4.5 Guildline CTD S/N 53501 Data

2.2.4.6 Sea-Bird CTD Data

2.3 Chemistry Sampling and Analyses

2.3.1 Laboratory Methods

2.3.1.1 Temperature

2.3.1.2 Salinity

2.3.1.3 Dissolved Oxygen

2.3.1.4 Nutrients

2.3.1.5 Chlorophyll a and phaeo-pigments

2.3.1.6 Total Suspended Solids

2.3.1.7 Particulate Organic Carbon and Nitrogen

3. LIST OF FIGURES

4. LIST OF TABLES

5. REFERENCES

6. APPENDICES

6.1 CTD Data: Tables and Plots

6.2 Chemistry Data Tables

6.3 Chemistry Profile Plots

6.4 Chemistry Property / Property Plots

6.5 Chemistry Tables - Filtration Samples

6.6 Profile Plots - Filtration Samples

7. Related Data Reports in the Canadian Data Report of Hydrography and Ocean Sciences Series

Abstract

Macdonald, R.W., M. O'Brien, E.C. Carmack, R. Pearson, F.A. McLaughlin,

D. Sieberg, J. Barwell-Clarke, D.W. Paton and D. Tuele 1995. Physical and Chemical Data Collected in the Beaufort, Chukchi and East Siberian Seas, August - September 1993. Can. Data Rep. Hydrogr. Ocean. Sci.: 139, 288pp.

During a cruise aboard the CCGS Henry Larsen water properties were measured in the Beaufort Sea, Canada Basin, Chukchi Sea and East Siberian Sea as part of a study  investigating the transport and fate of contaminants in the western Arctic Ocean.  The oceanographic support measurements reported here include conductivity-temperature-pressure  (CTD) profiles and chemical measurements made on water collected by a rosette sampler (salinity, dissolved oxygen, orthophosphate, reactive silicate, nitrate and nitrite, chlorophyll a, particulate organic nitrogen and carbon and total suspended solids).

Key words: Arctic, Beaufort Sea, Canada Basin, carbon, chlorophyll a, Chukchi Sea, CTD, East Siberian Sea, nitrogen, nutrients, oceanography, oxygen, salinity, temperature, total suspended solids.

Résumé

Macdonald, R.W., M. O'Brien, E.C. Carmack, R. Pearson, F.A. McLaughlin,

D. Sieberg, J. Barwell-Clarke, D.W. Paton and D. Tuele 1995. Physical and Chemical Data Collected in the Beaufort, Chukchi and East Siberian Seas, August - September 1993. Can. Data Rep. Hydrogr. Ocean. Sci.: 139, 288pp.

Au cours d’une mission à bord du CCGS Henry Larsen certaines propriétés hydrographiques des eaux de la mer de Beaufort, du bassin Canada, de la mer Chukchi et de la mer de Sibérie est ont été mesurées dans le cadre de travaux de recherche sur le transport et le devenir des contaminants dans l’océan Arctiques ouest. Les données océanographiques de support présentées dans ce rapport comprennent des profils conductivité-température-pression (CTD) et des mesures chimiques sur des échantillons d’eau obtenus avec une rosette (salinité, oxygène dissous, orthophosphate, silicate réactif, nitrate et nitrite, chlorophylle a, azote organique et carbone particulaire et solides totaux en suspension).

Mots-clés : Arctique, mer de Beaufort, bassin Canada, carbone, chlorophylle a, mer Chukchi, CTD, mer de Sibérie est, azote, nutriments, océanographie, oxygène, salinité, température, solides totaux en suspension.

Acknowledgments

We are very much indebted to the officers and men of the Canadian Coast Guard Ship Henry Larsen for their support in carrying out the sampling.  We especially note the willingness of Capt. S.A. Gomes to take advantage of the good ice conditions to extend the work as far east as the Makarov Basin, and the continuing efforts of Captain David Johns and Ivan Côté (Coast Guard Northern, Ottawa) to plan and organize icebreaker support for our program.  This work was jointly supported by the Arctic Environmental Strategy Green Plan Program and the PERD Climate Program. We thank S. Thomson for advice on style and text-editing.

# INTRODUCTION

The objectives of this cruise were (1) to recover, service and redeploy four moorings in the Beaufort Sea at 700m and 3300m and (2) to carry out an oceanographic section in the Canadian Basin comprising the Beaufort, Chukchi and East Siberian seas, measuring water properties and collecting sediment cores.  In part, this cruise was a preparation for a major oceanographic section to be carried out in 1994 (AOS-94).  Sampling included water, sediments, suspended particulates and zooplankton.  The data assembled in the present report includes only the standard supporting oceanographic determinations:  conductivity-temperature-pressure profiles; % transmissivity, fluorescence and  photosynthetically active radiation (PAR) profiles;  bottle measurements for salinity, temperature, nutrients, dissolved oxygen, Chl a, particulate organic carbon and nitrogen, and total suspended solids.

Additional measurements (not reported here) have been or will be made for water-column tracers (freons and carbon tetrachloride, d18O, helium/tritium, barium) and for contaminants (hexachlorocyclohexanes (HCHs), 137Cs, 129I, 239,240Pu, 90Sr).   Sediments were collected at several stations with a large box corer and were sectioned aboard ship and will be analyzed for contaminants (organochlorines, PAH, metals and radionuclides). Samples of zooplankton were collected for organochlorine determination and large volumes of water were filtered to collect suspended particles for organochlorine determination.

This project was carried out aboard the ice breaker CCGS Henry Larsen (IOS mission #9324).  The field work was divided into two legs: the first, from 15 to 25 August involved mooring activities in the Beaufort Sea and Canada Basin; the second, from 25 August to 25 September, focused on the oceanography of the Chukchi and East Siberian seas. This expedition was a cooperative mission involving scientists from Canada (Institute of Ocean Sciences and Bedford Institute of Oceanography), Russia (Institute of Nature Conservation) and the United States (University of Washington and Scripps Institution of Oceanography).

Within the context of global change and contaminants, the Larsen-93 expedition had as key scientific objectives :

- To continue time-series measurements at the Arctic Ocean Climate Station sites in the Canada Basin by recovering and redeploying moorings and obtaining ancillary oceanographic measurements.

- To investigate the regional oceanography of the Chukchi and East Siberian shelves.

- To measure contaminant burdens in the water column and in sediments and zooplankton.

## Overview of the Field Work

The sampling program for mission #9324 accomplished the following (data summarized in this report are outlined in bold font):

- Collect conductivity-temperature-depth (CTD) data.

- Collect fluorometer, transmissometer and PAR measurements.

- Collect water samples for the determination of nitrate plus nitrite, silicate, orthophosphate, salinity, temperature, dissolved oxygen, chlorophyll a, particulate organic carbon and nitrogen, total suspended solids, CFCs, total CO2, d13C, alkalinity, d18O, tritium/helium, barium, HCHs and radionuclides.

- Collect samples for organochlorine determination from particulate and dissolved phases of the water column (large volume filtration and in-situ pump sampling).

- Collect zooplankton samples for organochlorine analysis using vertical net hauls.

- Collect sediment box cores for a variety of analyses including contaminants (e.g, metals, PAH, organochlorines, radionuclides).

- Recover time series mooring at station A01-92 in the Canada Basin.

- Redeploy time series mooring for a period of 2 years at station A01-93.

- Recover mooring at station AM1-92 in the Mackenzie Canyon.

- Redeploy mooring at station AM10-93 for a period of 2 years at the edge of the Mackenzie Canyon.

- Deploy moorings in Herald Canyon.

- Collect data on current profiles (using Acoustic Doppler Current Profilers) over the Barrow Canyon and Chukchi Slope.

## Station Locations and Information

Figure 1(a-d) is a series of charts showing the station locations in the Beaufort, Chukchi and East Siberian seas during the 9324 mission.  The locations for Guildline or FSI CTD casts were determined using a Trimble Transpak GPS located near the CTD acquisition system with an antenna mounted on the starboard side of the ship two decks up from the boat deck.  The locations for Sea-Bird CTD casts were taken from the ship's navigation system on the bridge. The positions are expected to be within 100m of the true position.

Table 1 provides a chronological list of the CTD profile locations (start of cast), indicates the instrumentation used in the profiling, and whether or not water sampling was done for that cast. Where more than one cast was done at a station separate coordinates are given for each cast.

# METHODS

## Field Sampling - Rosette Casts

The water sampling and CTD casts were carried out from the starboard boat deck using a 24-bottle GO (General Oceanics) rosette sampler outfitted with 24 l0 L BIO-design bottles. In addition to the CTD instrumentation, the system also included a Sea Tech transmissometer, a Sea Tech Fluorometer, a dissolved oxygen probe, and a PAR sensor. The rosette package was lowered over the side using a crane and winch and the rosette was placed on a flat trolley once on deck and pulled by means of a come-along into a heated container for the water sampling. The instrumental data were collected on a PC located in the CTD/Rosette lab by means of a conducting cable.

## CTD (Conductivity-Temperature-Depth) Systems

Four CTDs were used during the cruise, two from Guildline, one from FSI and one from Sea-Bird. The FSI instrument was the main CTD and was attached to the rosette for the duration of the cruise.  For several of the rosette casts we attached one of the Guildline CTDs to the rosette (in place of a bottle) to obtain an intercomparison between the two CTDs on the rosette.  Guildline CTDs were also used on their own in cases where water sampling was not needed, and the Sea-Bird CTD was used for CTD-only casts in shallower water. Tables 2-8 list the CTDs, their serial numbers and configurations, external sensors used, intercomparisons made, and the uncertainties based on instrumental specifications, calibrations, and bottle intercomparison data.

The two Guildline CTDs and the FSI CTD transmitted their data in “real-time” via a 3500 m, three conductor sea-cable to their matching deck units.  For each of the three CTDs the data were transmitted at 25 samples per second.  The FSI data were logged on an ALR 486 desktop MSDOS computer and the Guildline data were logged on a Toshiba T5200 portable computer.   The GO 1016 rosette pylon also communicated real-time with its deck unit via the sea cable and was controlled by the T5200 portable computer.

The general configuration used for sharing the sea-cable between the various instruments was to run the FSI CTD through two of the conductors and to run the rosette underwater unit through the remaining conductor and the shield.  When an additional Guildline CTD was mounted on the rosette, it shared the rosette conductors.  Sharing a single conductor between the Guildline CTD and the rosette was accomplished by switching between the two manually; the Guildline CTD was powered on the downcast to collect data and then switched manually to the rosette to trip bottles on the up-cast.

The Sea-Bird CTD recorded internally and therefore needed no external conductors.  Sea-Bird casts were done by lowering the instrument on the end of a hydrowire.  The data were logged at  2 samples per second during the cast and then transferred to the ALR computer by serial cable after the cast was completed.

### Data Collection Methods

All rosette casts and casts involving the Guildline CTDs along were taken from the ship's starboard boat deck.  The rosette and CTDs were positioned over the side using the ship's starboard cargo crane and lowered using a hydraulic winch.

The Sea-Bird CTD casts were taken from the port side of the ship's foredeck.   The CTD was positioned over the side and lowered using a hydraulic winch and A-frame.

The temperatures on deck varied from about 5 °C to -10 °C.   To prevent freezing of the CTDs and rosette, we stored them in a specially designed "hanger" constructed from a pair of 8'x12' cargo containers.   The rosette and CTDs were kept in the container at room temperature until just before the cast when they were rolled out on deck and deployed.  The Sea-Bird CTD used from the foredeck was brought into the forecastle between casts.

### Instrument Problems

#### FSI Current Setting Too Low

The data from casts 4,5 and 6 taken with the FSI CTD were extremely spiky.  This apparently was caused by setting the current level too low for the FSI deck unit power supply.  After cast 6 we increased the current level and the FSI CTD performed virtually spike-free for the rest of the cruise.

#### Guildline RF Interference

Both Guildline CTDs suffered from sporadic bursts of noise spikes.  We were able to determine coincidence between these bursts and the ship’s radio transmissions.  A shielded deck-cable running from the winch to the CTD deck units grounded at the deck unit end did not appear to prevent this interference.  This has been a common problem with the Guildline CTDs, in our experience, which required data from every cast taken with the instrument to be despiked.

#### Pressure/Time Drift with Guildline Temperature & Conductivity Data

As discussed fully in sections 2.4.3 and 2.4.4, comparisons between the FSI and Guildline CTDs show significant drifts in Guildline temperature and conductivity readings over the duration of each cast.  There are also significant offsets in these channels between casts.

Evidence suggests that the drift is due to a difference between the CTD internal temperature and the water temperature. The most likely reason for this problem is that we kept the CTDs at room temperature until just before a cast and then soaked them at the surface only for 2 to 3 minutes before profiling.

#### Air Bubbles Trapped in Guildline Conductivity Cell

For many of the Guildline CTD casts we found that the conductivity readings did not change to surface water values when the CTD was first put in the water and remained close to zero for some time.  We attribute this problem to air bubbles trapped in the conductivity cell.  We were able to correct the problem by repeatedly lowering the CTD to some depth below 20 m and then bringing it back to the surface (thereby clearing the air bubble).  This problem persisted over the duration of the cruise despite efforts to keep the conductivity cells clean and wetted with alcohol at all times when the CTDs were out of the water.

In most cases this problem was identified at the start of the cast and eliminated before the actual cast was recorded.  In a few cases, however, the problem was not identified before the cast started or it was ignored because the Guildline CTD was on the rosette and we knew that we would recover good data from the FSI CTD.  Casts corrupted in this way have not been included in this report.

#### Sea-Bird Electronic Problems

On the last day of CTD casts (casts 103-115), the Sea-Bird CTD 1031 developed an electronics fault that caused it not to record large sections of data during the casts.  Because the Sea-Bird is an internally-recording CTD, we could not tell if this happened until after the cast was completed.  Generally, the lost records occurred during the first 50 to 80 m of the cast and then after that the instrument worked well.  As a result, we were able to reconstruct the corrupted profiles from the up-cast data.  However, this procedure introduces additional uncertainty in the temperature and salinity readings because sampling occurs in the wake of the CTD as it rises through the water producing turbulence and potentially contaminating temperature.

### Data Processing

The steps outlined below  were performed in the processing of each CTD cast.

#### FSI CTD Data

- The raw, binary, data collected from this CTD were converted to an ASCII format.  Since pre-cruise pressure, temperature, and conductivity calibration coefficients for this CTD were entered into the CTD itself  the values for these channels were in engineering units.  However, the values for the transmissometer, fluorometer, and PAR sensor were still raw at this point.

- The raw ASCII data were converted to engineering units for all channels and salinity was calculated.  For the transmissometer and PAR sensors we used pre-cruise calibration coefficients and for the fluorometer we used nominal coefficients based on the fluorometers selected sensitivity range.

- The casts were processed using an automatic despiking program to remove noise spikes.

- A 0.45 second delay exponential filter was applied to the conductivity data to match its time response with that of temperature.  To compensate time properly for some of the casts, it was also necessary to delay the temperature data by shifting it a number of records relative to the conductivity data.

- The pressure data were adjusted to match the post-cruise calibration using the following formula:

	Corrected Pressure = Pressure ´ 0.999185 + 2.0

- Based on comparisons with bottle data, the conductivity was corrected using the following formula:

	Corrected Conductivity = Conductivity ´ (Julian Day  ´ 0.5957 ´ 10-5  +0.99907)

This formula corrects the conductivity for a slow drift with time over the length of the cruise.

- All swells and up-casts were removed from the data.

- The data for each cast were decimated into one-decibar bins using a simple averaging process.

- The temperature and salinity data were filtered using a 1 decibar low-pass filter to remove spikes from the salinity that could not be adequately removed by time compensation.

- Derived oceanographic quantities were calculated from the pressure, temperature, and salinity data using the algorithms given by Fofonoff and Millard [1983].

#### Guildline CTD Data

- The raw, binary, data collected from this CTD were converted to an ASCII format.

- The raw ASCII data were converted to engineering units for all channels using pre-cruise calibration coefficients and salinity was calculated.

- The casts were processed using an automatic despiking program to remove any noise spikes.  For some casts, additional manual despiking was required.

- For two of the casts (68,69), the temperature data time response had to be adjusted to match that of the conductivity.  This was done by applying a 0.3 second exponential filter to the temperature data and then delaying the temperature by 2 records relative to the other channels.

- For the Guildline CTD5 only, the pressure data were adjusted to match the pressures measured by the FSI in casts involving the two CTDs using the formula:

	Corrected Pressure = Pressure ´ 1.006667  - 0.6667

- All swells and up-casts were removed from the data.

- The data for each cast were decimated into one-decibar bins using a simple averaging process.

- Based on comparisons with the FSI CTD, the conductivity and temperature were corrected using the coefficients A and B listed in Table 3 and applying the following formula:

	Corrected Value = Measure Value + A + B ´ Pressure

- The temperature and salinity data were filtered using a one-decibar low-pass filter to remove spikes from the salinity that could not be adequately removed by time compensation.

- Derived oceanographic quantities were calculated from the pressure, temperature, and salinity data using the algorithms given by Fofonoff and Millard [1983].

#### Sea-Bird CTD Data

1.	The raw, HEX-ASCII, data from this CTD were converted to an ASCII format.

2.	The raw ASCII data were converted to engineering units for all channels using pre-cruise calibration coefficients and salinity was calculated.

3.	Each cast was processed to compensate for the differences in the time responses of the temperature sensor and the conductivity cell.  This was done by delaying the temperature data by 3 records and the conductivity data by 1.5 records.

- The pressure, temperature, and conductivity data were filtered using a 3 Hz low-pass filter to smooth out sensor noise and match sensor response times.  Salinity was recalculated at this point.

- Casts 103 to 115 were reconstructed from the up-cast and available sections of the down-cast.

- All swells and up-casts were removed from the data.

- The data for each cast were decimated into one-decibar bins using a simple averaging process.

- Based on comparisons with the Guildline CTD S/N CTD5 data, the conductivity was corrected using the following formula:

Corrected Conductivity = Conductivity ´ 1.06249 - 0.0397

- The salinity data was filtered using a one-decibar low-pass filter to remove spikes from the salinity that could not be adequately removed by time compensation.

- Derived oceanographic quantities were calculated from the pressure, temperature, and salinity data using the algorithms given by Fofonoff and Millard [1983].

### CTD Data Validation

Data were validated in several ways including comparison of pre-cruise and post-cruise calibrations, comparison of CTD with bottle data, and intercomparisons between CTDs.

#### Summary of Accepted Uncertainties

From the following analysis, the accepted uncertainties in the various CTD sensor readings are as follows:

#### FSI CTD Data

The first step in validating the data was to assure that the FSI CTD data, as measured for all rosette casts, was accurate.  The FSI data were then used as the benchmark for the data from the other CTDs.

Pressure

The pressure sensor electronics in the FSI CTD includes an internal temperature sensor which is used in factory calibrations to determine the temperature effect on pressure readings.  Therefore, the factory calibrated pressure readings are supposed to be fully corrected for instrument temperature.

When we performed a post-cruise calibration in our labs, however, we found that the factory calibration agreed with ours at room temperature but not at freezing temperatures.  At 0 °C there was a two-decibar offset at zero decibars, decreasing to a zero offset at 3000 decibars.  Since our casts were primarily in water at 2 to -2 °C, we applied a correction to the FSI pressure data to adjust the calibrated readings to match the sensor's response (Corrected Pressure = Pressure ´ 0.999185 + 2.0)

Temperature

The pre-cruise temperature calibration was compared to the post-cruise calibration and found to agree to within 0.002 C° over the range -2 to 10 °C.   We also compared some of the deep potential temperature readings with historical data from the same areas finding agreement to within the accuracy of the instrument.  As a result, we considered the temperature values based on the pre-cruise calibration valid.

Conductivity/Salinity

The FSI conductivity data were validated by comparison with the conductivity calculated from the rosette bottle samples.  The bottle conductivities were calculated from the bottle salinities (measured using a Guildline Autosal laboratory salinometer) and the CTD temperature measured at the time the bottle was tripped.  Bottle and CTD conductivities were used from casts spanning the entire cruise.  However, only samples collected deeper than 300 m were used because relatively high gradients in the water column above that depth confound such an intercomparison.

Figure 2, the ratio of bottle conductivity/CTD conductivity vs. pressure, shows no apparent correlation between bottle conductivity/CTD conductivity and pressure.  However, Figure 3, a plot of the same bottle conductivity/CTD conductivity ratios vs. Julian Day, shows a general trend in the ratio between the two conductivities as a function of time. The best linear fit (Corrected Conductivity = Conductivity ´ (Julian Day  ´ 0.5957 ´ 10-5+ 0.99907)) was used to remove this drift from the FSI CTD data.  The CTD salinity was recomputed from the corrected conductivity and the differences between the CTD and bottle salinities were calculated (Table 5; Figure 4).

#### Transmissometer, Fluorometer, and PAR Data:

The transmissometer and PAR data are calculated using pre-cruise calibration coefficients.  Because no independent measurement exists for validating these measurements, the data are considered accurate to within the accuracy of the sensors.

The fluorometer was not calibrated prior to the cruise and therefore nominal coefficients have been used to produce chlorphyll concentrations as a function of fluorometer raw readings. The fluorometer will eventually be calibrated based on chlorophyll concentrations measured from the bottle samples.

#### Guildline CTD S/N CTD5 Data

The data from Guildline CTD S/N CTD5 was validated by comparing the 1 metre averaged CTD5 with the “benchmark” FSI CTD data for casts in which both the FSI CTD and CTD5 were lowered on the rosette together (see Table 6 for listing).

Pressure

Pressure readings from the two CTDs were compared by using the pressures associated with key features in the cast profiles including, for example, the maximum depth of the casts.  This comparison yielded an adjustment formula for the Guildline CTD S/N CTD5 (Corrected Pressure = Pressure  ´ 1.006667  - 0.6667) which was applied to all CTD5 data before further processing.

Temperature

The differences between the FSI temperature and CTD5 temperature were calculated and are shown in Figure 5 as a function of pressure.  Three things can be readily observed from this plot:

- The water column above 300 m is variable with depth and the high gradients mask the small temperature offsets sought in the intercomparison.

- There is a definite drift in the differences as a function of pressure  (or time during the cast).

- There is a varying offset in the temperature differences that does not appear to be a linear function of time. (The cast 75-76 difference is closer to the 52-53 difference, which is 4 days away, than the 77-78 difference which is 2 hours away.)

We believe that the differences between the FSI CTD and CTD5 are not simply a function of pressure but are due to variations in the CTD5 electronics as a function of the difference between internal CTD temperature and ambient temperature.  This hypothesis could be tested by performing a number of lab tests using dummy sensors and varying calibration bath temperatures.  In the absence of such testing, we have applied only a simple temperature correction as a function of pressure to the CTD5 data (Corrected Temperature = Temperature + Pressure ´ 0.35714 ´ 10-5).  Figure 6, a plot of the difference between the FSI CTD temperature and the CTD5 corrected temperature data vs. pressure, indicates that the corrected CTD5 temperatures agree with the FSI temperatures to within ±0.005 C° for the 3 intercomparison casts.

Conductivity/Salinity

As in the case of temperature, the differences between the FSI conductivity and CTD5 conductivity were calculated (Figure 7). The conductivity intercomparison shows trends and offsets similar to those shown in the temperature intercomparison (Figure 5) supporting the notion that the differences arise from a common source -- i.e., electronic drift as a function of internal temperature.

As with the temperature comparison, the best correction that could be applied to the CTD5 data without further extensive laboratory testing was a simple pressure correction (Corrected Conductivity = Conductivity + 0.9 ´ 10-4 + Pressure ´ 0.1 ´ 10-6) shown in Figure 8.

The corrected salinities for CTD5 were computed from the corrected temperatures and conductivities and the difference between the FSI salinities and CTD5 corrected salinities were calculated.  Figure 9, a plot of the salinity differences as a function of pressure, indicates that the CTD5 corrected salinities agree with the FSI salinities to within ±0.006 PSU.

#### Guildline CTD S/N 53501 Data

The data from Guildline CTD S/N 53501 were validated by comparing the 1 metre averaged 53501 and FSI CTD data from simultaneous casts.  Only two such casts were performed during the cruise in which both the FSI CTD and 53501 were lowered on the rosette together (Table 7).

Pressure

Pressure readings from the two CTDs, which were compared by matching the pressures associated with key features in the cast profiles, showed agreement to within 1 decibar at all depths.

Temperature

The differences between the FSI temperature and 53501 temperature were calculated.  Figure 10 shows an even more pronounced drift in temperature difference with pressure (or time during the cast) than was seen with CTD5, suggesting the problem to be common to both Guildline CTDs.  As with CTD5, only a simple pressure correction could be applied without further lab work (Corrected Temperature = Temperature - 0.0135 + Pressure ´ 0.1199 ´ 10-4).

Figure 11 indicates that the corrected 53501 temperatures agree with the FSI temperatures to within ±0.01 C° from 500 metres and below for the casts compared.

Conductivity/Salinity

As with the CTD5, the 53501 conductivity showed a drift in difference from the FSI conductivity as a function of pressure (or time during cast) which was corrected (Figure 13) using a simple formula (Corrected Conductivity = Conductivity - 0.08 ´ 10-4 + Pressure ´ 0.1 ´ 10-6).

The corrected salinities for 53501 were computed from the corrected temperatures and conductivities and the difference between the FSI salinities and 53501 corrected salinities were calculated (Figure 14).  The 53501 corrected salinities agree with the FSI salinities to within ±0.006 PSU from about 500 metres and deeper.

#### Sea-Bird CTD Data

Because there were no casts done using the FSI CTD and the Sea-Bird CTD simultaneously, the “corrected” Guildline CTD5 data were used as the benchmark for the Sea-Bird CTD readings.  The data from Sea-Bird CTD S/N 1031 were validated by comparing the 1 metre averaged 1031 and Guildline CTD5 data from a single simultaneous profile (casts #88,89). There are shortcomings associated with this procedure:

- The Guildline CTD5 has already been shown to have problems with sensor drift over time.

- The CTD5 cast had unreliable conductivity readings in the first 25 metres, probably the result of an air bubble trapped in the cell.

- The intercomparison cast was relatively shallow (175 m).  In all other intercomparisons we have found the water above approximately 300 m to contain too much environmental variance for sensitive intercalibration.

- There is only a single intercomparison cast with no way to determine instrument drift over time.

Despite these limitations, the following analysis will show the comparison of these two casts to be useful.  From the temperature and salinity profiles of cast 89 (Figure 15) we can see that the best regions (low gradients) for the intercomparison are from 30 dBars to 100 dBars and from 140 dBars to 175 dBars.

Pressure

Pressure readings from the two CTDs, compared by matching the pressures associated with key features in the cast profile, showed agreement to within 1 decibar at all depths.

Temperature

The differences between the CTD5 temperature and 1031 temperature were calculated.  From Figure 16 we can see that for the regions where there are not rapid temperature changes with depth, the two CTDs agree to within 0.01 C° which is the specified accuracy of the Sea-Bird CTD.  Therefore, the limited data do not suggest that any corrections need be applied to the 1031 temperatures.

Conductivity/Salinity

The differences between the CTD5 conductivity and 1031 conductivity were calculated (Figure 17).  The large difference at depths less than 25 m is probably the result of a bubble caught in the CTD5 conductivity cell.  However, the conductivity differences below 25 m are due to a calibration problem with 1031 which we corrected with a simple formula (Corrected Conductivity = -0.0397 + Conductivity ´ 1.06249) and the results are shown in Figure 18.  The corrected salinities for 1031 were computed from the corrected conductivities and the difference between the CTD5 salinities and 1031 corrected salinities are shown in Figure 19.  From 40 to 100 dBars and from 120 to 175 dBars, the 1031 corrected salinities agree with the FSI salinities to within 0.02 PSU.

## Chemistry Sampling and Analyses

Samples were drawn from the 10 L BOT bottles on the GO rosette inside a heated container. The order of sampling was: CFC; helium/tritium (copper tube); dissolved oxygen; total CO2; tritium (1L glass bottle); carbon isotopes; nutrients; oxygen isotopes; barium; salinity; chlorophyll a (if sufficient water remaining). Additional bottles were tripped at a given depth when sampling for HCH, TSS, POC/PON, CHLa and radionuclides due to the larger volumes of water required.  Each sample bottle on the rosette was given a unique identifier number of six digits.  Additional bottles tripped sequentially at the same depth were assigned the same identifier number plus a letter A, B, C or D.  CTD casts and water sampling were co-ordinated and the instrumental data at bottle trip depths is reported in Appendices 5.2 and 6.5 together with water chemistry data.  See Appendices 6.3 and 6.6 for chemistry profile plots and Appendix 6.4 for property-property plots.

### Laboratory Methods

The precision of the methods used was estimated by analysing replicates and is expressed as the pooled standard deviation sp, which is calculated as

sp=νisi2νi

and where νi=ni−1 degrees of freedom, and the ni and si, refer to the number of replicates and their standard deviation for the individual components used in the pooled standard deviation calculation.

#### Temperature

Temperatures at the bottle depths were obtained from the CTD records from the same cast.  The pressure was recorded on a log just prior to tripping the bottles and this pressure was used to match to the CTD records. The shallowest CTD records are at 2 decibars.  In the chemistry data tables, the CTD salinities and temperatures at the bottle depths are reported along with the chemistry data. See Section 2.2 for further details on the temperature data.

#### Salinity

Salinity samples were drawn into 200 mL glass salinity bottles, after 3 rinses, from 10 L BOT bottles. The samples were then tightly capped. and shipped back to IOS for analysis on a Guildline Autosal (model 8410).  Data are reported in practical salinity units (psu) [Lewis and Perkin, 1978].  The salinometer was standardized against Standard Sea Water of salinity 34.998 (K15 = 0.99994 which was obtained from Standard Sea Water Service, Institute of Oceanography, Wormley, Godalming, Surrey,  England.  For mission #9324, the overall precision of sampling and analysis was evaluated from a set of 28 duplicates (pairs collected from the same depth and bottle).  The pooled standard deviation (sp) was 0.0025 psu with 27 degrees of freedom.  In the data tables, the CTD salinities at the bottle depths are included.  See section 2.2 and Figures 2, 3 and 4 for a comparison of CTD and bottle salinities.  Differences between salinometer and CTD salinities exhibit more scatter above 300 meters and are not included in this comparison.

#### Dissolved Oxygen

Dissolved oxygen samples were “pickled" immediately in the field and taken to the laboratory on the ship for determination by the Micro-Winkler technique [Carpenter, 1965].  Analyses were carried out within 24 hours of collection.  Calibration of the thio-sulphate solution was performed daily with each titration set by using a Sagami primary standard KIO3.  Precision of the method was monitored during calibration and with blind replicates sampled from the same BOT bottles.  The pooled standard deviation for 37 pairs was 1.73 mM.

#### Nutrients

Water samples for nutrient determination were collected into glass and polystyrene test tubes (2 glass and 2 polystyrene tubes per sample) after three rinses.  Nutrients (silicate, nitrate plus nitrite and orthophosphate) were determined in the laboratory on the ship using Technicon Autoanalyzer II components.  Reactive silicate and nitrate plus nitrite were determined according to Technicon Industrial Methods No. 186-72 W and 158-71 W respectively, and soluble orthophosphate was determined using a modified Technicon method [Brynjolfson, 1973].  Sagami standards were used to calibrate secondary standards which were prepared daily in 30.5 g/L NaCl solutions.  Most of the water samples were analyzed in duplicate and the average is reported in the tables. The precision of the determinations, based on these duplicates, was found to be: silicate, sp =0.38 mM, degrees of freedom =222; nitrate sp =0.10 mM, degrees of freedom =220; phosphate sp =0.02 mM, degrees of freedom =222.

#### Chlorophyll a and phaeo-pigments

Chlorophyll samples were taken down to a maximum depth of 325 m and were subsampled directly from the 10 L BOT bottles into 1 L polybottles. Onboard ship, 1 to 2 L samples were filtered onto 24 mm GF/F filters using a low vacuum filtration.  The filtration castles were rinsed down with 3.2% NaCl and about one mL of a 1% suspension of MgCO3 was squirted onto the filter just before the filtration was complete.  The filters were then folded in half and wrapped in a Whatman filter, labelled, placed in a dark bottle containing dry silica gel and placed in a -20 °C freezer until analysis.  During the filtration and storage, the samples were kept dark as much as possible.  At IOS, chlorophyll a and phaeo-pigment levels were determined fluorometrically with a Turner Design fluorometer (model 10-AU-005) which was standardized with pure Chlorophyll [Strickland and Parsons, 1972].  No deep samples were taken for chlorophyll. Chlorophyll a and phaeo-pigment values were corrected for filter blanks which were treated in exactly the same way as samples and the average filter blank was subtracted from each sample as an equivalent weight (mg) of chlorophyll or phaeo-pigment per filter.

Filter blanks:

0.00089 ± .00091 mg Chla per filter, n=5

0.00369 ± 0.00269 mg Phaeo-pigment per filter, n=5

There were no duplicate samples filtered and analyzed for chlorophyll on this cruise.  The precision, based on duplicates from the 1992 data from this series [Pearson et al, 1994] was for chlorophyll a, sp = 0.11 mg/L and for phaeo-pigment, sp = 0.0059 mg/L for 3 duplicate pairs.

#### Total Suspended Solids

The TSS samples were sampled directly from the 10 L BOT bottles into acid cleaned 2 L polybottles.  Onboard ship, the samples of 4 to 9 L were filtered onto 47 mm, 0.4 mm polycarbonate nuclepore filters, rinsed 3 times with a 3% ammonium carbonate solution, placed on a Petri-slide and stored in a -20 °C freezer.  The nuclepore filters were acid cleaned, DMQ rinsed, dried at 50 °C and pre-weighed on a Mettler M-3 balance to 0.001 mg.  The filters were transported frozen back to IOS where they were dried overnight (12 hours) at 50 °C and weighed on a Mettler M3 balance.  Large differences in TSS values from samples taken on different days and different casts  (station A01) indicate that there may be quite a bit of patchiness in the surface layers.  TSS samples were taken at stations A01, B01, TA, TC, E01 and F09 only.

#### Particulate Organic Carbon and Nitrogen

Samples of 4 to 10 L were collected in 4.5 L glass jugs from the 10 L BOT bottles and filtered as soon as possible using a low vacuum onto 47 mm quartz fibre filters (precombusted at 300 °C for half an hour) and the glass filtration castle was rinsed 3 times with 3.2 % NaCl.  After filtration, the samples were stored frozen at -20 °C in glass petri dishes and brought back to IOS for analysis. In preparation for analysis, the samples were dried at 50 °C overnight and then exposed to fuming HCl in a glass dessicator for 12 hours.  The samples were then again left overnight in a 50 °C oven.  Fresh HCl was used for each batch and seven filter blanks were also exposed to HCl and analysed in the same manner as the samples.  The samples were analysed for carbon and nitrogen on a CE440 Elemental Analyzer.  The edges of the filters were carefully cut off and the filters were cut in half using clean scissors.   The double-drop feature of the instrument was used; each half of the filter was pressed into a nickel sleeve and both halves were dropped onto the ladle and combusted together.  The instrument was calibrated daily with acetanilide and acetanilide was also run as an unknown to check the operation of the instrument.  Filter blanks were processed exactly the same as for the samples except that they were not rinsed with 3.2% NaCl in the field.  Six filter blanks were run with average carbon and nitrogen results as follows :

Carbon blank    =  19.00   mg C  ±  10.15   (n = 6)

Nitrogen blank  =    0.89   mg N  ±   1.05   (n = 6)

The CN ratios appear to be high for station A01 and also, data from the two casts at station A01 show marked differences from each other.

# LIST OF FIGURES

Fig. 1 (a-d) - 9324 Cruise track and station locations

Fig. 2 - Bottle conductivity/ FSI CTD conductivity vs. Pressure

Fig. 3 - Bottle conductivity/ FSI CTD conductivity vs. Julian Day

Fig. 4 - FSI CTD Salinity - Bottle Salinity vs. Pressure

Fig. 5 -  FSI Temperature - CTD5 Temperature vs. Pressure

Fig. 6 - FSI Temperature - CTD5 Corrected Temperature vs. Pressure

Fig. 7 - FSI Conductivity - CTD5 Conductivity vs. Pressure

Fig. 8 - FSI Conductivity - CTD5 Corrected Conductivity vs. Pressure

Fig. 9 - FSI Salinity - CTD5 Corrected Salinity vs. Pressure

Fig. 10 - FSI Temperature - Guildline 53501 Temperature vs. Pressure

Fig. 11 - FSI Temperature - Guildline 53501 Corrected Temperature vs.

Pressure

Fig. 12 - FSI Conductivity - Guildline 53501 Conductivity vs. Pressure

Fig. 13 - FSI Conductivity - Guildline 53501 Corrected Conductivity vs. Pressure

Fig. 14 - FSI Salinity - Guildline 53501 Corrected Salinity vs. Pressure

Fig. 15 - Temperature and Salinity Profile for Cast 89

Fig. 16 - Guildline CTD5 Temperature - Sea-Bird Temperature vs. Pressure

Fig. 17 - Guildline CTD5 Conductivity - Sea-Bird 1031 Conductivity vs. Pressure

Fig. 18 - Guildline CTD5 Conductivity - Sea-Bird Corrected Conductivity vs.

Pressure

Fig. 19 - Guildline CTD5 Salinity - Sea-Bird 1031 Corrected Salinity vs. Pressure

# LIST OF TABLES

Table 1 - 9324 CTD Cast Details

Table 2 - CTD Configurations

Table 3 - Guildline Temperature and Conductivity Correction Coefficients

Table 4 - List of Accepted Uncertainties in CTD Pressure, Temperature, and

Salinity Data

Table 5 - FSI CTD Salinity - Bottle Salinity Statistics

Table 6 - Casts Used in FSI vs. Guildline CTD5 Intercomparison

Table 7 - Casts Used in FSI vs. Guildline 535015 Intercomparison

Table 8 - Casts Used in Guildline CTD5 vs. Sea-Bird Intercomparison

# REFERENCES

Brynjolfson, S.J., 1973.  A modification of the Technicon methodology for the determination of orthophosphate in seawater,  Paper presented at Water Resources Services, Vancouver, B.C.

Carpenter, J.H., 1965. The Chesapeake Bay Institute technique for the Winkler dissolved oxygen method, Limnol. Oceanogr., 10, 141-3.

Fofonoff, N.P., and R.C. Millard, 1983.  Algorithms for computation of fundamental properties of seawater, UNESCO Technical Papers on Marine Science, 44, 141-3.

Lewis, E.L., and R.G. Perkin, 1978. The Practical Salinity Scale 1978: conversion of existing data. Deep Sea Res., 28 A, 307-328.

Macdonald, R.W., R. Pearson, D. Sieberg, F.A. McLaughlin,  M.C. O'Brien,

D.W. Paton, E.C. Carmack, J.R. Forbes and J. Barwell-Clarke, NOGAP B.6, Physical and chemical data collected in the Beaufort Sea and Mackenzie River delta, April-May 1991, Can. Data Rep. Hydrogr. Ocean Sci., 104, 154 pp.

R. Pearson, M. O'Brien, D. Sieberg, F.A. McLaughlin, D.W. Paton, D. Tuele,

J. Barwell-Clarke, E.C. Carmack, R.W. Macdonald and M. Galbraith 1994., NOGAP B.6, Physical and chemical data collected in the Beaufort Sea and Mackenzie River delta, April-May and September, 1992, and ice core data collected in 1991-1992. Can. Data Rep. Hydrogr. Ocean Sci.: 129, 199pp

Strickland, J.D.H. and T.R. Parsons, 1972. A practical handbook of sea water analysis;  Bulletin 167, Fisheries Research Board of Canada, Ottawa, 310pp.

# APPENDICES

## CTD Data: Tables and Plots

The following appendix contains data plots for each CTD cast taken on the 9324 cruise.  Each page displays the data for one cast in four sections:

Header

The header appears in the upper left quadrant of the page.  Each header has 6 lines that appear as follows:

NOGAP 1993

Henry Larsen

STATION             : CS01

REFERENCE NO: 93-24-001

DATE/TIME         : 24/08/93 19:30 UTC

POSITION            : 70-17.5N 133-37.2W

The first two lines are the title and sub-title and are the same for each page.  The remaining four lines are defined as follows:

STATION		-	This is the 4 character station name.  These names match the names given in Table 1 and in Figure 1.

REFERENCE NO	-	This number appears as 93-24-###, where ### is the consecutive cast number as referred to in the report text and listed in Table 1.

DATE/TIME		-	This is the date and time (UTC) at the start of the cast

POSITION		-	This is the position at the start of the cast.

Profile Plot  In the lower left quadrant is the profile plot.  Each profile plot shows Potential Temperature, Salinity, and (if available) Transmissivity, as a function of Pressure.

TS Plot  In the lower right quadrant is the TS plot.  This plot shows Potential Temperature as a function of Salinity.  Also shown on each profile are three lines of constant potential density at 25, 26, and 27 kg/m3, and the freezing point line.

Data Table  In the upper right quadrant is a data table showing various measured and derived properties at selected standard depths.   Units and abbreviations used in the data table are summarized in a table on the next page.

Also included in this appendix are profile plots of % Transmissivity, Chlorophyll, and PAR as a function of Pressure for the FSI CTD casts.

| Quantity | Abbreviation | Units | Notes |
| --- | --- | --- | --- |
| Pressure | Pres | DBAR |  |
| Temperature | Temp | Deg C |  |
| Potential Temperature | Theta | Deg C |  |
| Salinity | Sal | PSS-78 |  |
| Depth | Dept | metre |  |
| Potential Density | Gam-th | kg/m3 |  |
| Geopotential Anomaly | GPA | J/Kg |  |
| % Transmissivity | % Tr | %/metre | Units used for all FSI Transmissometer Data |
| % Transmissivity | %/5cm | %/5cm | Units used for all Sea-Bird Transmissometer Data |
| Chlorophyll | Chl | mg/m3 | Chlorophyll has been measured using a Fluorometer and calculated using nominal coefficients. |
| Photosynthetically Active Radiation | PAR | mE/Sm2 |  |

Units and Abbreviations used in APPENDIX Data Tables

## Chemistry Data Tables

The following appendix contains chemical data from water samples collected on the 9324 cruise.  The tables include the station name, cast number, date and time of cast, the latitude (LAT) and longitude (LONG) recorded at the beginning of the cast, and a record of the instrumentation on the rosette at the time of sampling. The records from this instrumentation (CTD, Transmissometer, Fluorometer, and PAR Sensors), that coincide with the time of bottle trip are included in the tables.

## Chemistry Profile Plots

The following Appendix contains profile plots for the following parameters :

- Bottle salinity

- Bottle salinity vs CTD salinity

- Dissolved oxygen

- Temperature

- Silicate

- Orthophosphate

- Nitrate

- Nominal chlorophyll a (Fluorometer)

- Photosynthetically active radiation (PAR sensor)

- Percent transmission

The profiles are plotted in groups of stations as follows :

- A01, B01, AM10, FM01

- TA, TC, D01, F09, C01

- E01, E04

Each group of stations is plotted twice as follows :

- Surface to 200 or 350 metres to show surface detail

- Complete profile

## Chemistry Property / Property Plots

The following Appendix contains property / property plots as follows :

- Temperature - Salinity (TS plots)

- Oxygen - Salinity

- Silicate - Salinity

- Nitrate - Salinity

- Phosphate - Salinity

- Temperature - Oxygen

## Chemistry Tables - Filtration Samples

The following Appendix contains the data tables for the following filtered samples:

- Particulate organic carbon

- Particulate organic nitrogen

- CN ratios

- Total suspended solids

- Extracted chlorophyll a

- Extracted Phaeo-pigments

Also included in the tables is instrument data collected at the bottle trip depth :

- CTD temperature

- Percent transmission (Transmissometer data)

- Nominal chlorophyll a (Fluorometer data)

- Photosynthetically active radiation (PAR sensor)

The letters in the last three columns indicate which bottle the sample was drawn from (see section 2.3 ).

## Profile Plots - Filtration Samples

The following Appendix contains profile plots of the following data :

- Particulate organic carbon

- Particulate organic nitrogen

- CN ratios

- Total suspended solids

- Extracted chlorophyll a

- Extracted Phaeo-pigments

- Percent transmission (Transmissometer data)

- Nominal chlorophyll a (Fluorometer data)

- Photosynthetically active radiation (PAR sensor)

# Related Data Reports in the Canadian Data Report of Hydrography and Ocean Sciences Series

Macdonald, R.W., K. Iseki, E.C. Carmack, D.M. Macdonald, M.C. O'Brien, and F.A. McLaughlin, 1988. Data Report: NOGAP B.6; Beaufort Sea Oceanography, September, 1986. Can. Data Rep.  Hydrogr. Ocean Sc: 58, 68pp.

Cuypers, L.E., A.W. Blaskovich, E.C. Carmack, and R.W. Macdonald, 1988. NOGAP B.6; Physical data collected in the Beaufort Sea, September, 1986. Can. Data Rep. Hydrogr. Ocean Sci.: 59, 149pp.

Macdonald, D.M., L.E. Cuypers, D. McCullough, E. Carmack, and R.W. Macdonald, 1988, NOGAP B.6: Volume 2: Physical data collected in the Beaufort Sea, March-June 1987. Can. Data Rep. Hydrogr. Ocean Sci.: 60(2), 157 pp.

Macdonald, R.W., K. Iseki, M.C. O'Brien, F.A. McLaughlin, D. McCullough, D.M. Macdonald, E.C. Carmack, H. Adams, M. Yunker, G. Miskulin and S. Buckingham, 1988. NOGAP B.6; Volume 4: Chemical data collected in the Beaufort Sea, Summer 1987. Can. Data Rep. Hydrogr. Ocean Sci: 60(4), 103pp.

Macdonald, R.W., K. Iseki, M.C. O'Brien, F.A. McLaughlin, D. McCullough, D.M. Macdonald, E.C. Carmack, M. Yunker, S. Buckingham, and G. Miskulin, 1988. NOGAP B.6; Volume 5: Chemical data collected in the Beaufort Sea, and Mackenzie River delta, March-July 1987. Can. Data Rep. Hydrogr. Ocean Sci: 60(5), 55 pp.

Carmack, E.C. et al., 1989.  NOGAP B.6: Volume 6: Physical data collected in the Beaufort Sea, Summer 1987. Can. Data Rep. Hydrogr. Ocean Sci.: 60(6), 219pp.

Macdonald, R.W., E.C. Carmack, M.C. O'Brien, F.A. McLaughlin, B.G. Minkley, and K. Berger-North, 1990. Oceanographic Data Collected from the Sir John Franklin in the Beaufort Sea, September 1989. Can. Data Rep. Hydrogr. Ocean Sci.: 80, 100pp.

Macdonald, R.W., E.C. Carmack, F. A. McLaughlin, D. Sieberg, M. C. O'Brien, D. Paton, R. Pearson, Yang Liangfeng, C. Gobeil, 1991, Oceanographic Data Collected from the Henry Larsen in the Beaufort Sea, August-September 1990. Can. Data Rep. Hydrogr. Ocean Sci.: 97, 142 pp.

Macdonald, R.W., R. Pearson, D. Sieberg, F.A. McLaughlin, M.C. O'Brien, D.W. Paton, E.C. Carmack, J.R. Forbes, J. Barwell-Clarke, 1992, Physical and Chemical Data Collected in the Beaufort Sea and Mackenzie River Delta, April - May 1991. Can. Data Rep. Hydrogr. Ocean Sci.: 104, 155pp.

Macdonald, R.W., D. Sieberg, R. Pearson, D. Paton, M. C. O'Brien, F.A. McLaughlin, and E.C. Carmack, 1992, Oceanographic Data Collected from the Henry Larsen in the Beaufort Sea, September 1991. Can. Data Rep. Hydrogr. Ocean Sci.: 112, 108 pp.

Pearson R., M. O'Brien, D. Sieberg, F.A. McLaughlin, D.W. Paton, D. Tuele, J. Barwell-Clarke, E.C. Carmack, R.W. Macdonald and M. Galbraith 1994., NOGAP B.6, Physical and chemical data collected in the Beaufort Sea and Mackenzie River delta, April-May and September, 1992, and ice core data collected in 1991-1992. Can. Data Rep. Hydrogr. Ocean Sci.: 129, 199pp