Centre for Ocean Climate Chemistry

Cruise 9402

Freon Data Report

Wendy Richardson

Ocean Science and Productivity Division

Institute of Ocean Sciences

Sidney, B.C., Canada

December 1999

Introduction

Methods

Sampling

Water

Air

Analysis

Calibration

Standards

Daily Calibration

Results

Station Data

Station P04 - May 12, 1994

Station P12 - May 14, 1994

Station P16 - May 15, 1994

Station P20 - May 16, 1994

Station P24 - May 17, 1994

Station P26  Small Rosette May 19, 1994

Station P26  Large Rosette May 19, 1994

Station P26  (Combined Data)

Tests

Bottle blanks

Air

Freon Levels of Air

Discussion

Comparison of earlier data

Comparison of F11 at Station Papa over time

Comparison of F12 at Station Papa over time

Conclusions

Recommendations

References

# Introduction

The May 1994 COCC cruise to Station Papa was the fifth use of the Freon Analysis System in the field.  The level of contamination during the previous cruise on a Russian research vessel was a concern, and it was necessary to ensure the recovery of the freon analysis system.

# Methods

## Sampling

### Water

Freon samples were taken from 10 liter Niskin bottles only since the large bottles have a better O-ring surface area/water volume ratio than smaller Niskins.  These bottles were well aged, but no specific precautions had been taken to lessen contamination of Freon from O-rings or elastic tubing in these Niskin bottles.  They were however, stored closed with the remnants of the previous sample out on deck where the air is not as heavily contaminated with CFC's.  The Freon samples were the first samples drawn from the Niskin since they are the most easily contaminated by air entering the head space of the bottle.  The samples were drawn directly into 100 ml precision ground glass syringes.

Once the Niskin bottle was on deck, the integrity of the bottle was verified by pushing in the drain valve while the vent remained closed.  If there was no leakage, the syringe valve was inserted into the drain valve and the vent opened.  The syringe was filled with about 10 ml of seawater.  The hydrostatic pressure of the water in the Niskin bottle forced the syringe plunger outward. This reduces the chance of drawing an air bubble into the syringe which is likely if the plunger was pulled out.  The syringe was then removed while the water was still flowing and emptied into the outflow (but not into the Niskin).  This 10 ml rinse was repeated three times.  Then the syringe was filled with approximately 100 ml and the plunger was rinsed with the niskin outflow while expelling the first 80 ml of the large draw.  The three 10 ml rinse technique was repeated and the syringe then filled to about 100 ml, and stopcock closed.  If an air bubble was drawn into the syringe while sampling, the rinses were started all over again.  An elastic was put on the syringe to maintain a positive pressure on the water sample.  The syringe was placed in a sample bucket filled with fresh seawater at least one inch above the syringe barrel.  The bucket of samples was kept in a lab sink, with a continual overflow from the seawater loop manifold.

Hydrocast Freon profiles were carried out at stations P04, P12, P16, P20, P24 and P26.  Bottle blanks were evaluated with 2000 meter water at Station R8 at 51 30 degrees North: 135 00 degrees West.

Carrier blanks, stripper blanks, and restripped samples were also analyzed throughout the cruise.

### Air

Syringe air samples were taken from above the bridge, sampling deck and inside the lab.  The designated air syringe was flushed at least ten times, then a sample drawn, the stopcock closed, and tension elastic put in place.

## Analysis

Analyses are carried out using the Chloroflorocarbon Analytical System built by Glenn Smith, Centre for Ocean Climate Chemistry, Institute of Ocean Sciences.  The COCC analytical system is a modified version of the one described by Bullister and Weiss (1988).  The system has been modified so that the pre column is in its own temperature controlled oven, separate from the Gas Chromatograph oven.  Also an additional valve was automated.

## Calibration

### Standards

Working standard tank number 63102 was used throughout this cruise.  (F11, 530.56 ppt, standard deviation 1.81, n=6, and F12, 370.21 ppt, standard deviation 2.27, n=7)  This standard was made up of outside air and UHP nitrogen using the COCC facilities.  The tank was calibrated against COCC's  lab standard tank number 63088 (F11,  457.59 ppt, standard deviation 0.55; and F12,  263.13 ppt, standard deviation 0.76).  This lab standard was calibrated by Bullister in October 1993.

### Daily Calibration

Calibration curves used to determine sample concentrations were generated by multiple injections of predetermined volumes of standard gas.  Verifying the carrier blank would be the first procedure of the day.  This was followed by a suite of S(small loop)*1, S*2, S*3, S*4, and L(large loop)*1, L*2, L*3.  If these areas looked reasonable based on the previous days' results, a stripper blank was analyzed and then water sample analyses begun.  A L*1 standard was run every few samples, approximately hourly, throughout the day to monitor drift in detector sensitivity.

# Results

Over three hundred Freon analysis runs were carried out during this cruise.  A total of a hundred and five water samples were analyzed.

Data reduction was carried out using the Ocean Chemistry/NOGAP PC adapted computer program based on the Scripps computer program (Weiss).  This program requires salinity and temperature for calculations; the former was taken from Salinometer data; and the latter was read from the sample bucket when the syringe was removed and attached to the extraction system.

Woce Quality Definition

| Value | Definition |
| --- | --- |
| 1 | Sample not analyzed |
| 2 | Acceptable measurement |
| 3 | Questionable measurement |
| 4 | Bad Measurement |
| 5 | Not reported |
| 9 | Sample not drawn from bottle |

Results that do not have a quality of 2 (acceptable) have been omitted from the subsequent plots.

## Station Data

### Station P04 - May 12, 1994

| Sample Number | Pressure dbar | F11 pmol/Kg | F12 pmol/Kg | F11 WOCE Quality | F12 WOCE Quality |
| --- | --- | --- | --- | --- | --- |
| 153 | 1304 | 0.065 | 0 | 3 | 2 |
| 144 | 1200 | 0 | 0 | 2 | 2 |
| 142 | 1000 | 0 | 0 | 2 | 2 |
| 141 | 800 | 0.141 | 0 | 2 | 2 |
| 140 | 600 | 0.203 | 0.540 | 2 | 3 |
| 139 | 400 | 0.790 | 0.421 | 2 | 2 |
| 138 | 300 | 1.384 | 0.691 | 2 | 2 |
| 137 | 200 | 2.663 | 1.308 | 2 | 2 |
| 136 | 150 | 3.209 | 1.535 | 2 | 2 |
| 135 | 100 | 3.831 | 1.981 | 2 | 2 |
| 134 | 50 | 4.142 | 2.311 | 2 | 2 |
| 133 | 10 | 4.000 | 2.006 | 2 | 2 |

### Station P12 - May 14, 1994

| Sample Number | Pressure dbar | F11 pmol/Kg | F12 pmol/Kg | F11 WOCE Quality | F12 WOCE Quality |
| --- | --- | --- | --- | --- | --- |
| 170 | 2400 | 0.158 | 1.087 | 4 | 4 |
| 169 | 2199 | 0.077 | 0.276 | 4 | 4 |
| 168 | 1999 | 0.690 | 0.475 | 4 | 4 |
| 167 | 1799 | 0 | 0.242 | 2 | 4 |
| 166 | 1599 | 0 | 0.357 | 2 | 4 |
| 165 | 1399 | 0 | 0.104 | 2 | 4 |
| 164 | 1200 | 0 | 0.234 | 2 | 4 |
| 163 | 1000 | 0.129 | 0.745 | 4 | 4 |
| 162 | 799 | 0.056 | 0.121 | 2 | 2 |
| 161 | 598 | 0.236 | 0.188 | 2 | 2 |
| 160 | 399 | 1.021 | 0.915 | 2 | 2 |
| 159 | 299 | 1.914 | 1.207 | 2 | 2 |
| 158 | 202 | 3.053 | 1.963 | 2 | 2 |
| 157 | 147 | 3.968 | 2.256 | 2 | 2 |
| 156 | 99 | 4.059 | 2.062 | 2 | 2 |
| 155 | 50 | 4.304 | 2.271 | 2 | 2 |
| 154 | 9 | 4.002 | 2.252 | 2 | 2 |

Both mole sieves became F12 saturated just before Station 12 samples were analysed, this is why the results for deep water were coded as poor.

### Station P16 - May 15, 1994

| Sample Number | Pressure dbar | F11 pmol/Kg | F12 pmol/Kg | F11 WOCE Quality | F12 WOCE Quality |
| --- | --- | --- | --- | --- | --- |
| 199 | 3675 | 0.101 | 0 | 4 | 2 |
| 191 | 1999 | 0 | 0 | 2 | 2 |
| 190 | 1799 | 0 | 0 | 2 | 2 |
| 189 | 1601 | 0 | 0 | 2 | 2 |
| 188 | 1398 | 0 | 0 | 2 | 2 |
| 187 | 1199 | 0 | 0 | 2 | 2 |
| 186 | 999 | 0 | 0 | 2 | 2 |
| 185 | 800 | 0.079 | 0.113 | 2 | 2 |
| 184 | 598 | 0.296 | 0.241 | 2 | 2 |
| 183 | 398 | 1.400 | 0.714 | 2 | 2 |
| 182 | 298 | 2.631 | 1.334 | 2 | 2 |
| 181 | 197 | 3.302 | 1.705 | 2 | 2 |
| 180 | 150 | 3.504 | 1.847 | 2 | 2 |
| 179 | 99 | 4.523 | 2.132 | 2 | 2 |
| 178 | 48 | 5.919 | 2.287 | 2 | 2 |
| 177 | 8 | 5.589 | 2.080 | 2 | 2 |

### Station P20 - May 16, 1994

| Sample Number | Pressure dbar | F11 pmol/Kg | F12 pmol/Kg | F11 WOCE Quality | F12 WOCE Quality |
| --- | --- | --- | --- | --- | --- |
| 222 | 4025 | 0.064 | 0 | 3 | 2 |
| 213 | 2001 | 0 | 0.164 | 2 | 4 |
| 212 | 1801 | 0 | 0.097 | 2 | 4 |
| 211 | 1598 | 0 | 0 | 2 | 4 |
| 210 | 1400 | 0 | 0 | 2 | 2 |
| 209 | 1200 | 0 | 0 | 2 | 2 |
| 208 | 1000 | 0 | 0 | 2 | 2 |
| 207 | 800 | 0.105 | 0.128 | 2 | 2 |
| 206 | 600 | 0.322 | 0.228 | 2 | 2 |
| 205 | 399 | 1.215 | 0.576 | 2 | 2 |
| 204 | 301 | 2.182 | 1.109 | 2 | 2 |
| 203 | 200 | 3.131 | 1.697 | 2 | 2 |
| 202 | 100 | 5.865 | 2.278 | 2 | 2 |
| 201 | 50 | 6.501 | 2.228 | 2 | 2 |
| 200 | 15 | 6.506 | 2.313 | 2 | 2 |

### Station P24 - May 17, 1994

| Sample Number | Pressure dbar | F11 pmol/Kg | F12 pmol/Kg | F11 WOCE Quality | F12 WOCE Quality |
| --- | --- | --- | --- | --- | --- |
| 243 | 3603 | 0 | 0.110 | 2 | 3 |
| 231 | 1200 | 0 | 0 | 2 | 2 |
| 230 | 998 | 0 | 0 | 2 | 2 |
| 229 | 798 | 0.065 | 0.209 | 2 | 2 |
| 228 | 598 | 0.470 | 0.317 | 2 | 2 |
| 227 | 398 | 1.452 | 0.741 | 2 | 2 |
| 226 | 199 | 2.992 | 1.510 | 2 | 2 |
| 225 | 99 | 5.152 | 2.300 | 2 | 2 |
| 224 | 49 | 6.503 | 2.301 | 2 | 2 |
| 223 | 9 | 6.486 | 2.350 | 2 | 2 |

### Station P26  Small Rosette May 19, 1994

| Sample Number | Pressure dbar | F11 pmol/Kg | F12 pmol/Kg | F11 WOCE Quality | F12 WOCE Quality |
| --- | --- | --- | --- | --- | --- |
| 257 | 2001 | 0 | 0.087 | 2 | 2 |
| 256 | 1400 | 0.159 | 0 | 4 | 2 |
| 255 | 801 | 0.100 | 0.119 | 2 | 2 |
| 254 | 400 | 1.595 | 0.701 | 2 | 4 |
| 253 | 399 | 1.881 | 0.900 | 4 | 2 |
| 252 | 199 | 3.135 | 1.336 | 2 | 2 |
| 251 | 152 | 3.692 | 1.678 | 2 | 2 |
| 250 | 101 | 5.241 | 2.075 | 2 | 2 |
| 249 | 75 | 7.005 | 2.499 | 2 | 2 |
| 248 | 51 | 6.674 | 2.303 | 2 | 2 |
| 247 | 26 | 6.697 | 2.492 | 2 | 2 |
| 246 | 11 | 7.785 | 2.723 | 3 | 3 |

### Station P26  Large Rosette May 19, 1994

| Sample Number | Pressure dbar | F11 pmol/Kg | F12 pmol/Kg | F11 WOCE Quality | F12 WOCE Quality |
| --- | --- | --- | --- | --- | --- |
| 278 | 3999 | 0 | 0 | 2 | 2 |
| 268 | 1998 | 0 | 0 | 2 | 2 |
| 267 | 1800 | 0 | 0 | 2 | 2 |
| 266 | 1598 | 0 | 0 | 2 | 2 |
| 265 | 1399 | 0 | 0 | 2 | 2 |
| 264 | 1196 | 0 | 0 | 2 | 2 |
| 263 | 999 | 0 | 0 | 2 | 2 |
| 262 | 998 | 0 | 0.229 | 2 | 4 |
| 261 | 600 | 0.508 | 0.311 | 2 | 2 |
| 260 | 299 | 2.228 | 1.005 | 2 | 2 |
| 259 | 99 | 5.368 | 2.240 | 2 | 2 |
| 258 | 10 | 6.758 | 2.381 | 2 | 2 |

### Station P26  (Combined Data)

The best quality data from the two casts were combined in the following table and plot to show a detailed profile of freon results.  If the data was good for a specific depth sampled during both casts a mean value was used.

| Sample Number | Pressure dbar | F11 pmol/Kg | F12 pmol/Kg | F11 WOCE Quality | F12 WOCE Quality |
| --- | --- | --- | --- | --- | --- |
| 263 | 999 | 0 | 0 | 2 | 2 |
| 255 | 801 | 0 | 0.199 | 2 | 2 |
| 261 | 600 | 0.508 | 0.311 | 2 | 2 |
| 254 | 400 | 1.595 | 0.701 | 2 | 2 |
| 260 | 299 | 2.228 | 1.005 | 2 | 2 |
| 252 | 199 | 3.135 | 1.336 | 2 | 2 |
| 251 | 152 | 3.692 | 1.678 | 2 | 2 |
| 250 | 101 | 5.241 | 2.075 | 2 | 2 |
| 249 | 75 | 7.005 | 2.499 | 2 | 2 |
| 248 | 51 | 6.674 | 2.303 | 2 | 2 |
| 247 | 26 | 6.697 | 2.492 | 2 | 2 |
| 258 | 10 | 6.758 | 2.381 | 2 | 2 |

## Tests

### Bottle blanks

Sampling Location: 51 30 N  135 00 W

May23, 1994 Depth 2000M

| Niskin ID | Sample Number | Syringe | Salinity | Run # | F11 Pmol/Kg | F12 Pmol/Kg |
| --- | --- | --- | --- | --- | --- | --- |
| 22 | 282 | 7746 | 34.580 | 4415 | .064 | .050 |
| 20 | 284 | 7677 | 34.586 | 4418 | .041 | 0 |
| 19 | 285 | 9828 | 34.586 | 4419 | .030 | *.272 |
| 18 | 286 | 9823 | 34.586 | 4421 | 0 | 0 |
| 17 | 287 | 7626 | 34.587 | 4422 | 0 | 0 |
| 16 | 288 | 7633 | 34.587 | 4423 | 0 | 0 |
| 15 | 289 | 7749 | 34.585 | 4424 | 0 | 0 |
| 14 | 290 | 2829 | 34.586 | 4425 | .011 | .188 |
| 13 | 291 | 5576 | 34.586 | 4427 | .035 | .174 |
| 11 | 293 | 5547 | 34.588 | 4428 | .012 | 0 |
| 10 | 294 | 7631 | 34.588 | 4429 | .005 | .007 |
| 9 | 295 | 7676 | 34.589 | 4433 | .063 | .169 |
| 8 | 296 | 7622 | 34.588 | 4434 | *.183 | 0 |
| 7 | 297 | 9838 | 34.587 | 4435 | .074 | 0 |
| 6 | 298 | 9836 | 34.586 | 4436 | .093 | .116 |
| 5 | 299 | 5 | 34.589 | 4437 | .075 | .091 |
| 4 | 300 | 3 | 34.588 | 4439 | .081 | .130 |
| 3 | 301 | 100 | 34.587 | 4440 | .097 | .144 |
| 2 | 302 | 13 | 34.586 | 4441 | .044 | 0 |
|  |  |  |  | Mean | .040 | .059 |
|  |  |  |  | Std | .035 | .074 |
|  |  |  |  | LD | .105 | .222 |

The means and standard deviations were taken from the bottle blank results excluding the two high values with asterisks.  The limit of detection is given as three times the standard deviation of the blanks.

## Air

### Freon Levels of Air

|  | Boat | Deck |  |  | Aft | Deck |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| STN | Run # | F11 ppt | F12 ppt | STN | Run# | F11 ppt | F12 ppt |
| P04 | 4174 | 294.53 | 510.34 | P16 | 4275 | 287.05 | 518.12 |
| P20 | 4315 | 292.53 | 512.35 | P20 | 4316 | 281.83 | 591.77 |
| P24 | 4347 | 289.00 | 494.24 | P26 | 4387 | 281.31 | 532.07 |
| P26 | 4386 | 280.43 | 520.31 |  |  |  |  |

|  | Lab | Air |  |
| --- | --- | --- | --- |
| STN | Run # | F11 ppt | F12 ppt |
| P04 | 4175 | 337.74 | 1000.2 |
| P16 | 4276 | 356.30 | 1630.1 |
| P20 | 4317 | 306.69 | 866.9 |
| P24 | 4348 | 333.54 | 1089.9 |
| P26 | 4388 | 367.42 | 931.6 |

# Discussion

## Comparison of earlier data

Comparison plots of F11 and F12 results from Station Papa over several years are shown on the following pages.  These data are from Gammon (1982), Schwarz (1988) and the data presented in the COCC Cruise 9201 Freon Data Report (Richardson 1993), COCC Cruise 9303 Freon Data Report (Richardson 1993), and this Cruise 9402.  Data Report results are in raw form with no corrections.

### Comparison of F11 at Station Papa over time

### Comparison of F12 at Station Papa over time

# Conclusions

## Recommendations

Considering the space restrictions of the John P. Tully, which result in many people and analyses being carried out in one lab, it is mandatory that no spray cans be used in the lab.  Even without Freon propellants, di-chloro and tri-chloro ethane can cause major contamination problems.  These problems were also encountered and described by U.Schwarz (1989).

Fortunately, during this cruise people refrained from using spray cans, and no work was done to the ships refrigeration systems as on a previous cruise.  With this cooperation it is possible to get excellent quality Freon data results while doing analysis in the main lab on the John P. Tully.

# References

Bullister J.L., Weiss R.F. (1987)  Determination of CCl3F and CCl2F2 in Seawater and Air. Deep-Sea Research Vol. 35 No. 5, 839-853.

Gammon R.H., Cline J.D., Wisegarver (1982)  Chlorofluoromethanes in the Northeast Pacific Ocean: Measured Vertical Distributions and Application as Transient Tracers of Upper Ocean Mixing. J. Geophys. Res. 87(C12), 9441-9454.

Richardson W. (1993)  Centre for Ocean Climate Chemistry Cruise 9201 Freon Data Report, Unpublished manuscript, Institute of Ocean Sciences. Sidney, B.C. Canada.

Richardson W. (1993)  Centre for Ocean Climate Chemistry Cruise 9303 Freon Data Report, Unpublished manuscript, Institute of Ocean Sciences. Sidney, B.C. Canada.

Richardson W. (1995)  Centre for Ocean Climate Chemistry Cruise 9402 Freon Data Report, Unpublished manuscript, Institute of Ocean Sciences. Sidney, B.C. Canada.

Schwarz U. (1988)  Freon data for OC-85-OS-02 and OC-85-OS-03, Unpublished manuscript, Institute of Ocean Sciences. Sidney, B.C. Canada.

Schwarz U. (1989)  Freon Manual, Unpublished manuscript, Institute of Ocean Sciences, Sidney, B.C. Canada.

Weiss R.F.  Freon Lab Manual, Unpublished manuscript, Scripps Institute of Oceanography, San Diego, California, USA