## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 11 April 2024 | ### Changed pad value for DO mass for sample 4. SH |
| 13-Jan-2014 | Added underway pCO2 data from Sophia Johannessen’s Excel files prepared for The Carbon Dioxide Information Analysis Center (CDIAC). The file is located in the cruise .DOC directory. |
| 12-Jun-2013 | Merged DMS data to bottle cast numbers 14, 47, 80 and 90 from Wendy Richardson’s DMS file located in the \DOC directory. For information on methods see file "Method 1996 to 1999 for Dimethylsulfide Analysis.doc” in directory \\OSD_Data_Archive\Cruise_Data\Documents\Analysis Reference Papers\. |
| 13-Feb-2012 | - Added one productivity cast containing Primary Productivity, Chlorophyll, Nutrients, POC and PON, from Frank Whitney’s Productivity spreadsheet. - Merged DIC and Total Alkalinity data to bottle cast numbers 47, 69, 80 and 90 from Jim Christian’s DIC file located in the \DOC directory. For more information, including quality control analysis see John Page’s spreadsheet files in directory \\OSD_Data_Archive\Cruise_Data\Documents\Station Papa Data Archiving\Alkalinity and DIC. |
| 31-Jul-2006 | File 1996-01-0047.che CTD and bottle salinity seriously disagreed. Log book notes that conductivity sensor was bad on this day so the CTD salinity was removed. J.L. |

Processing Notes         August 13, 1996

Cruise 96-01

Location: N.E. Pacific, Line P

Project: Papa

Chief Scientist: Frank Whitney

Vessel: Tully

Processed by: Germaine Gatien

Date of Processing: July 11, 1996 - August 13, 1996

Number of Original Casts: 89 (Numbered 1-92 with 19-22 missing)

Number of Casts Processed: 41

------------------------------------------------------------------------------

INSTRUMENT SUMMARY

------------------------------------------------------------------------------

The Following Instrument Was Used For Casts: 1,11,13,44,46,68,71,79,89

Instrument Type         : Guildline WOCE

Instrument Model        : WOCE

Instrument Serial Number: 59901

The Following Instrument Was Used For Casts: 5,7,9,15,22,24,26,28,30,32,

34,36,38,40,42,48,52,54,56,58,60,62,64,66,73

Instrument Type         : Guildline 8715

Instrument Model        : DEL433DE

Instrument Serial Number: 58483

The Following Instrument Was Used For Casts: 75,77,81,83,85,87,91

Instrument Type         : Guildline 8715

Instrument Model        : DEL433DE

Instrument Serial Number: 53,501

-----------------------------------------------------------------------------

SUMMARY OF QUALITY AND CONCERNS

----------------------------------------------------------------------------

Of the three CTDs used during this cruise only the Guildline #53501

performed well for all casts. (There were 7).

Four of the 9 casts using the Guildline WOCE CTD had very poor

salinities; the S channel had to be removed from 3 of these.

There were conductivity cell problems in the second half of the cruise.

The log indicates that there was a problem with the copper thermometer

during cast 46 and at cast 89 the conductivity cell failed and

thermistors failed on both channels. The log also mentions problems

with cast 71 with very noisy salinity.

The Guildline #58483 was used for 25 casts. The log mentions trouble

with noisy salinity in casts 48 and 64 . (While noisy, the average

values are close to expected ones.) At cast #66 there was trouble with

the copper thermometer and at #73 it failed at depth. There was a

problem with the pressure records at #30, and cast #56 was full of

spikes in T, S and Tr.

The salinity channel had to be removed from 7 casts in the second half

of the cruise: 46,48,64,66,71,73,89.

The salinity is very noisy in casts: 13 and 28.

Cast #56 was heavily edited.

------------------------------------------------------------------------------

PROCESSING SUMMARY

------------------------------------------------------------------------------

1. Preliminary Steps

The log and calibration information were found. These, and the Header

Check and Header Summary were examined for errors.

A few errors were corrected in the .RAW files:

-	The transmissometer calibration was wrong in casts #71 & #72.

This was corrected in the .RAW files.

-	Due to GPS failure there was no position for cast #90; this was

an upcast so the position from the downcast was substituted.

SURFACE was run to establish which were downcasts, from which lists

were prepared of downcasts for each CTD. SURFACE was rerun for each

CTD establishing that no correction of surface pressure was needed.

Two positions from the Cruise Track were checked in the log, and found

to be correct.

The data was converted to IOS HEADER format

The following casts were deleted:

Cast 3 was very noisy and was a repeat of cast 1

Casts 17, 50 and 70 were very shallow (10m).

Casts 2,4,6,8,10,12,14,16,18,23,25,27,29,31,33,35,37,39,

41,43,45,47,49,51,53,55,57,59,61,63,65,67,69,72,74,76,78,

80,82,84,86,88,90,92 were upcasts.

2. Calibrations (program CALIB)

Guildline WOCE CTD calibrations from Dec. 3, 1995 were used for

casts 1,11,13,44,46,and 68:

Name                      Units Fmla Pad    Coefficients

------------------------- ----- ---- ------ ------------

Pressure                  n/a     10 n/a    (0 1 0)

Temperature               n/a     10 n/a    (-5.93725 .38917E-04

.48229E-13)

Conductivity_Ratio        n/a     63 n/a    (-.519E-03 .16918E-05

.288885E-16)

Transmissivity            n/a     10 n/a    (.3297225 .31978E-02)

Temperature:thermistor1   n/a     34 n/a    (3592.774 -.6542941E-01

.1017114E-02 .2943959E-03

.1568311E-06 .93E-02)

Temperature:thermistor2   n/a     34 n/a    (3629.344 -.766553E-01

.1045541E-02 .2903017E-03

.1588842E-06 .91E-02)

The same Guildline WOCE calibrations were used for casts 71,79 and 89

EXCEPT that the transmissivity calibration were as follows:

Name                      Units Fmla Pad    Coefficients

------------------------- ----- ---- ------ ------------

Transmissivity            n/a     10 n/a    (.310275 .31234E-02)

Guildline 58483 calibrations from Jan. 18, 1995 were used for

cast #s 5,7,9,15, 22,24 and 26

Name                 Units Fmla Pad    Coefficients

-------------------- ----- ---- ------ ------------

Pressure             n/a     10 n/a    (-4 3000)

Temperature          n/a     10 n/a    (.403934E-01 1.000189)

Conductivity_Ratio   n/a     63 n/a    (-.961E-03 1.000727)

Transmissivity       n/a     10 n/a    (112.785 -2.25659)

The same Guildline 58483 calibrations were used for

cast #s 28,30,32,34,36,38,40,42,48,52,54,56,58,60,62,64,66,73

EXCEPT for the tranmissivity coefficients, which were as follows:

Name                 Units Fmla Pad    Coefficients

-------------------- ----- ---- ------ ------------

Transmissivity       n/a     10 n/a    (100.785 -2.01731)

Guildline 53501 calibrations from ???? were used for casts 75,77,81,83,

85,87,and 91

Name                 Units Fmla Pad    Coefficients

-------------------- ----- ---- ------ ------------

Pressure             n/a     10 n/a    (0 2000)

Temperature          n/a     10 n/a    (-.3886E-02 1.0006)

Conductivity_Ratio   n/a     63 n/a    (.12967E-02 .9985384)

Transmissivity       n/a     10 n/a    (-100.785 2.01731)

3. Despiking (program DESPIKE)

Advance Despike was run on all casts using the following DESPIKE

parameters:

DESPIKE TABLE

CHANNEL              FIT   OVER-  MIN     MAX    MIN     MAX     SPIKE

WIDTH LAP    VALUE   VALUE  STDDEV  STDDEV  TOL   REP.

-------------------- ----- ----- ------- ------- ------- ------- ----- ----

Pres                    55     5   -2.00 4500.00  .10000 999.000  3.00   T

Temp                    55     5    1.00   11.00  .00500 999.000  3.00   T

Cond                    55     5     .00    1.50  .00010 999.000  3.00   T

Tran                    55     5     .00   80.00  .10000 999.000  3.00   T

4. Time Compensation

TIMECOMP was run for all casts using Temp. Probe Dist = .07

Sample Period    = .04

5. Delete

DELETE was first run for all casts using the following paramaters:

Surface Record Removal: Low Sal & Last Press Min

Maximum Surface Pressure (relative):      20.00

Minimum Surface Salinity:     25.000

Surface Swell Pressure Tolerance:       1.00

Pressure filtered over width:    21

Swells deleted. Warning message if pressure difference of      2.00

Drop Rate < 0.03 (calculated over 61 points) will be deleted.

Sample interval = 0.0

DELETE was rerun on casts #30 and #56 after 30.tim and 56.tim had

been hand-edited to remove a jump in pressure, and many large spikes

respectively. On the second run of DELETE pressure filter widths of 5

and 11 were used and Low Drop Rate was not deleted because records had

been removed.

6. Filtering

The salinity was filtered using a median BOXCAR. The lowest

filter width that would remove the spike in S in cast #60 was

21. This was then applied to all casts.

After discussion with Frank Whitney it was decided to use the

same filter for transmissivity. It is hoped that this will leave

evidence of "biological" spikes, while eliminating instrument noise.

Exceptions: for cast #1 a filter width of 51 and for cast #26 a width of

61 was applied to both Tr and S to remove spikes.

7.Bin Average

Bin Average was run on all casts (Parameters:DEL,PRES,1,0,NO,NO)

8.Hand Editing and Test Plots

Page plots with T-S plot were produced for all casts.

Examination of these plots showed that many casts needed

hand editing which was done using CTDEDIT.

The following casts had salinity edited: 1,13,15,22,30,32,34,

38,44, and 56.

Temperature and Transmissivity were also edited in casts #30 and 56.

9.Comparison with Bottle Data

The .BOT files were calibrated as in step #2 to produce .CAB files;

then sample numbers were added to produce .SAM files

which provided the secondary source for the Comparison routine.

Bernard Minkley prepared .HYD files as the primary source.

There were 7 hydro files available from Rosette sampling with

the WOCE CTD. The associated CTD casts from 4 of these had salinities

that were too noisy to use with the COMPARE routine.

The three remaining casts were entered into COMPARE and the resulting

file WOCE.CSV was imported into the COMPARE Macros in Excel.

The results were not useful. Cast #2 was very shallow.

Casts #69 and #80 showed very different results.  The conductivity

cell on the WOCE CTD was behaving very badly for the later casts and

the results are not reliable enough to justify using the fit

calibrations derived with COMPARE.

10. Intercomparisons

Comparisons were made of repeat casts at a station. In all cases the

transmissivity was quite different.

a.) casts 11,13,15 - 11 was shallow. 13 and 15 were very similar.

b.) casts 36,87,89 - cast 89 had a very noisy salinity profile.

Cast 36 was much like 87 with slight lower salinities. Though they

occurred 12 days apart they were much closer than 87 and 89 which

occurred sequentially.

c.) casts 44,46,48,85 - in this group 46 stands out as odd. The

S and T profiles are noisy and S is much higher than at the other

stations 46 is also noisy in S. Cast 48 is also very noisy.

d.) casts 66 and 81 - these look totally different. These stations

were compared with nearby casts. 81 is very similar to 77 and 79,

while 66 has much higher salinities.

e.) casts 68,71,73,75,77,79 - cast 73 has lower S. The other casts

look similar.

So casts 46, 66, 73 and 89 are suspect. Casts 46 and 89 used the

WOCE CTD while 66 and 73 used the Guildline #58483.

Plots of several profiles from nearby stations were made to look for

other irregularities and no other large discrepancies were found except

in transmissivity.

11. Remove channels

Conductivity was removed from all casts.

Salinity was removed from casts: 46,48,64,66,71,73,89.

12. Produce Final Files

Files were renamed .CTD and notes were made in the headers if salinity

was noisy or the salinity channel was removed.

A cross-reference listing was produced.

Data archive on tapes DLT017 and DLT018 in saveset 9601_all.bak.

13. Additional corrections Dec 12/96 RMBROWN

96010007.CTD -  remove record #1 - bad SALINITY , low REC COUNT

-   delete repaeting records at Pres=113.04 (bottom of

cast)

96010038.CTD - remove record #1 bad surface TEMP, low REC COUNT

14. Review of data for profiles 11,13 and 15 (p04)

file:971500015.DOC

date: Oct 02,1997

Temperature anomaly at Station P04 (profiles 11, 13,15)

-------------------------------------------------------

re: 960010015

Claude Savenkoff noted a large region (75 metres to 120 metres depth) where

the temperature seemed anomalously high, when compared to previous stations

at this location:

96010011 P04 21/02/96 22:30 UTC WOCE CTD/Rosette

96010013 P04 22/02/96 00:50 UTC WOCE CTD/Rosette

96010015 P04 23/02/96 03:08 UTC Guildline Standard CTD only

I have reviewed the data and can find no reason to EXCLUDE these data.

Although a different CTD was used, the agreement is good in the surface mixed

layer and at depth amongst these casts. Furthermore, a simple error in

temperature measurement should cause a large error in slainity, which

does not seem to be the case. Imspection of the T-S plots does not

reveal unstable water structure.

RMB

# Moving files to new archive. May 8, 2003 – Joe Linguanti

- Chlorophyll data from Frank Whitney’s 9601-HYD.XLS spreadsheet was merged with the bottle data for 4 files. Frank’s original file is in the DOC directory.

- Program HEADER EDIT was used to change non-standard channel names and units and add additional metadata to the CTD and Rosette files.

- Program Clean was used to;

- reset number of records, minimum and maximum values

- delete empty channels where applicable

- set channel name initial letters to uppercase

- Files .CLN were copied to new archive and renamed to the new file convention “YYYY-CC-NNNN.* “.

- The loop file was obtained from Frank Whitney. The latitude and longitude values were computed to 3 decimal places. The spreadsheet file was then converted to IOS Header format with meta-data added to it. The original .XLS file is in the DOC directory.

- The SAIL files were copied from the TRACKOBS directory in the OSAP Data Archive.

- Channels not needed were removed.

- Program HEADER EDIT was  used to change non-standard channel names and units and add additional metadata to the files.

- IOS Header Date and Time channels were added.

- Program Clean was used to set the start and end times and the geographical area.