Cruise 9608: Line C and Line BP, Nutrient Data

Comments

File storage and analysis methods

Files are:

000202.aaz, .anz, .xls: data for samples analyzed Feb 2/00.

Nutrients were analyzed following procedures outlined in Barwell-Clarke and Whitney, 1996.

These samples had been frozen in isopropanol immediately after collection and stored in a freezer until analyzed.  The long storage time may affect the quality of the data.

Nitrate/silicate and phosphate were analyzed separately; the former in plastic test tubes, the latter in glass.

Standards and Pipettes

Standards were prepared using an Eppendorf 1-10 mL maxipettor which had been calibrated on

Jan 26/00.  Although each tip delivered close to the dialed volume, the concentrations of the standards have been corrected for actual volume delivered.

The nitrate standard, #2000-2,  was a 2500.4 uM solution of KNO3, prepared from preweighed, desiccated chemical on Jan 28. Silicate  standard #98-2 was a 5000.3 uM solution of Na2SiF6, prepared from oven-dried preweighed chemical on Jan 27/00.  The phosphate standard, #99-7, was a 2500 uM solution of KH2PO4 which was prepared from preweighed chemical on Jan 28/00. It was further diluted 100:1000 to a final concentration of 250 uM.  All standards were compared to Sagami standards and/or previously prepared standards on Jan 28/00.

All comparisons were within acceptable limits.

Standards for this day's analysis were prepared in plastic volumetric flasks.

Variability

No duplicate samples were taken nor analyzed.  A sample of water taken at 125m at Station MP4 in June, 1999 is measured periodically as an external measure of drift.

Variability of P4 water measured as an unknown:

nitrate		silicate		phosphate

P4	27.0777	39.34383	2.000983

P4	26.75466	39.27697	1.993096

mean	26.92		39.31		2.00

stddev	0.23		0.05		0.01

The variability is a bit high for nitrate.

Analysis notes:

Phosphate peaks oddly shaped for samples:  96-16, 100-16.

Suspect alcohol in samples:PO4: 174-10, 180-5.

Salty around lid area in many of the samples; see original data notes. Particularly, N: 102-11, 10; P: 174-2.

Insoluble crystals in most samples; see original data notes.

Lid loose on phosphate samples: LC09-16, 8, 7, 6; cast 102-4, 3; LBP2-6, 5; LBP4-11, 7; LC97-6.

Fill level was low for all samples, Casts 96 (LC11), 19 (LBP4), 178 (LC05), 180 (LC04).

Temperature on top of manifold was about  26.7-27.0 C

The phosphate channel for this day's analysis was surging significantly, resulting in some oddly shaped peaks.

The phosphate molybdate reagent connector kept slipping off, requiring repeated measurements for the affected samples.

Phosphate channel had been giving poor quality data on previous days' runs.