Cruise 9614: Line C Nutrient Data

# Comments

File storage and analysis methods:

Files are:

000201.aaz, .anz, .xls: data for samples analyzed Feb 1/00.

Nutrients were analyzed following procedures outlined in Barwell-Clarke and Whitney, 1996.

These samples had been frozen in isopropanol immediately after collection and stored in a freezer until analyzed.  The long storage time may affect the quality of the data.

Nitrate/silicate and phosphate were analyzed separately; the former in plastic test tubes, the latter in glass.

# Standards and Pipettes

Standards were prepared using an Eppendorf 1-10 mL maxipettor which had been calibrated on

Jan 26/00.  Although each tip delivered close to the dialed volume, the concentrations of the standards have been corrected for actual volume delivered.

The nitrate standard, #2000-2, was a 2500.4 uM solution of KNO3, prepared from preweighed, desiccated chemical on Jan 28. Silicate standard #98-2 was a 5000.3 uM solution of Na2SiF6, prepared from oven-dried preweighed chemical on Jan 27/00.  The phosphate standard, #99-7, was a 2500 uM solution of KH2PO4 which was prepared from preweighed chemical on Jan 28/00. It was further diluted 100:1000 to a final concentration of 250 uM.  All standards were compared to Sagami standards and/or previously prepared standards on Jan 28/00.

All comparisons were within acceptable limits.

Standards for this day's analysis were prepared in plastic volumetric flasks.

Variability:

No duplicate samples were taken, nor were duplicates intentionally read. However,

where measurements had to be repeated in one channel, the other channels' data

are reported.

A sample of water, collected at Station P4 at 125m in June, 1999 (Cruise 9910) is routinely measured as an external measure of drift.

Deviation of a sample of water run several times throughout the analysis day:				N            Si        P

P4	27.1	39.2	2.03

P4	27.4	39.6	2.03

P4	27.1	39.2	2.08

average27.2	39.3	2.04

stddev	0.16	0.23	0.030

Analysis notes:

There were no samples for: LC11 - 1, LC12-1, LC10-1, LC9-1, LC9-4, LC2-4.

Phosphate peaks oddly shaped for samples: LC11-10, LC11-6, LC11-2, LC10-10, LC7-7, LC7-6, LC7-5, LC7-1, LC3-5, LC3-4.

There was significant surging in the phosphate channel.

There may have been some debris in the PO4 channel; cleaning at the end of the day seemed to improve the peaks for 000202.

Suspect alcohol in samples:LC7-7, LC7-6, LC3-2.

Salty around lid in many of the samples (see original analysis notes); particularly in:PO4: LC10-10, LC10-6, LC3-5, LC3-4, LC3-1.

Insoluble crystals in most samples; see original data notes.

Lid loose on phosphate samples:LC11-5, LC7-8, LC4-2, LC3-3.

Temperature on top of manifold was about  26 C