Cruise 9638: Laparousse Bank.  Nutrient Data.

File storage and data processing:

970312	.aaz	.anz	.wk4		data analyzed March 12, 1997

970314	.aaz	.anz	.wk4		analyzed March 14

970318	.aaz	.anz	.wk4		analyzed March 18

970319	.aaz	.anz	.wk4		analyzed March 19

970325	.aaz	.anz	.wk4		analyzed March 25

970326	.aaz	.anz	.wk4		analyzed March 26

970327	.aaz	.anz	.wk4		analyzed March 27

9638nutr				.xls	combined nutrient file

Samples were measured using an Autoanalyzer II, following the procedures outlined in Barwell-Clarke

and Whitney (1997).  Nitrate/silicate samples were defrosted about 24 h before analysis.  Phosphate

samples were rapidly defrosted in moving room temperature air or warm water.  Lab temperature was

noted periodically, and may be found on the original data sheets; it was generally around 26C.

Calibration:

Standards used for these samples had been previously prepared and compared to Ocean Science

International commercially prepared standards on January 31, 1997 (file 970131.aaz).  The nitrate

standard concentration was approximately 99.4% of the OSI standard; silicate and phosphate were

slightly more concentrated than the OSI standards, 110.65 % and 100.48%, respectively. A minor

correction may be required for the silicate data after this calibration has been verified.

The silicate standard (#960813-1) was a 7500.27 uM (1.4105 gm/l) solution of Na2SiF6 prepared from

preweighed chemicals on Jan 17, 1997.  Nitrate standard, #96-3, was a 2500.3 uM (0.2528 gm/l) solution of KN03

prepared on Jan 17, 1997.  Phosphate standard, #96-2, was a 2499.8 uM (0.3402 gm/l) solution of KH2PO4,

prepared in Nov. 1996 and diluted to 249.98 uM at that time.  The 249.98 uM solution was used to prepare

the working standards.

An Eppendorf maxipettor (Whitney's #3) was used to pipette volumes greater than 1 ml.  The pipettor

was calibrated on Jan 15, 1997 and again on March 24, 1997 for each pipette tip.  The delivery rate had

remained more or less constant for the second calibration, and was very close to the volume dialed.

The same tips were used throughout the run,although they were acid washed periodically.

An Eppendorf 100-1000 uM pipettor (Ocean Chem, #66061) was used to pipette volumes less than 1.0 ml.

Variability Estimates

Estimates of machine variability, based on standards run as unknowns:

Low std.	Nitrate	Silicate	Phosphate

Expected	12.00	45.00	0.900

Mean	12.19	44.33	1.066

Std. deviation	0.17	0.20	0.941

Medium std

Expected	24.00	90.00	1.800

Mean	24.24	90.05	1.779

Std. deviation	0.26	0.45	0.031

High std.

Expected	36.00	135.00	2.700

Mean	36.14	135.34	2.691

Std. deviation	0.17	0.95	0.034

Nitrate standards were generally higher than expected.  The variability was fairly constant for

nitrate and phosphate, but increased with concentration for silicate.

PO4 low standard variability was probably overestimated due to one errant measurement.

970325, 970326: PO4 standards which were run as unknowns were lower than expected. This may be because

a different batch of NaCl was used in the rinse water than was used to prepare the standards.  Although

care was taken to establish the baseline using the old NaCl, it may have been influenced by the rinse water, which

was about 15 pts higher.

No duplicate samples were taken, therefore there is no estimate of sample variability.  It is strongly

recommended that at least 10% of the samples be duplicated, in order to assess variability, which can

be particulary high when the samples are rapidly frozen in isopropanol.

Factors affecting data quality:

970314: N/Si: 121-129 unable to analyze samples on 970312, were stored refrigerated, capped, until 970314.

The data look acceptable, but the nitrate may be a little low. Phosphate were samples analyzed in the usual manner.

Nitrate standards and peaks for this series may be more variable than usual due to minor

instability of the trace.  Smoothing was applied to peak analysis for nitrate.

Some nitrate surface samples were lower than the NaCl baseline.  In such cases, the baseline was

adjusted to assume that the lowest sample was 0.

These samples had been rapidly frozen in isopropanol, and as is usual with the method, many of the samples

were contaminated with alcohol.  The worst stations, LJ02, LJ01, LBP2 and LI10 had enough alcohol in

the PO4 tubes to pour off.  There was also alcohol in the bags.  To maintain good quality data, it is imperative

to monitor the level of alcohol in the coolant bath, and to drain and wipe the tubes dry.