## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 3-Feb-2014 | The CTD salinity channel was removed from the Bottle files since the salinity data from the rosette files was not processed or calibrated. For details see section “Rosette Files, April 30, 2003” in file “\\OSD_Data_Archive\Cruise_Data\Documents\CS Wong’s Archive\ Station Papa Data Arching 2002.doc”. J.L. |
| 13-Jan-2014 | Added underway pCO2 data from Sophia Johannessen’s Excel files prepared for The Carbon Dioxide Information Analysis Center (CDIAC). The file is located in the cruise .DOC directory. |
| 14-Jun-2013 | Merged DMS, DOC and DON data to bottle casts from Wendy Richardson’s spreadsheet file located in the \DOC directory. For information on methods see file "Method 1996 to 1999 for Dimethylsulfide Analysis.doc” in directory \\OSD_Data_Archive\Cruise_Data\Documents\Analysis Reference Papers\. |
| 13-Feb-2012 | - Added one productivity cast containing Primary Productivity, Chlorophyll, Nutrients, POC and PON, from Frank Whitney’s Productivity spreadsheet. - Merged DIC and Total Alkalinity data to bottle cast numbers 10, 34, 62, 87, 123 and 150 from Jim Christian’s DIC file located in the \DOC directory. For more information, including quality control analysis see John Page’s spreadsheet files in directory \\OSD_Data_Archive\Cruise_Data\Documents\Station Papa Data Archiving\Alkalinity and DIC. |
| 18-Jul-2011 | Jim Christian noted that the bottle and CTD salinities were offset by about 0.02 in the rosette files (.CHE). The processing notes mentioned that the CTD files were recalibrated by subtracting about 0.022 for CTD #0443 and adding about 0.018 for CTD #437. There was no evidence in the rosette files that the CTD salinities had been recalibrated. Therefore since all the rosette files were taken with CTD #0443, the CTD salinities in the rosette files were recalibrated using the same coefficients as those used to recalibrate CTD #0443. As mentioned above, 2 CTD instruments were used 0443 and 0437. The Instrument Serial number in all the files contained both numbers. For all the .CTD files I modified the Instrument section in the header to contain only the CTD Serial Number that was used for that particular cast. J.L. |
| 08-Mar-2011 | Corrected a 10X offset in Phosphate for bottle casts 164 and 166. J.L. |
| May 7, 2003 | Added chlorophyll and ammonium data from Frank Whitney’s spreadsheet file, 9711-HYD.xls, to 20 bottle files. J.L. |
| May 6, 2003 | Moved files to new archive. |
| 5-Jun-2000 | Removed chlorophyll channel since it did not contain chlorophyll as stated, rather we suspect it contained PO4 as umol/kg. J.L. |
| 26-May-2000 | Removed Ammonia channel from .CHE files since data was of questionable quality. J.L. |

PROCESSING NOTES

Cruise: 9711

Agency: OSAP/JGOFS

Project: Line P - Station Papa

Chief Scientist: Philip Boyd

Platform: J P Tully

Processed by: Germaine Gatien

Date of Processing: 10 July 1997 - 15 August 1997

Number of original casts: 166

Number of casts processed:  81

INSTRUMENT SUMMARY

Two SeaBird Model SBE 911+ CTDs were used: #0443 AND #0437. Both were used with the rosette sampler.

SUMMARY OF QUALITY AND CONCERNS

There are problems in the calibrations for conductivity for both CTDs.  The data from each sensor is consistent showing similar values for repeat casts and at adjacent sites. But the primary and secondary sensors are very different for CTD #0443 and none of the salinity values from either CTD compare well with the bottles.  The errors in salinity are approximately as follows:

CTD #0443    Primary Salinity         0.022 PSU too high

Secondary Salinity     0.042 PSU too low

CTD #0437    Primary Salinity        0.023 PSU too low

Secondary Salinity     0.018 PSU too low

After re-calibrations the difference in salinity at depth for a repeat cast using  the different CTDs was less than 0.002 PSU around 2500 db.

The temperature sensors are very consistent.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained and  note was made of problems that occurred during the cruise.

The salinity bottle data sheets were found.

The cruise summary sheet was completed.

3. Conversion of Raw Data

This step was completed by Bernard Minkley using the calibration control files 97110001.con for CTD #0443 and 97110017.con for CTD #0437.

4. Preparing  a set of data for further processing

Files *.cnv were copied into *.org; this set of files was left unprocessed.

The *.cnv files for upcasts were deleted and processing continued on downcasts only.

5. Test Plots

To check the agreement of the pairs of temperature and conductivity sensors plots were made of a selection of casts for each of the instruments used.

a.) For CTD #0443 casts examined were: 1,11,61,83,122,149. At depth temperature variations were <0.001 C.deg. Conductivity varied by about 0.006 units resulting in salinity variations on the order of 0.07 PSU.

b.) For CTD #0437 casts examined were: 23,53,77,125. At depth temperature variations were <0.001, conductivity <0.001 units and salinity <0.008 PSU.

These comparisons indicate a problem with the conductivity calibration for CTD #0443. The temperature sensors were in good agreement as was the conductivity for CTD #0437.

6. WILDEDIT

Program WILDEDIT was used to remove spikes in T, C, S and Tr. This stage of processing is to flag very bad points. Parameters used were:	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

7. CELLTM

The conductivity cell thermal mass correction was done using alpha = 0.03 and 1/beta = 7.0.

8. FILTER

The pressure was low-pass filtered with a time constant of 0.15 seconds to increase pressure resolution for LOOPEDIT.

9. LOOPEDIT

This routine was used to remove scans for which the CTD was moving at less than the minimum velocity of 0.33 m/s or traveling backwards due to ship roll. In future processing this step will be dropped.

10. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

11. Checking Headers

The first check of headers indicated that the transmissivity data was missing from many casts and the conversion had to be re-run. To avoid going so far in the processing before finding such an error, I would suggest that in future a random check be done during step 5 to make sure that all expected channels are present in the casts chosen for test plots. After the data was reconverted and put through steps 6 through 10 the check of headers turned up no errors.

12. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative):      20.00

Surface Swell Pressure Tolerance:        .50

Pressure filtered over width:   11

Swells deleted. Warning message if pressure difference of      2.00

Drop rates<    .33m/s (calculated over  11 points) will be deleted

Sample interval =       .04 seconds.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

14. THIN

The following parameters were used for program THIN:

Reference Channel: press

Using Standard Depths for Thin Values

Tolerance for finding values:       .500

Average of Values within Tolerance will be Output

15. Test Plots

Profiles of all casts were examined (on screen) for large inconsistencies between primary and secondary values and for any evidence of problems with the processing. As noted earlier there is a larger discrepancy than expected between S0 and S1 for CTD #0443. The secondary salinity looks quite messy for casts 77 and 124 and both primary and secondary looks messy for cast #133 in the upper 130 metres.

A cursory intercomparison was done to choose which values to use for the page plots. For CTD #0443 there are bottle salinities; these indicate that the primary values are best.  For CTD #0437 there were no bottles so a choice was made by comparing casts at one site using the 2 CTDs and this showed that the secondary values were closer to the bottle values. Primary values will be processed for #0443; secondary, for #0437 except for casts #77 and 124.

Page plots were produced using primary or secondary values as determined above.

16. CTDEDIT

The following casts were cleaned only near the surface: 1,5,7,13,15,17,31,57,63,86,118,141,

The following casts had more extensive editing: 3,11,21,23,27,29,53,55,59,67,71,81,83,90,100,108,122,126,128,134,139,147,153,155,165.

Explanatory notes were made in any edited file under the REMARKS heading.

The edited files (*.edt) and those not edited (*.avg) were combined in a set identified as *.plt.

17. Preparing files for intercomparison

The *.ros files were converted to IOS header files and renamed as *.bot.

A cast number pairs file and a pairs list were prepared and used to add bottle sample numbers to the *.bot files producing the *.sam files.

18. COMPARE

The COMPARE routine was used to produce a file 9711COMP.CSV which was imported into EXCEL using the COMPARE macros. The salinity differences were plotted against pressure excluding shallow data. The CTD salinity data was  consistently about 0.022 PSU higher than the bottle values.  COMPARE was used to prepare  recalibration file 9711comp.ctr. This was applied to all *.sam files to produce files that were called *.sac. This corrected data was once again compared with the bottle data in COMPARE and produced differences of less than 0.004 PSU below 200 dbars with most values within 0.002 PSU. The recalibration was then applied to all data from  CTD #0443.

There were no bottle casts with CTD #0437 so COMPARE could not be used to correct the salinity.

19. Other intercomparisons

The following intercomparisons were made of salinity values before the COMPARE recalibration was applied:

1. Repeat casts at P10 using CTD #0437 showed variations of 0.02 PSU at 1000m. and 0.01 at 2000m.

2. Repeat casts at P26 using CTD #0443 showed variations of 0.02 PSU at 1000m. and <0.01 at 2000m.

3. Casts at adjacent sites (Z06,Z07, Z08) using CTD #0437 showed variations of 0.01 PSU.

4. Repeat casts at site Z09 using the two different CTDs showed variations of 0.05 PSU. After the recalibration of CTD #0443 the variation was about 0.03 PSU.

5. To quantify the error in CTD #0437 theta-S plots were examined for 2 casts at adjacent stations at which different instruments were used (casts # 106 and 108).These plots were very similar in shape so a plot was prepared of salinity error for constant values of theta versus pressure. This indicates that a salinity correction of 0.018 PSU will correct the differences to within 0.002 PSU below 600 db. This correction was applied to secondary salinity  for casts using  CTD #0437. For the casts  #77 and  #124  primary salinity must be used and the salinity correction is 0.023 PSU.

20. Final Plots

Page plots were prepared using the recalibrated data.

21. REMOVE

The following channels were removed from all casts: Scan_Number, Conductivity:Primary, Conductivity:Secondary and Flag.

The following channels were removed from casts 1,3,5,7,9,11,13,15,31,33,35,37,45, 47,49,51,59,61,63,65,77,79,81,83,86,88,90,92,96,98,100,106,118,120,122,124,126,128,131,134,135, 137,139,141,143,145,147,149,151,153,155,157,159,161,163,165: Temperature:Secondary and Salinity:T1:C1.

The following channels were removed from casts 17,19,21,23,25,27,29,39,41,43,53,55,

57,67,71,73,75,94,102,104,108,110,112,114,116. Temperature:Primary and Salinity:T0:C0.

22. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

23. Particulars

In general the primary temperature and conductivity are to be used for CTD #0443 and the secondary values for CTD #0437.

Cast # 69 - The data was bad and this cast was deleted.

Cast # 77 - Though CTD # 0437 was used for this cast the secondary S,T data looks bad near the bottom, so the primary values were used.

Cast # 85 - This was a shallow, messy cast. Delete.

Cast # 86 - The log notes errors in deep rosette data; huge spikes occurred below 200 m. After processing these spikes disappeared.

Cast # 124 - Though CTD #0437 was used for this cast the secondary S,T data looks bad near the surface, so the primary values were used.

Cast #133 - The data above 150 dbars was unreliable. So the upcast , #134, was used.

Casts # 45 and 59 contained irradiance data.

# June 17, 1999 – Joe Linguanti

Cast 17 – changed station from MP04 to MP05.

# Moving files to new archive. May 6, 2003 – Joe Linguanti

- Program HEADER EDIT was used to change non-standard channel names and units and add additional metadata to the CTD and Rosette files.

- Program Clean was used to;

- reset number of records, minimum and maximum values

- delete empty channels where applicable

- set channel name initial letters to uppercase

- Files .CLN were copied to new archive and renamed to the new file convention “YYYY-CC-NNNN.* “.

- The loop file was obtained from Frank Whitney. The latitude and longitude values were computed to 3 decimal places. The spreadsheet file was then converted to IOS Header format with meta-data added to it. The original .XLS file is in the DOC directory.

- The SAIL files were copied from the TRACKOBS directory in the OSAP Data Archive.

- Channels not needed were removed.

- Program HEADER EDIT was  used to change non-standard channel names and units and add additional metadata to the files.

- IOS Header Date and Time channels were added.

- Program Clean was used to set the start and end times and the geographical area.