## REVISION NOTICE TABLE

# PROCESSING NOTES

Cruise: 1997-20, 1997-21, 1997-22

Agency: JOIS (12 US and Canadian agencies)

Project: JOIS

Geographic Area: Hudson Strait, Davis Strait, Jones Sound, Nares Strait

Chief Scientist: Dr. Kelly Kenison Falkner

Platform: CCGS Louis S. St. Laurent

Processed by: Germaine Gatien

Date of Original Processing: 6 March - 31 March 1998

Date of Reprocessing: 18 January 2002 – 22 March 2002

Number of original casts: 167

Number of casts processed:  165 (88 ICTD, 77 MCTD)

# Note that the data from cruises 1997-20, 1997-21 and 1997-22 used the same equipment and were analyzed as a group.

# INSTRUMENT SUMMARY

Two Falmouth Scientific Inc.  CTDs were used:

Instrument Type:  FSI ICTD

Instrument Serial Number: 1329

Instrument Type:  FSI MCTD

Instrument Serial Number: 1534

# FIELD USE (This section was written by H. Melling)

The I-CTD was deployed from the CCGS Louis S St Laurent using the A-frame and hydraulic winch on the boat deck. Profiling speed was well controlled at about 1 m/s. Seawater samples were acquired with the rosette stopped at standard depths on the upcast. The useable CTD profiles were acquired during the downcast of the instrument.

The nominal sampling rate of the I-CTD as configured was 4 Hz (ST = 0). The manual indicates that for this rate of sampling, the sensor outputs are integrals over 105 ms. The actual sampling interval is only approximately related to the tau set-up parameter. The actual sampling rate (3.64 Hz) was slower than the nominal value, presumably because the instrument sampled more channels than the primary data stream (pressure, temperature and conductivity). The instrument recorded data from the following sensors: pressure, platinum thermometer (slow), C-cell, thermistor (fast), touch switch, transmissometer and fluorometer.

# SUMMARY OF QUALITY AND CONCERNS

There are uncertainties in the pressure values for the ICTD.  For the down-cast the pressure gauge seems to take a while to equilibrate with values decreasing as the CTD sits near the surface. Sometimes there are no values with zero conductivity so that it is impossible to determine what the pressure gauge would have read at the surface, and where a value can be determined the pressure may still be reading a little high if equilibration is not complete. The values appear to be about 3db on average with values ranging from 1.2 to 5.6 for 9721. For the up-casts the surface pressure reads 1.1db on average with a spread of 0.1 to 2.6. It is impossible to determine how much of the difference is due to hysteresis and how much to errors in the down-cast values. During 9722 the CTD was lowered to about 30m for many of the casts and then returned to near the surface for a complete cast; for these casts the pressure is likely to be more reliable. A correction of –1.1db was made in calibrating the casts.

There are similar uncertainties in pressure for the MCTD. For the up-casts the surface varies from -4.9db to +0.7db. For the down-casts the surface varies from -1 to -3. A pressure correction of +2.0 dbars was included in calibration.  There remains an error bound of  ±3db.

The temperature and salinity data for the MCTD has been recalibrated to match the ICTD; the quality of this adjustment is limited by the configuration of the deployment for the intercalibrations; the two instruments were not sampling at exactly the same depth nor at the same rates.

The time increment in the raw section of the headers has been corrected in the files to be archived but is incorrect in files from early stages of processing (before the filtering stage.) It should be 0.275s for the ICTD and 0.616s for the MCTD.

## PROCESSING SUMMARY

This data was reprocessed in February/March 2002. The preliminary steps from the original processing were reused.

1.Preliminary Steps (This step was done in March 1998)

The Log Books were not available. Spreadsheets detailing times and positions of casts were obtained.

The salinity bottle data was obtained.

The cruise summary sheet was completed.

The cast numbers were changed to IOS format. Casts using the ICTD used designations such as D3; these were renamed 97**00**.raw. Casts using the MCTD used designations such as M5; these were renamed 97**10**.raw. Thus the fifth position in the cast number  indicates which CTD was used. Consecutive numbers were assigned to each instrument independently, so you can not determine from the numbers when two instruments were deployed on the same cast. There was a problem using the page plot routine in IOSSHELL as it would not print the reference number properly if there was a number other than 0 in the 5th position. The cast numbers were written by hand on those plots. Two raw files named D111NG and M110B were named 97223111 and 97223110, and processed to determine what was in them; they proved to be repeats of other casts and somewhat noisier so were deleted.

For the ICTD the data (*.raw and *.bt1) was converted to IOS Headers in IOSSHELL using “CONVERT FSI RAW FILES” with 1329.cnf as configuration file. There were many errors in headers (Lat and Long  format) which had to be fixed before this conversion was successful.  Comments about data quality found on spreadsheets were entered in the IOS Headers. Station names were entered into the headers after conversion. (The station names are entered correctly in the IOS Headers, but in the header summary and header checks only 4 digits are printed, so the last number or letter is dropped from casts #8 and #36-42 of leg 1.) Header Check and Header Summary were run after calibration since the converted files do not contain depth information.

Before conversion of  MCTD files to IOS HEADERS the first line of many of the *.raw files had to be removed. Many files required further editing before conversion was successful. The first line of data frequently contained only zeros and the last line was often blank; such lines had to be removed.  Corrupted data points were replaced by interpolation. Frequently there was an opening section with an incorrect time, followed by several zero lines and then the data with the correct time began. Such opening sections had to be removed, but no useful data was lost because those sorts of records occurred while the instrument was kept near the surface for equilibration. The configuration file used was 1534.cnf. Station names were added as part of  conversion which was done using ”CONVERT FSI MICRO CTD FILES”.

2. Preliminary Editing and Despiking (This work and all that follows was done in January-March 2002)

The calibrations of P, T and C are done internally by the FSI CTDs.

Conversion to IOS HEADER format was done and record numbers were added using ADD TIME CHANNEL to enable time series plotting.

Time series plots were produced for all casts and were reviewed to guide editing.

ICTD

A text editor was used to remove initial partial downcast records before the full cast as follows:

9720: 26.

9722: 103-110,112-121.

The despike routine in Viewedit was used to remove spikes in the following casts:

9720: 19,25,33,36,38

9721: 51,56,67

9722: 104.105,113,114,119,121,123,125

For most casts the despiking was only applied to the conductivity channel, but for cast 97200019 all channels contained spikes.

MCTD

Before plotting a text editor was used to remove initial partial downcast records and a few records with near-zero conductivity records at the beginning or end of the cast. Reviewing the plots led to further editing of 3 casts. More records were removed from the beginning of cast #101 to eliminate a spike, and the upcast sections of casts #116 and 129 were removed to avoid spikes at the bottom of the casts.

The times shown in the headers for casts #1 through #7 are in UTC in some places and UTC - 7 hours in others (although in all entries they are said to be in UTC). The times were corrected to UTC as required.

VIEWEDIT was used to remove a few spikes mostly at the beginning and end of casts. In running program DELETE (step #5) it was found that there were sudden drops in pressure values too fast to be real; the temperature and conductivity varied smoothly at these points. These casts were revisited in VIEWEDIT and interpolation was used to smooth the pressure.

3. Pressure Study

The end of the upcast section of each ICTD cast was examined to find the first pressure for which the CTD appeared to be at the surface. Where conductivity fell suddenly to low values was judged to be the surface. The average value for the 3 legs was 1.1db and for the individual legs the averages were 0.9, 0.8 and 1.4db. The range was from 0.1 to 2.6db.

For the MCTD casts the average for the three legs were –1.3db with the range from +0.7 to –4.9db. There is an obvious dependence on the maximum pressure of the cast with an average of –0.3db (values of from –1.2db to +0.2db) for casts deeper than 400m. The surface pressure for downcasts is generally about –3.6db with no obvious dependence on air or surface water temperatures. Were the pressure dependence of the upcast surface pressures due to hysteresis then we would expect that the change would increase with maximum pressure. The change does increase until the casts reach a depth of about 600m. But casts deeper than 600m do not show any further increase; if anything there is a slight decrease with deeper casts. There is no obvious temporal variation. Since the error is gradually changing through each cast (from –3.6db to –0.3db for a deep cast) recalibration requires either an extremely complicated scheme or a compromise. It was decided to pick a value that might be considered typical of the mid-depth of the donwcast section of a 200m-400m deep cast. The MCTD pressure will be recalibrated by adding 2db to all values, but it should be noted that the uncertainty in pressure will be ±3db.

4. TIME COMPENSATION

a.) Development of the routine (notes by H. Melling)

Various problems with the design and operating configuration of the FSI CTD’s conspire to yield measured sequences of ocean temperature and conductivity that are asynchronous and have different bandwidths. If these are combined directly in the computation of salinity, the precision of the result is low and there are large systematic errors. This problem results from a serious error in the design of the FSI CTD’s.

An ad hoc (but only partial) solution to this problem was implemented in the processing of these data. The bandwidth of the conductivity channel was reduced by convolving the data series with a impulse response function that had been designed to make the time response characteristic for conductivity the same as that for the slower temperature sensor. The impulse response function, determined empirically via plunge trials, is not a simple analytic function. It is available only as a 10-second sequence of discrete values at intervals of 5 ms (ICTD-ImR.lis, for the ICTD). For use with field data, the function was sub-sampled to the sampling interval of the CTD using the ad hoc programme ICTD-Adj.exe.

Following convolution of the conductivity time series with the sub-sampled impulse response (C-Wgts.lis), the time series of temperature and salinity are optimally matched for the calculation of salinity. The match is not perfect for two reasons: a) the conductivity signal is aliased because the sampling frequency (approximately 4 Hz) is much lower than the Nyqvist frequency (20 Hz) of the signal; b) spatial averaging over the sampling volume for conductivity is ignored; c) differences in fall speed are ignored.

Because of differing sensor-response characteristics, the pressure, temperature and conductivity time series are not synchronous. Moreover, the numerical convolution introduces an additional delay in the conductivity response, which can be calculated from the impulse response for all CTD variables except temperature. The relative timing of conductivity and temperature outputs were determined by trial and error to achieve best results in computing salinity.

A QuickBasic45 programme (ICTDelay.exe) was used to determine the delay in temperature that was optimal for salinity calculation. We use the platinum sensor as the source of temperature data, since the output of the fast-response thermistor drifts badly (> 30 mC) over times measured in tens of minutes. The programme works from the file ICTDImR.lis that contains the empirical impulse response correction sampled at 5 ms intervals. The programme calculates salinity for four different delays in temperature. Delays for temperature in the range of 100-500 ms are suggested. A graphical examination of the salinity series is the easiest route to choice of the ‘best’ delay.

Because of aliasing during sampling, processing with the ‘best’ delay will not in general provide a useable profile of salinity. The computed profile must be smoothed using a running average. Near the surface where vertical gradients in salinity are steep, an average over about 2 seconds will generally produce the desired monotonic increase in salinity (typical of Arctic waters) without flattening the gradient. At greater depth, where the salinity gradient is weaker, the computational noise in salinity dominates the background gradients and a longer averaging period is desirable. Using the output of ICTDelay.exe, various running averages of the salinity time series were examined graphically, to determine the ‘best’ filter length. Averaging intervals may range from 9 points (2.5 s) near the surface to 23 points (6.3 s) at depth.

Data from the M-CTD were adjusted for optimal congruence in response to changing temperature and conductivity using an analogous procedure. Because the two models of CTD have different temperature sensors, the convolution functions for the two instruments are different.

b.) Application of the time compensation routines

The edited files from step 2 were used for the time compensation step.

Programs ICTD-ADJ and MCTD-ADJ were applied to the conductivity to match the time response of the temperature sensors. For the ICTD the sample interval was set to 0.275s and the time response to 0.325s and for the MCTD the values were 0.616s and 0.275s. These programs were prepared by H. Melling.

CLEAN was run to fix the headers.

For the ICTD the IOSSHELL routine REMOVE was used to create two set of files containing:

adjusted Pressure, adjusted Temperature and adjusted Conductivity

adjusted Pressure, adjusted Transmissivity and adjusted Fluorescence

The latter set of file will be left unprocessed and note should be made that the values are nominal, raw values.

For the MCTD, REMOVE was used to create only one set of files containing:

adjusted Pressure, adjusted Temperature and adjusted Conductivity

5.  DELETE

##### ICTD

Before running DELETE the Padj,Tadj,Cadj files output from the TIME COMPENSATION program were edited to remove initial records as follows:

records with zero conductivity, and

the first 36 records with non-zero conductivity.

This editing was not necessary for the casts that had been edited in Step 2. A few bad records were removed from the end of the upcast of cast #32.

Then DELETE was run using the following parameters:

Surface Record Removal: None

Pressure filtered over width:  9 (9x0.275s=2.48s)

Swells deleted. Warning message if pressure difference of 2.00

Delete Slow Drop Rate: Yes

Min Drop Rate (m/sec): 0.5 	Drop Rate Width (samples): 8

Sample interval =  0.275s

Reviewing the DELETE log indicated no problems. The LAST DEPTH in the DEL files was checked against the maximum sampling depth from the header files and found to be reasonably close for all casts. Where bottle records existed the depth of the deepest bottles were found to be less than the LAST DEPTH except for casts #97200042, 97220104, 97220108, 97220112, 97220114 and 97220127.  In those cases the descent rate was very low in the bottom few metres and DELETE operated correctly in removing the data.

##### MCTD

DELETE was run using the following parameters:

Surface Record Removal: None

Pressure filtered over width:  5 (5x0.616s=3.08s)

Swells deleted. Warning message if pressure difference of 2.00

Delete Slow Drop Rate: Yes

Min Drop Rate (m/sec): 0.5 	Drop Rate Width (samples): 5

Sample interval =  0.616s

Reviewing the warnings led to the conclusion that there were problems in the pressure channel for 3 casts (97201005, 97201018 and 97221138); those casts were put through VIEWEDIT again to repair the pressure and then put through steps #4 and 5 again. The only warnings remaining pertain to surface, bottom, upcasts or areas of low descent rate.

6. Calibration of the ICTD files

The instrument was calibrated at Falmouth Scientific on 23 May 1996, and internal coefficients have not been changed since that time. The instrument converts observations to scientific units through use of a calibration table stored within the instrument. These values are scaled for recording as hex integers. The temperature and pressure sensors were calibrated at Scripps (Bon van Hardenberg) in December 1998. At this time, it was noted that the output from the fast thermistor was unstable, drifting as much as 30 mC in 45 minutes. German users of the I-CTD have noted the same problem. This sensor is judged unusable.

(The above paragraph was written by H. Melling.)

Calibration with file 1997new.ccf was used on the ICTD files to convert conductivity to conductivity ratio and correct conductivity for effects of P and T on the dimensions of the C cell, to apply the post-cruise calibration of temperature, to apply a –1.1db offset to the pressure and to calculate salinity.

The BOT files were calibrated using 1997bot.ccf (like the ICTD CCF file except that these files have not been through the time compensation step). Plots were made of the data but no editing was found necessary. The files were bin averaged on bottle number and named RAC.

The header summary was run after calibration to check depths.

7. Initial Calibration of the MCTD files

Calibration with file 1997mctd.ccf was used on the MCTD files to convert conductivity to conductivity ratio and correct conductivity for effects of P and T on the dimensions of the C cell, to apply the post-cruise calibration of temperature, to apply a +2db offset to the pressure and to calculate salinity.

- Filtering of ICTD files

- For the ICTD salinity was run through a box car filter for all casts using the following parameters:

Name of filter size reference channel: Pressure:Adj

Filter sizes:

up to:   0.100000E+03 :    9 (2.58s)

up to:   0.300000E+03 :   13 (3.58s)

up to:   0.500000E+04 :   23 (6.32s)

up to:   0.600000E+04 :   23 (6.32s)

up to:   0.700000E+04 :   23 (6.32s)

HEADER EDIT was run to correct the time increment in the RAW section of the headers to reflect the value found in section 4, namely 0.275s.

Multi-cast T-S plots were produced to compare deep sections of nearby casts as a quality control check. The variations are within the range expected for this instrument. The ICTD data was then analyzed by Humfrey Melling to compare it with that from bottles; he determined that no correction to conductivity was needed. For details see the following files prepared by H. Melling:

•1997-20-22 ICTD conductivity calibration bottle levels.doc.

•ICTD Conductivity Calibration.doc.

- Comparison of MCTD with ICTD and recalibration of the MCTD data

There were 20 casts during which the ICTD AND MCTD were used simultaneously. COMPARE was run using thinned ICTD and MCTD files (9720COM1.xls) for the 11 of those casts which sampled below 200db. Large differences were found in the salinity, but there are also differences in temperature which varied in sign and magnitude. To quantify the temperature differences the maxima or minima of features were compared at a few casts. This again showed that the differences vary in sign. A study was made of the differences at 6 casts in regions of low temperature gradient and at the bottom of another 18 casts. Plotting differences against cast #, pressure, temperature and temperature gradient suggests that it is temperature that is the controlling factor. A trendline was found from the fit of temperature difference vs temperature in 9720com1.xls and used to recalibrate the MCTD temperature (MCTDtemp.ccf); salinity was recalculated.

After this step COMPARE was rerun (9720com2.xls) and show that the MCTD salinity is too high by 0.0665psu using data from below 500db. The salinity was recalibrated using MCTDsal.ccf and COMPARE was run again (9720com3.xls) and the results were found to be satisfactory.

- Filtering of MCTD data

- For the MCTD salinity was run through a box car filter for all casts using the following parameters:

Name of filter size reference channel: Pressure:Adj

Filter sizes:

up to:   0.100000E+03 :    5 (3.08s)

up to:   0.300000E+03 :    7 (4.62s)

up to:   0.500000E+04 :   11 (6.78s)

up to:   0.600000E+04 :   11 (6.78s)

up to:   0.700000E+04 :   11 (6.78s)

HEADER EDIT was run to correct the time increment in the RAW section of the headers to reflect the value found in section 4, namely 0.616s, to add the name of the chief scientists and the ship name to the headers and to add a warning that the pressures are to be considered ±3db.

- Multi-cast T-S plots were produced to compare deep sections of nearby casts as a quality control check. The variations are within the range expected for this instrument.

- Graphical Editing

The page plots were used to decide which casts needed hand editing using CTDEDIT.

##### ICTD

Salinity was cleaned for most ICTD casts in the top 20m. For a few casts records were removed and temperature was cleaned near the surface. Deeper instabilities were left unedited as there was no clear spiking or shed wakes to justify editing. The only casts that were not edited were:

9720: 6,9,11,16,19,29,31

9721: none

9722: 109,111,116,118,124,126

##### MCTD

For the MCTD about half the casts needed editing. The following casts were edited near the surface (above 20m) or at the bottom:

9720: 1-3,5,7,8,11,15,19,20,22,23,25 - 27,29,30.

9721: none

9722: 101,102,103,105-109,113-115,118-122,130,134,139,140.

More extensive editing was done for the following casts:

9720: 10,14.

9721: none

9722: none

Notes of any editing done were made in the individual files.

13.  Final Plots

DERIVED QUANTITIES was run to calculate Theta, Depth, SVA, Gamma and Sound Speed. THIN was then run to obtain values at standard depths. Using edited files and the thinned files final page plots were produced and checked to ensure that the editing was satisfactory.

16. Remove Channels

The following channels were removed from the ICTD and MCTD casts: Conductivity:Adj.

14. Produce Final File

HEADER EDIT was used to add to the headers of all ICTD files:

• the chief scientist’s name

• the ship name

Particulars

ICTD:  	97200001- fluorometer on high sensitivity setting.

- instability near the bottom indicated in T-S plot, also found by MCTD.

File D111NG.raw was renamed 97203111. It has no bottle file and contains down, up and down sections. It is at the same site as 97201111 and was deleted.

97200032 – An error was found in position and time in the header for this cast after processing was underway. It was fixed in all files.

MCTD:	97201013 - no data

File MCTD was renamed 97211050 because it should have been M50 originally.

File M110B was named 97223110 and converted to determine its content. It contained repeated up and down casts apparently at the same site as 97221110 and was deleted.

97201030 had a data line that was corrupted. The temperature and pressure were missing and were replaced by interpolated values.

97221108 had a data line that was corrupted. The pressure was replaced by an interpolated value.

97221118 is missing the first 1000 metres of down-cast.

- RECOMMENDATIONS

When using the FSI CTDs it is important to choose a consistent way of writing the latitude and longitude in the header. The IOSSHELL routine for converting to IOS headers could be edited to reflect whatever choice is made. Similarly, if a decision is made to enter the station identifier in a field (ex. Operator or Direct) the conversion program could be used to read this into the IOS headers. These changes would make the processing of the data more efficient.

Establishing the surface pressure appears to require that the CTD be held in water below the surface for a few minutes and then be raised until a conductivity of value 0 is achieved. The pressure gauge seems to require some time to equilibrate.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	9720, 9721, 9722

Dates: 	Start:	August 1, 1997			    End:	20 August, 1997

Location:	Hudson Strait, Davis Strait, Jones Sound, Nares Strait

Vessel:		CCGS Louis S. St. Laurent

Chief Scientist:	Dr. Kelly Kenison Falkner

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | FSI | ICTD | 1329 | yes |  |
| 2 | FSI | MCTD | 1534 |  |  |

NOTE: Record Sensor Calibration using the form “CTD Calibration Information”, and attach.

Salinity (analyst:  Leah Foechterle)

Samples run after equilibrating to room temperature for at least 12 hours using a model 8400A Autosal.  Precision is approximately + 0.003%

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	FSI/ICTD/1329

Cruise ID#:	9720-9721-9722

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure |  | 21 May 1996 | * factory |  |  |
| Temperature |  | “ | “ | Jan 1998 | **** IOS |
| Conductivity |  | “ | “ |  |  |
| Transmissometer | 598 | 13 Feb. 1996 | lab - Bernard Minkley |  |  |
| Fluorometer | 93S |  |  |  |  |

* The FSI ICTD was calibrated at the factory on 21 May 1996. The calibration coefficients are stored internally and the output of T, C and P is calibrated.

**** A post-cruise temperature calibration was done for the ICTD by Bon van Hardenberg at IOS as described at the end of the next section.

| Calibration Coefficients |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Channel | Formula | Coefficients |  |  |  |  |  |
| Name | # | C1 | C2 | C3 | C4 | C5 | C6 |
| Temperature - post cruise | 10 | - 0.0241006 | 1.001510 |  |  |  |  |
| Transmissometer | 10 | -0.08 | 1.008654 |  |  |  |  |

Sensor Calibration Notes:

The log incorrectly states that the transmissometer was # 234.

The transmissivity output from the ICTD is a 14-bit number which has been converted from the analog signal. The raw values have not been converted to %TR. (This can be done by dividing by 16383 (2**14 - 1) and multiplying by 100. The results imply that the 25cm path length has been accounted for in the conversion.)

The transmissometer coefficents were derived from the measurements made by Bernard Minkley at IOS as follows:

Trans (actual) = 20* Voltage(actual) = 20* [ A/B *(Voltage measured-Blank Voltage)]

= A/B * Trans (measured) - 0.08

= 1.008654 * Trans (measured) - 0.08

where A = Air Voltage measured at the factory (4.6620); B= Lab calibration (4.6220); Blank voltage = 0.0040.

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	FSI/MCTD/1534

Cruise ID#:	9720-9721-9722

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure | 1534 | May 1996 | * factory |  |  |
| Temperature | “ | “ | “ | Jan 1998 | **** IOS |
| Conductivity | “ | “ | “ |  |  |

| Calibration Coefficients |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Channel | Formula | Coefficients |  |  |  |  |  |
| Name | # | C1 | C2 | C3 | C4 | C5 | C6 |
| Temperature - post cruise | 10 | - 0.0241006 | 0.999785 |  |  |  |  |

* The FSI MCTD was calibrated at the factory in May 1996. The calibration coefficients are stored internally and the output of T, C and P is calibrated.

* A post-cruise temperature calibration was done for the MCTD by Bon van Hardenberg.