## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 	97-25

Date:	May 27 - 30, 1997

Location: Strait of Georgia

Project:	Georgia Strait

Chief Scientist: Ken Denman

Platform: Vector

Processed by: Joseph Linguanti

Date of Processing :22 August 1997 - 16 September 1997

Number of original casts: 70

Number of casts processed:  34

INSTRUMENT SUMMARY

Two SeaBird Model SBE 911+ CTDs were used: #0443 AND #0437. CTD 0443 was used in conjunction with the 24 bottle rosette and CTD 0437 was used alone.

SUMMARY OF QUALITY AND CONCERNS

The primary and secondary conductivity sensors are very different for CTD #0443.  The primary sensor produced salinities approximately 0.06 PSU higher than the secondary sensor. Comparison with bottle data showed that the primary sensor was more accurate, however both  differed considerably from the bottle values. The primary sensor produced salinities about 0.0278 PSU higher and the secondary sensor produced salinities about 0.04 PSU lower. The primary conductivity sensor was used in processing salinity. Doug Anderson was notified of these problems.

The temperature sensors for both instruments are very consistent.

There were no calibration samples taken for CTD 0437. For the June 1997 Line-P cruise (9711) Germain Gatien noted almost the exact same problems encountered for this data set. Since she was able to compare 0437 salinities with bottle salinities I decided to use the corrections calculated for cruise 97-11.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained and  note was made of problems that occurred during the cruise.

The salinity bottle data sheets were found.

The cruise summary sheet was completed.

3. Conversion of Raw Data

This step was completed by Bernard Minkley.

4. Preparing  a set of data for further processing

Files *.cnv were copied into *.org; this set of files was left unprocessed.

5. Test Plots

To check the agreement of the pairs of temperature and conductivity sensors plots were made of a selection of casts for each of the instruments used.

a.) For CTD #0443 all casts were checked . At depth temperature variations were <0.001 C.deg. Conductivity variations resulted in salinity variations on the order of 0.06 PSU. Both salinity channels from CTD cast # 17 were compared with the deep (>200 db) bottle values from station NI03. The primary salinity was about 0.022 PSU higher and the secondary salinity was about 0.05 PSU lower. Since the primary was more accurate, it was used in the processing of salinity.

b.) For CTD #0437 all casts were examined. At depth temperature variations were <0.001, conductivity <0.001 units and salinity <0.008 PSU. Primary sensors were used for processing.

6. WILDEDIT

Program WILDEDIT was used to remove spikes in T, C, S and Tr. This stage of processing is to flag very bad points. Parameters used were:	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

7. CELLTM

The conductivity cell thermal mass correction was done using alpha = 0.03 and 1/beta = 7.0.

8. FILTER

The pressure was low-pass filtered with a time constant of 0.15 seconds.

9. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

10. Checking Headers

Most casts were OK. For casts 5 and 6 the station id was changed from D8A to D7A.

11. DELETE

The following DELETE parameters were used:

Surface Record Removal: Low Salinity and Last Press Min

Maximum Surface Pressure (relative):      20.00

Minimum Surface salinity:    0.0

Surface Swell Pressure Tolerance:        .50

Pressure filtered over width:   11

Swells deleted. Warning message if pressure difference of      2.00

Drop rates<    .2m/s (calculated over  11 points) will be deleted

Sample interval =       .04 seconds.

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

13. Test Plots

Page plots were produced for all casts from the meter averaged files. Data appeared to be very clean with the exception of cast 60 which showed unstable water in the upper 20 db. Plots of the secondary channels did not show the instability therefore the secondary channels were used for this cast.

14. CTDEDIT

All casts were OK.

15. Preparing files for intercomparison

The *.ros files were converted to IOS header files and renamed as *.bot. A cast number pairs file and a pairs list were prepared and used to add bottle sample numbers to the *.bot files producing the *.sam files.

16. COMPARE

Bottle salinities were taken only with CTD 0433. The COMPARE routine was used to produce a file 0433.CSV which was imported into EXCEL using the COMPARE macros. The salinity differences were plotted against pressure excluding shallow data. The CTD salinity data was  about 0.0288 PSU higher than the bottle values. The following correction was applied to salinity for casts 1 to 45,

new salinity = old salinity - 0.0288

There were no bottle casts with CTD # 0437 and so COMPARE could not be used to correct salinity. An extensive comparison was done by Germain Gatien for the June cruise 97-11. Since cruise 97-11 immediately followed this one, it was decided to use the corrections determined for cruise 97-11. The following corrections were applied to salinity for casts 46 - 71.

Casts 46-59 and casts 61-71,  new salinity = old salinity + 0.023

Casts 60,                                new salinity = old salinity + 0.018  Secondary salinity channel used.

17. Final Plots

Page plots were prepared using the corrected data.

18. REMOVE

The following channels were removed from all casts: Scan_Number, Conductivity:Primary, Conductivity:Secondary and Flag.

The following channels were removed from casts 1 - 59, 61-71: Temperature:Secondary and Salinity:T1:C1.

The following channels were removed from cast 60. Temperature:Primary and Salinity:T0:C0.

19. Producing final files

a.) The final files were renamed *.ctd.

20. Particulars

In general the primary temperature and conductivity are to be used for CTD #0443 and the secondary values for CTD #0437. The only exception was cast 60 where the secondary channels were used since they did not show any instability.

Casts 3,5  - Transmissivity was bad so it was deleted.

Cast  60    - See note at beginning of this section and section 13.

Data archived on tapes DLT023 and DLT024 in saveset 9725_all.bak