PROCESSING NOTES

Cruise: 97-38

Agency: OSAP

Project: Inlets, Strait of Georgia

Chief Scientist: Dario Stucchi

Platform: Vector

Geographic Area: WCVI/Strait of Georgia

Date: 6 October 1997 - 12 October 1997

Processed by: Joseph Linguanti

Date of Processing : 15 February 1998 - 7 March 1998

Number of original CTD casts: 95

Number of casts processed: 55

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0437) was used with a rosette. A transmissometer was attached.

SUMMARY OF QUALITY AND CONCERNS

All CTD casts were in good order with the exception of cast 3 that had untrustworthy transmissometer data. Comparisons between the primary and secondary sensors were consistent with what was noted for cruise 97-37 (see notes under ‘Summary of quality and Concerns’ for cruise 9737). The only anomaly observed was for cast 21 (an up-cast) where the conductivity difference ranged from 0.08 to 0.2 S/m. This anomaly is not present on any of the subsequent casts. The logbook notes that bottom was hit on this cast and the large conductivity difference is probably due to some debris around the sensors.

Comparison with bottle salinities found the primary salinity to be low by approximately 0.025 PSU. Again this was consistent with the previous cruise, 97-42 and the later cruise, 97-37. The primary channels of temperature and conductivity were used in processing.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Logbook was obtained and note was made of problems that occurred during the cruise.

The salinity bottle data sheets were found.

The cruise summary sheet was completed.

3. Conversion of Raw Data

Conversion of the raw data was completed by Bernard Minkley using conversion files ‘97380001.con’.

4. Preparing  a set of data for further processing

Files *.cnv were copied into sub-directory ORG; this set of files was left unprocessed. All processing was done in directory PROCESS and COMPARE.

5. ALIGNCTD

The secondary conductivity channel was advanced by 0.073. (The primary channel is aligned at acquisition by the deck unit).

6. Test Plots and Channel Check

A sample of casts (#1,3,4,9,20,30,40,50,60,70,80,90) was plotted to check for agreement between the pairs of  T and C sensors. Differences were consistent with what was found for cruise 97-37.

7. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

8. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03 and 1/beta = 7.0.)

9. STRIP

It was necessary to run this routine to remove the salinity channels, so that they could be recomputed to take into account the editing routines (steps 5, 7 and 8).

10. DERIVE

Program DERIVE was run on all casts to calculate secondary salinity and descent rate.

11. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

12. Checking Headers

A cruise plot, header summary and header check were produced. No errors were found.

The surface check was run. The mean surface pressure = .808. No correction was made.

.

13. DELETE

The following DELETE parameters were used:

Surface Record Removal:  Last Press Min

Maximum Surface Pressure (relative):      20.00

Surface Swell Pressure Tolerance:        .50

Pressure filtered over width:   11

Swells deleted. Warning message if pressure difference of      2.00

14. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

15. Test Plots

No such problems were found. Page plots were produced using T0, S0 and Tr for all casts. These plots were examined for spikes and instabilities.

16. CTDEDIT

Minor editing was done on 3 casts, 20, 32 and 85.

.

17. Preparing files for intercomparison

Bottle salinities were taken during 46 casts. The *.ros files were converted to IOS header files and renamed as *.bot. A cast number pairs file and a pairs list were prepared and used to add bottle sample numbers to the *.bot files producing the *.sam files.

18. COMPARE

The COMPARE routine was used to produce a file 0437.CSV that was imported into EXCEL using the COMPARE macros. The CTD salinities are about 0.025 PSU low. A recalibration file (9738CORR.CCF) was prepared in COMPARE and applied to all casts. The *.sam files were also recalibrated and COMPARE run again using those new files (*.cor) to check that the correction was satisfactory. This second run of COMPARE gave results mostly within +/- 0.003 PSU.

For all files 0.02546 PSU was added to the primary salinity

19. Final Plots

Page plots were prepared using the corrected data.

20. REMOVE

The following channels were removed from all casts: Secondary Temperature, Secondary Salinity, Scan_Number, Conductivity:Primary, Conductivity:Secondary Descent Rate and Flag.

21. Producing final files

The final files were renamed *.ctd.

24. Particulars

Cast 3 – 	Transmissivity removed because it was untrustworthy.

Casts 87, 94 –	Inversions look suspicious. Both pairs of sensors show the same inversions, therefore it was decided to leave as is because no other explanation available.

Data archived on tapes DLT046 and DLT047 in saveset 9738_ALL.BAK

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:		97-38

Dates: 	Start:		6 October 1997		    End:	12 October 1997

Location:	West Coast of Vancouver Island, Strait of Georgia

Vessel:		Vector

Chief Scientist: 	Dario Stucchi

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SeaBird | 911+ | 0437 | yes | yes |

Notes:  A transmissometer, PAR sensor and fluorometer were attached to the CTD #0437.

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SeaBird/911+/ 0437

Cruise ID#:		97-38

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure | 63502 | 1-Nov-1996 | Factory |  |  |
| Temperature | 2106 | 29-May-1997 | Factory |  |  |
| Conductivity | 1748 | 30-May-1997 | Factory |  |  |
| Secondary Cond. | 1764 | 6-Jun-1997 | Factory |  |  |
| Secondary Temp. | 2062 | 29-May-1997 | Factory |  |  |
| Transmissometer | 192D | 31-Oct-1996 | Factory |  |  |

Sensor Calibration Notes:

Listing of Configuration files used is attached which include the temperature, conductivity and pressure sensor calibrations.