PROCESSING NOTES

Cruise: 9808

Agency: IOS/UBC

Project: COPRA

Chief Scientist: Steve Romaine

Platform: CCGS Vector

Date: 30 March 1998 – 31 March 1998

Processed by: Germaine Gatien

Date of Processing : 14 May 1999 – 15 May 1999

Number of original CTD casts: 5

Number of casts processed: 5

INSTRUMENT SUMMARY

A SeaBird Model SBE 19 SEACAT CTD  (S/N#197957 - 1294) was used.

SUMMARY OF QUALITY AND CONCERNS

The data is of good quality and required only slight editing.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

The Log sheets were obtained; there was no mention of any problems with the CTD.

There were no salinity bottles.

The cruise summary sheet was completed.

- Conversion of Raw Data

- The raw data was converted using conversion file 9808ctd.con. The calibrations used for temperature and conductivity were from a post-cruise calibration done April 29, 1998. The pre-cruise calibration was done 12 February 1993. The drift over that 5-year period was 0.0003psu/month and 0.00001C°/year.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature. The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

5. ALIGNCTD

A test run was done on one cast comparing the results of advancing the temperature channel by 0.4, 0.5 and 0.6s relative to pressure. All three improved the alignment of T and C with no notable difference between the three choices so the value suggested in the manual, 0.5s, was used for all casts. .

6. DERIVE

Program DERIVE was run to calculate salinity.

7. Test Plots

The casts were plotted on-screen; up- and down-casts were in reasonable agreement

8.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files (911+) was used to convert the SEA-Bird data to IOS Headers.

9. Checking Headers

A header summary was produced and no errors found. There was a disagreement between the log notes and the header as to the station name for cast 104, but the log appears to be in error.

The surface check was run. The average surface pressure is –0.5dbars. No adjustment is necessary.

The header check was produced.

The cruise track was plotted and looks reasonable.

11. DELETE

Because the drop rate was fairly low the minimum was chosen to be 0.1m/s to minimize data loss. For the most part the drop rate was steady and the only warning was for cast 100 near the bottom when the CTD reversed direction briefly during the down-cast.

The following DELETE parameters were used:

Surface Record Removal:  Last Press Min

Maximum Surface Pressure (relative):      20.00

Surface Swell Pressure Tolerance:        .50

Pressure filtered over width:   11

Swells deleted. Warning message if pressure difference of      2.00

Drop rates<    .10 m/s (calculated over 4 points) was deleted.

Sample interval =       .04 seconds.

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

13. Test Plots

Profiles of the files before and after DELETE were produced.

Page plots were produced and were examined for spikes and instabilities to guide the use of CTDEDIT.

14. CTDEDIT

CTDEDIT was used to clean noise in S and T at the surface in the following casts: 100,101,102 and 104.  Note was made of the editing details in the relevant files

- Intercomparisons

There were no bottle casts for this cruise nor for 9748 in December 1998 when the same CTD was used.

Repeat casts – There were no repeat casts.

Historic ranges -  Casts #103 and 104 were plotted with the ranges superimposed; all data fell within the ranges, but the latter are very wide in this area and depths.

18. Final Plots

Page plots were prepared using the edited data.

19. REMOVE

The following channels were removed from all casts: Scan_Number, Conductivity and Flag.

20. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

Cast #:  100 – Identified in the log as cast 1

101 – Identified in the log as cast 2

102 – Identified in the log as cast 3

103 – Identified in the log as cast 4

104 – Identified in the log as cast 5; station name given as BB03 in log, but BB01 in header. There is a correction to station names on the previous page that is assumed to apply to this cast as well. Other activities at the same lat/long are identified as occurring at BB01 elsewhere in the log.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	9808

Dates: 	Start:	30 March 1998	    End:	31 March 1998

Location:	Strait of Georgia

Vessel:	CCGS VECTOR

Chief Scientist:		Steve Romaine

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19 SEACAT | 1294 | No | Yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/19 SEACAT / 1294

Cruise ID#:		9808

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 1294 | 12 Feb 1993 | Factory | 29 Apr 1998 | factory |
| Conductivity | 1294 | 12 Feb 1993 | Factory | 29 Apr 1998 | factory |
| Pressure Sensor | 163223 | ? | ? | 7 May 1998 | factory |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.