## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 30-Mar-2017 | Renamed fluorometer channel name from "Fluorescence:URU" to "Fluorescence:Calibrated" since the data from the fluorometer (on the CTD) has been converted to Chlorophyll-a using linear regression on observed extracted chl-a. Joe Linguanti |
| 12-Jan-2004 | Move files to new archive. See note below. |
| 15-Jun-2000 | Cast 143 – corrected start time. J.L. |
| 23-Feb-2000 | Doug Yelland applied calibrations to the Chlorophyll fluorescence channel. Original CTD files were copied to the WORK directory. J.L. |
| 17-Jun-1999 | Cast 17 – changed station name from MP07 to MP06. J.L. |
| 15-Apr-1999 | Reprocessed cast 39 since the first 15 metres was missing from the file. The 15-minute wait at the surface must have confused the DELETE program. I deleted the first 12,000 records and then ran program DELETE. J.L. |

PROCESSING NOTES

Cruise: 9810

Agency: IOS

Project: La Perouse, GLOBEC

Party Chief: Dave Mackas

Platform: J.P.Tully

Date:11 May 1998 - 25 May 1998

Processed by: Germaine Gatien

Date of Processing : 18 June 1998 - 25 August

Number of original CTD casts: 84

Number of casts processed:  81

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was used. A second CTD (#0437) was on board.

A PAR sensor, SeaTech fluorometer and transmissometer were on board.

SUMMARY OF QUALITY AND CONCERNS

•The secondary conductivity sensor (#1766) did not perform well. The same sensor was used during cruise 9809 and the differences at 1000dbars between this sensor and sensor #1763 varied from more than 0.003 at the beginning of 9809 to less than 0.002 units by the end of 9810.

•The primary conductivity sensor behaved poorly during cast #64 so corrected secondary values were used for that cast.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained and  note was made of problems that occurred during the cruise.

The salinity data was obtained

The cruise summary sheet was completed.

- Conversion of Raw Data

The raw data was converted using conversion file 9810flo.con.

4. Preparing  a set of data for further processing

Files *.cnv and *.ros were copied into sub-directory CONVERT; this set of files was left unprocessed. All casts contain both up and down traces. Checks of a few casts show all expected channels present and suggest that the difference between primary and secondary conductivity is about 0.002 units, somewhat higher than expected.

- ALIGNCTD

The secondary conductivity channel was advanced by 0.073. (The primary channel was aligned at sea by the deck unit). A test cast shows that the two salinity channels track better after this step.

6. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

7. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03 and 1/beta = 7.0.)

8. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

9. Test Plots and Channel Check

A sample of casts (#6, 64,113,152 and 169) was plotted to check for agreement between the pairs of  T and C sensors. The temperature differences are ~0.0015 C° around 1000 dbars while the conductivity  differences varied from ~0.0022 units at the beginning to ~0.0019 units near the end;  the salinity varied from ~0.025psu at the beginning to ~0.022psu near the end. These results are similar to those found during cruise 9809 which used the same sensors. A cursory check of bottle data suggests that the primary sensors were the more accurate. However, for cast #64 there was a noisy section for the primary C near 750dbars.

10. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

11. Checking Headers

A header summary was produced and no errors found.

The surface check was run. The average surface pressure is 1.929dbars.

The header check was produced.

The cruise track was plotted.

12. DELETE

The following DELETE parameters were used:

Surface Record Removal:  Last Press Min

Maximum Surface Pressure (relative):      20.00

Surface Swell Pressure Tolerance:        .50

Pressure filtered over width:   11

Swells deleted. Warning message if pressure difference of      2.00

Drop rates<    .33 m/s (calculated over  11 points) was deleted.

Sample interval =       .04 seconds.

Due to warning messages the following casts were rerun with Slow Drop Rate NOT Deleted: #39,44,180.

Bad records had to be deleted from the beginning of cast #143, and DELETE rerun.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

14. Test Plots

Profiles of all casts were examined for large inconsistencies between primary and secondary values and for any evidence of problems with the processing.

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

15. CTDEDIT

CTDEDIT was used to clean noise in S and T near the surface in the following casts:18,33,92,105.

CTDEDIT was used to clean noise in S only near the surface in the following casts: 6,21,45,94, 114,117,141,143,157,159,187,191.

CTDEDIT was used for more extensive cleaning in the following casts: 5,25,39,48,49,64,112,183.

Note was made of the editing details in the relevant files.

16. COMPARISON WITH BOTTLES

There were bottle samples taken at 51 stations. COMPARE was run using file 9810comp.csv. The CTD salinities were higher than the bottle salinities with differences ranging from +0.0014 to +0.0019 at 1000dbars depending on the choice of what data to include. The upper end of the range looks like the best choice; this value is close to the results for cruise 9809 (+0.002psu at 1000dbars) which used the same sensors. It was decided to correct both sets of data by applying a linear offset of -0.002psu. COMPARE was rerun after applying the correction, and the differences between CTD and bottles is now <0.003 for most values below 200dbars.

17. Other intercomparisons

Repeat casts -  There were no repeat casts.

Historic ranges - About 12 casts were checked against the historical records and found to be within the normal range, including the 7 casts along Line P.

Sensor history -	All four sensors were used during cruise 9809 during which #1766 was found to be low by an amount that varied through the cruise from >0.003 units at the beginning to <0.002units at the end.

T:2095 - used in 9813 - differed from 2142 by 0.002 C°

2142 - used in 9813 -       “         “   2095 “       “

C:1763 - used during 9802 - no problems noted; diff from other C sensor by 0.0002.

1766 - used during 9707 - only used for 2 casts; diff from other C sensor by 0.001.

and during 9713 -  diff from other C sensor by 0.0005 units

18. Final Plots

Page plots were prepared using the edited data.(The salinity is uncorrected in the page plots).

19. REMOVE

The following channels were removed from all casts except cast 64: Temperature:Secondary, Salinity:Secondary, Scan_Number, Conductivity:Primary, Conductivity:Secondary and Flag.

The following channels were removed from cast 64: Temperature:Primary, Salinity:Primary, Scan_Number, Conductivity:Primary, Conductivity:Secondary and Flag.

20. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

Cast #1 - on deck - deleted

Casts #5,6,13 - use con file #14 to convert.

Cast #14 - on deck - deleted

Cast #17 - air pressure lost, stopped around 500m, then continued to 1000m. No obvious problem in data.

Cast #39 - sat at surface for 15 minutes because of a dry cell.

Cast #62 - DELETE removed no records because the pressure was less than 20.

Cast #64 - primary temperature and conductivity look strange around 750dbars. Secondary values were used for this cast after recalibration of the salinity.

Cast #65 - log note: PAR valve using DEV calibration - needs rerunning. Data looks o.k.

Cast #112 - DELETE removed no records because the pressure was less than 20.

Cast #116 - large wire angles

Cast #143 - bad header file - fixed before conversion to IOS Headers. The first 48 records were removed from the converted file (using a text editor) because the Descent Rate had overflowed.

Cast #163 - contains only one line of data  - Deleted

Cast #168  - Log mentions data errors at 320db; no errors obvious in converted data.

Cast #188 - heavy rain, large swell.

The following rosette files were deleted because there were no salinity bottles from the casts: 5,15,21,62,82,112,146,152,163.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:		9810

Dates: 	Start:		11 May 1998	    End:	25 May 1998

Location:	West Coast Vancouver Island

Vessel:		J.P.Tully

Chief Scientist:		Dave Mackas

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SeaBird | 911+ | 0443 | yes | yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SeaBird/911+/0443

Cruise ID#:		9810

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 26/09/96 | factory |  |  |
| Conductivity | 1763 | 30/09/97 | factory |  |  |
| Secondary Temp. | 2142 | 23/07/96 | factory |  |  |
| Secondary Cond. | 1766 | 23/07/96 | factory |  |  |
| Transmissometer | 182D | 11/13/96 | factory |  |  |
| Pressure Sensor |  |  |  |  |  |

Irradiance (PAR) - S/N 4495

fluorometer - SeaTech

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.

January 12, 2003 – Moving files to new archive.

The CHE files from the old archive did not contain all of the bottle salinity data. Melanie must of omitted the data somehow when she merged the nutrient data. I obtained the original rosettes files with the oxygen and bottle salinity data merged (HYD files) from Bernard’s saved disk and merged the nutrient data from the CHE files. The corrections applied to salinity and fluorescence data in the CTD files were applied to the salinity and fluorescence data in the new CHE files.