PROCESSING NOTES

Cruise:  9811

Agency: IOS

Project: La Perouse, El Niño, Georgia Strait

Party Chief: Tom Juhász

Platform: J.P.Tully

Date: 25 May 1998 - 31 May 1998

Processed by: Germaine Gatien

Date of Processing : 18 June 1998 - 26 August 1998

Number of original CTD casts: 23

Number of casts processed:  23

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0443) was used. A second CTD (0437) was on board.

SUMMARY OF QUALITY AND CONCERNS

Transmissivity channels from casts 1 and 2 contained only zero values.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained and  note was made of problems that occurred during the cruise.

The salinity data was obtained and the cruise summary sheet was completed.

- Conversion of Raw Data

The raw data was converted using conversion file 9811new1.con

4. Preparing  a set of data for further processing

Files *.cnv were copied into sub-directory CONVERT; this set of files was left unprocessed. All casts contain both up and down traces and all expected channels are present. The differences between T and C sensors are ~0.002C° and ~0.001units respectively.

- ALIGNCTD

The secondary conductivity channel was advanced by 0.073. (The primary channel was aligned at sea by the deck unit). A test of one cast shows that the two salinity channels track better after this step.

6. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

7. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03 and 1/beta = 7.0.)

8. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

9. Test Plots and Channel Check

A sample of casts (#3,15,22) was plotted to check for agreement between the pairs of  T and C sensors. The temperature variations at cast 3 are <0.0005 C° around  1000 dbars while the conductivity differences <0.0005 units. Salinity values varied by about 0.006psu. The salinity differences get slightly larger with depth; this fits observations made by Ron Perkin during cruise 9815 which used the same sensors. Ron applied a pressure correction to the conductivity channels using file 9815corr.ccf. The other casts checked were quite shallow and showed variations consistent with the cast #3 values.

The secondary salinity is noisy in casts #1,8,10 and 11. The primary salinity is spiky near the surface in cast #4.  A quick check of bottle salinities suggests that the values lie between the primary and secondary salinities. Since the secondary traces were sometimes noisy, the primary values were chosen and a correction will be applied to the primary conductivity later in the processing.

The transmissivity channels are present for casts 1 and 2 but contain only zero values.

10. Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

11. Checking Headers

A header summary was produced and no errors found.

The surface check was run. The average surface pressure is  1.578 dbars.

The header check was produced.

The cruise track was plotted.

12. DELETE

The following DELETE parameters were used:

Surface Record Removal:  Last Press Min

Maximum Surface Pressure (relative):      20.00

Surface Swell Pressure Tolerance:        .50

Pressure filtered over width:   11

Swells deleted. Warning message if pressure difference of      2.00

Drop rates<    .33 m/s (calculated over  11 points) was deleted.

Sample interval =       .04 seconds.

Due to a warning message cast # 13 was rerun with the SLOW DROP RATE feature turned off.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

- RECALIBRATION

The recalibration used by Ron Perkin in 9815 to correct for a pressure effect on the conductivity was applied to the metre-averaged files and the rosette files. (9811corr.ccf)

15. Test Plots

Profiles of all casts were examined (on screen) for large inconsistencies between primary and secondary values and no evidence was found of problems with the processing.

Page plots were produced using T0, S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

16. CTDEDIT

CTDEDIT was used to clean noise in S and T near the surface in the following casts:13,22

CTDEDIT was used to clean noise in S near the surface in the following casts:6,9,11,12, 15,19

CTDEDIT was used to clean noise in S at the bottom in the following casts: 1

Note was made of the editing details in the relevant files.

- COMPARE

Salinity bottles were taken at casts 1,2,3,5,15,17,19,21,22. COMPARE was run and shows that the primary salinity is about 0.0064psu higher than the bottle salinities. The results are in 9811comp.xls.

Recalibration of the data using this offset was done using file 9811comp.ccf. Running COMPARE again showed most values below 500dbars within 0.002psu of the bottle values. Results of the second comparison are in 9811rcal.xls.

- Other intercomparisons

Repeat casts - There were no repeat casts

Historic ranges - Casts 1,2,3,4,5,22 and 23 were checked and had T and S within the historic ranges.

Sensor history - All four sensors were recalibrated just prior to the cruise. They were subsequently used on 9815.

19. Final Plots

Page plots were prepared using the edited data.

20. REMOVE

The following channels were removed from all casts: Temperature:Secondary, Salinity:Secondary, Scan_Number, Conductivity:Primary, Conductivity:Secondary and Flag.

Transmissivity was removed from the following casts: 1 and 2.

21. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

Cast #1- 3 - used con file 9811new later found to be wrong and replaced by 9811new1.

Cast 1 - zero transmissivity values The transmissivity channel was deleted in REMOVE step.

Cast 2 - zero transmissivity values. The transmissivity channel was deleted in REMOVE step.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:		9811

Dates: 	Start:		25 May 1998	    End:		31 May 1998

Location:		West Coast Vancouver Island

Vessel:		J.P.Tully

Chief Scientist:		Tom Juhász

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SeaBird | 911+ | 0443 | yes | yes |

A second CTD system (#0437) was on board but was not used.

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SeaBird/911+/0443

Cruise ID#:	9811

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 01/05/98 | factory |  |  |
| Conductivity | 1729 | 01/05/98 | factory |  |  |
| Secondary Temp. | 2038 | 07/05/98 | factory |  |  |
| Secondary Cond. | 1762 | 01/05/98 | factory |  |  |
| transmissometer | 182D | 11/13/96 | factory rebuild |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.