## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 10-Jan-2006 | Surface loop data was added to the archive. The data was acquired from John Morris at PBS. The original spreadsheet file from John and more detailed processing notes can be found in the “Cruise_Data\Documents” directory. Any questions regarding this data should be directed to John Morris. J.L. |

# PROCESSING NOTES

Cruise: 9819

Agency: PBS

Project: High Seas Salmon

Geographic Area: N.E. Pacific, Gulf of Alaska

Scientific Party Chief: John Morris

Platform: W. E. Ricker

Date of Cruise: 5 October 1998 - 22 October 1998

Processed by: Germaine Gatien

Date of Processing :29 January 1999 - 8 February 1999

Number of original casts: 142

Number of casts processed:  70

# INSTRUMENT SUMMARY

The CTD used during this cruise was:.

Instrument Type:  Guildline 8715 CTD

Instrument Serial Number: 43825

# SUMMARY OF QUALITY AND CONCERNS

During most of this cruise the CTD pressure range was set incorrectly so that the casts were only half as deep as planned. After cast #126 this was corrected.

# PROCESSING SUMMARY

- PRELIMINARY STEPS

The Log Book was obtained and note was made of problems that occurred during the cruise.

The cruise summary sheet was completed.

The header summary and header check were examined. There were a number of cases where the date in the log is one day behind that of the header; the headers appear to be correct and were left unchanged. It is assumed that this error was due to using the date in local time, not UTC. The times all appear to be correct.

The surface check gave an average surface pressure of 2.25db, but the variation in values was large. Examining a few casts in detail suggests that the first values are not right at the surface. The CTD was held near the surface for a while and the shallowest good salinity values appear to be close to 0db. No correction was made.

There were separate up and down casts. Only down-casts were processed.

- CONVERSION

- The data was converted using the IOS SHELL program ACQCONV.

- CALIBRATION

There are pre-cruise (9819cal.ccf) and post-cruise calibrations (9819post.ccf) available for this CTD  done in July 1998 and November 1998, respectively. Casts #70, 94 and 122, for which there are deep salinity samples, were calibrated using the two sets, and a cursory examination led to the choice of the post-cruise calibrations as being slightly better.

The data was calibrated using file 9819post.ccf which is attached.

4. DESPIKE

- The following DESPIKE parameters were used:

- CHANNEL        FIT   OVER-  MIN       MAX       MIN         MAX         SPIKE

- WIDTH LAP    VALUE   VALUE  STDDEV  STDDEV     TOL.

- -------------------- ----- ----- ------- ------- ------- ------- -----

- Pres 		55     5        -2.00       2000.00	.10000        999.000    3.00

- Temp   	55     5         0.00           15.00	.00500        999.000    3.00

- Cond    	55     5           .00            1.50	.00010        999.000    3.00

5. TIME COMPENSATION

- Temp. probe Dist (m):   .0700  Sample period (sec):   .0400

6. DELETE

The following DELETE parameters were used:

Surface Record Removal: Low Sal & Last Press Min

Maximum Surface Pressure (relative):      20.00

Minimum Surface Salinity:     10.000

Surface Swell Pressure Tolerance:       1.00

Pressure filtered over width:   11

Swells deleted. Warning message if pressure difference of      2.00

Drop rates<    .30m/s (calculated over  11 points) will be deleted

Sample interval =       .04 seconds.

Due to warnings casts #34,70,80,88,94 and 126 were rerun with the Low Drop Rate feature turned off. There were warnings for casts #22 and 24 that no records were skipped because the pressure range was less than 20.

7. FILTER

- A test run was done (cast #10) to see if filtering improved the date. It removed some spikes so a median box car filter with filter width 21 was applied to the salinity channel for all casts.

8. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

9. HAND EDITING AND TEST PLOTS

The following casts were edited for T and/or S in the top 10m only: 6,18,20,24,32,36,40,54,56,68,70, 74,86,102,112,114,128,132,140,142.

The following casts were edited more extensively: 16,26,34,38,44,48,58,60,62,66,72,80,82,84,88,90, 92,96,98,100,104,106,116-126,130,136,138,144.

Details are noted in the comment section of the individual files.

10. COMPARISON WITH BOTTLE DATA

There were three deep calibration samples taken during casts #70/71, 94/95 and 122/123. The differences between the bottle salinities and those from the CTD were -0.0024, +0.0042 and +0.0090psu. The samples were from 500, 300 and 1000db respectively. This CTD was used during cruise 9817 (July 1998) when the CTD was found to be within +/-0.011psu of bottle values. The instrument was recalibrated between these two cruises.

11. OTHER INTERCOMPARISONS

Historic ranges were superimposed on the profiles of casts 34,48,68,70,72,74,76,126,128,130,132          and 134 and showed that the salinity and temperature were within those ranges.

- There were no repeat stations.

- RECALIBRATION

There is insufficient data to consider recalibration.

13. FINAL PLOTS

Page plots were prepared using the edited data.

14. REMOVE CHANNELS

The following channel was removed from all casts: Conductivity_Ratio.

15.	PRODUCE FINAL FILES

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

16. PARTICULARS

1,2,3 - Test casts.

50/51 - Demonstration cast - deleted.

70/71 - Salinity calibration at 500m. The log says 1000m due to error in pressure range setting.

94/95 - Salinity calibration at 300m. The log says 600 due to error in pressure range setting.

122/123 - Salinity calibration at 992m. The log says 1980. Error discovered in pressure readings.

128/129 - Calibration for pressure range was changed before this cast.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	9819

Dates: 	Start:	 October 5, 1998			    End:	October  22, 1998

Location:	N.E. Pacific,Gulf of Alaska

Vessel:		W. E. Ricker

Scientific Party Chief:    John Morris

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
|  | Guildline | 8715 | 43825 | no | yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	Guildline /8715/43825

Cruise ID#:	9819

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure |  |  |  |  |  |
| Temperature | 57915 | Nov 98 | IOS |  |  |
| Conductivity | 58814 | Nov 98 | IOS |  |  |

| Calibration Coefficients |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Channel | Formula | Coefficients |  |  |  |  |  |
| Name | # | C1 | C2 | C3 | C4 | C5 | C6 |
| Temperature | 10 | .0123857 | .998904 |  |  |  |  |
| Conductivity | 10 | .000452327 | 1.00047 |  |  |  |  |