# PROCESSING NOTES

Cruise: 9821

Agency: PBS

Project: Hake Offshore Research

Geographic Area: N.E.Pacific

Scientific Party Chief: Mark Saunders

Platform: W. E. Ricker

Date of Cruise: 4 August 1998 - 24 August 1998

Processed by: Germaine Gatien

Date of Processing: 8 February 1999 - 11 February 1999

Number of original casts: 104

Number of casts processed:  47

# INSTRUMENT SUMMARY

The CTD used during this cruise was a Guildline 8715 / Serial Number: 43825

# SUMMARY OF QUALITY AND CONCERNS

During this cruise the CTD pressure range was set incorrectly so that the casts were only half as deep as planned.

# PROCESSING SUMMARY

- PRELIMINARY STEPS

The Log Book was obtained and note was made of problems that occurred during the cruise. The CTD used during this cruise is known to have had the pressure range set incorrectly, so that depths mentioned in the log are twice the actual depths.

The cruise summary sheet was completed.

The header summary and header check were examined. A note in the log says that the times are out by one hour. This will be corrected later. Note that the times given in the COMMENT section in some headers will still be in error.

There are errors in cast numbers 48 through 53 in the CTD log. The log gives casts #49 and #51 as the downcasts at stations 18 and 19. In fact, casts #48 and #50 are the downcasts for those stations. Cast #52 given as an upcast at station 19 is actually a downcast at station 20 and was aborted. All the information in the headers appear to be correct as do all the other cast numbers in the log.

The surface check gave an average value of 3.2db. A closer look at the files shows that the CTD generally started below the surface and the pressure at which the first good salinity values appear is about 1.0db. In the calibration step an offset of -1.0db will be applied to the pressure.

There were separate up and down casts. Only down-casts were processed.

- CONVERSION

- The data was converted using the IOS SHELL program ACQCONV.

- CALIBRATION

There are pre-cruise (9821cal.ccf) and post-cruise calibrations (9821post.ccf) available for this CTD,  done in July 1998 and November 1998, respectively. The pre-cruise calibration was found to be best for cruise 9817 which was run just before this cruise. By November 1998 when cruise 9819 was run, the post-cruise calibration was closer to bottle values. A test was run on two casts and the pre-cruise calibrations produced results closer to the bottle samples. The data was calibrated using file 9821cal.ccf which is attached. A pressure offset of -1.0db was applied.

4. DESPIKE

A test cast showed that running ADVANCED DESPIKE removed many salinity spikes so all casts were put through the routine, using the following parameters:

CHANNEL        FIT   OVER-  MIN       MAX       MIN         MAX         SPIKE

WIDTH LAP    VALUE   VALUE  STDDEV  STDDEV     TOL.

-------------------- ----- ----- ------- ------- ------- ------- -----

Pres 		55     5        -2.00       2000.00	.10000        999.000    3.00

Temp   	55     5         0.00           15.00	.00500        999.000    3.00

Cond    	55     5           .00            1.50	.00010        999.000    3.00

5. TIME COMPENSATION

Temp. probe Dist (m):   .0700  Sample period (sec):   .0400

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Low Sal & Last Press Min

Maximum Surface Pressure (relative):      20.00

Minimum Surface Salinity:     10.000

Surface Swell Pressure Tolerance:       1.00

Pressure filtered over width:   11

Swells deleted. Warning message if pressure difference of      2.00

Drop rates<    .30m/s (calculated over  11 points) will be deleted

Sample interval =       .04 seconds.

Due to warnings casts #6,21,23,53 and 120 were rerun with the Low Drop Rate feature turned off.

After this step SURFACE was rerun and an average value of 0.56db was found.

- FILTER

A test run was done (cast #4) to see if filtering improved the date. It removed some spikes so a median box car filter with filter width 21 was applied to the salinity channel for all casts.

- BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

Page plots were produced to aid in editing.

9. HAND EDITING AND TEST PLOTS

All casts except for #25, 42,88,101,111,117 and 120 required some editing. For most casts this entailed fairly extensive editing of salinity in the top 30m with only light editing below that. The following casts required more extensive editing: 4,6,19,36,53,57,61,82,90 and 96.

Details are noted in the comment section of the individual files.

10. COMPARISON WITH BOTTLE DATA

There were five deep calibration samples available for casts #6,36,63,96 and 101. (Several bottles including one deep sample were unusable because they were overfilled.) The average differences between the bottle salinities and those from the CTD were +0.002psu with a range of -0.007 to +0.008. The samples were from 500, 300, 500,390 and 500db respectively. (See 9821comp.xls for details.) This CTD was used during cruise 9817 which immediately preceded cruise 9821; at that time the CTD was found to be within +/-0.011psu of bottle values.

11. OTHER INTERCOMPARISONS

Historic ranges were superimposed on the profiles of casts #4 through #36 and showed that the salinity and temperature were well within those ranges.

Casts #44 and 46 were close and about 30 minutes apart, but the latter cast was shallow. At 100db the differences in temperature were ~0.1C° and those in salinity were ~0.02psu.

- RECALIBRATION

There is insufficient information to consider recalibration.

13. FINAL PLOTS

Page plots were prepared using the edited data.

14. REMOVE CHANNELS

The following channel was removed from all casts: Conductivity_Ratio.

At this time the time error mentioned in step 1 was corrected using routine ADTIMCH.

15.	PRODUCE FINAL FILES

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

16. PARTICULARS

1.-Aborted. Delete

7.-Aborted. Delete

38/39.-Aborted. Delete

52.-Aborted. Delete

63.-Aborted. Delete

86.-Lids left on, bad salinity data. Delete

88.- Repeat of cast #86

96.-Log Note: Bad wire angle, poor salinity data in top 100m.

97.-Aborted. Delete

120.-Cast to 12m only.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	9821

Dates: 	Start:	 August 4, 1998			    End:	August  24, 1998

Location:	N.E. Pacific,Gulf of Alaska

Vessel:		W. E. Ricker

Scientific Party Chief:    Mark Saunders

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
|  | Guildline | 8715 | 43825 | no | yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	Guildline /8715/43825

Cruise ID#:	9821

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure |  |  |  |  |  |
| Temperature | 57915 | July 98 | J L Tech Services | Nov 98 | IOS |
| Conductivity | 58814 | July 98 | J L Tech Services | Nov 98 | IOS |

| Calibration Coefficients |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Channel | Formula | Coefficients |  |  |  |  |  |
| Name | # | C1 | C2 | C3 | C4 | C5 | C6 |
| Temperature | 10 | .0104305 | .9999472 |  |  |  |  |
| Conductivity | 10 | .001018911 | .998756 |  |  |  |  |