PROCESSING NOTES

Cruise: 9822

Agency: IOS

Location: Georgia Strait

Project: 12KHz Sidescan

Chief Scientist: Mark Trevorrow

Platform: CCGS Vector

Date: 4 September 1998 – 16 September 1998

Processed by: Germaine Gatien

Date of Processing : 17 May 1999 – 19 May 1999

Number of original CTD casts: 28

Number of casts processed: 28

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0506) was used.

SUMMARY OF QUALITY AND CONCERNS

Comparison with salinity bottles was impossible because of errors in collecting salinity.

Other than surface pressure spikes the data was generally of good quality.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained and no problems were noted in it.

The salinity data was deemed invalid by Bernard Minkley due to the absence of liner inserts and overfilling of bottles.

The cruise summary sheet was completed.

The configuration file was obtained. In checking the calibration coefficients a very small error in the secondary temperature was found and corrected. The corrected file is 9822ctd.con.

- Conversion of Raw Data

The raw data was converted (without salinity) using conversion 9822ctd.con.

There were header errors in files 1,2,4,5,9,13,14,15,16,17,18,19,20,22,23,25,27 and 28 associated with surface pressure spikes. The errors (interchange of NMEA latitude, longitude and time) were corrected.

A preliminary check of the data shows that for casts #1 and 2 the time interval was 1.3755 seconds while for all other casts it was 0.25 seconds.

4. ALIGNCTD

The secondary conductivity channel was advanced by 0.060. The primary channel was aligned at sea by the deck unit. Comparing cast #10 before and after this step showed that it was effective at removing spikes from the secondary salinity where the temperature gradient was large. A test cast was done advancing the secondary conductivity by 0.073; this had much the same result as 0.060 so the latter was chosen since it has proven the better choice on other data sets.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on casts #10 and 21 to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

8. Test Plots and Channel Check

A sample of casts (#10 and 21) was plotted to check for agreement between the pairs of  T and C sensors. The temperature variations are very noisy but ignoring spikes are ~+0.0008C° below 150db while the conductivity  differences are about -0.003units. Salinity values varied by <0.004psu. These sensors were first used in June 1998 during cruise 9814 at which time the variations were on the order of 0.0002C°,     -0.0002units and -0.002psu. Given that there were much deeper casts during 9814 than during 9822 the two are reasonably consistent.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

10. Checking Headers

A header summary was produced and no problems found.

The surface check was run. The average surface pressure is 1.1db, but most values were outside the range used for the average, with values of –11 to –20db. All the casts with negative surface pressures are ones associated with the surface pressure spikes and faulty headers. Examining these casts closely shows that while the temperature and salinity values stay relatively constant the pressure rapidly changes from –15 to 0db. The conclusion is that the pressures are probably correct from 0 upwards, but the first 10 or 20 data records have faulty pressures.

The header check was produced and no problems found.

The cruise track was plotted and looks reasonable.

- DELETE

The descent rate was generally low averaging about 0.6m/s so the minimum drop rate was chosen to be 0.1m/s to minimize data loss.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0

Pressure filtered over width: 15

Swells deleted. Warning message if pressure difference of 2.00

Drop rates< 0.10 m/s (calculated over 6 points) was deleted.

Sample interval =       .25 seconds.

- BIN AVERAGE

The BIN AVERAGE routine would not run until the faulty surface pressure records were removed; this was done using a text editor on casts # 1, 2, 4, 5, 9, 13-20, 22, 23, 25, 27 and 28.

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =   1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

13. Test Plots

Profiles of all casts were examined (on screen) for any evidence of problems with the processing. The secondary sensors were extremely noisy during cast #5 and values fell outside the historic ranges. Primary sensors were selected for editing.

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

14. CTDEDIT

CTDEDIT was used to clean noise in S and/or T near the surface for the following casts: #1,2,6, 7,8,10,13,17,20,23 and 24.

CTDEDIT was used for more extensive cleaning in the following cast: #5

Note was made of the editing details in the relevant files.

15. Intercomparisons

Previous experience with these sensors -These sensors were used during cruise 9814 in June 1998 and bottle comparisons found the salinity to be low by 0.004psu and no recalibration was done. Primary sensors were found to be closest to bottle salinities.

COMPARE – Four hydro bottles were collected but the salinities were deemed invalid due to the absence of liner inserts and overfilling. The results were examined and three of the bottles were higher than the CTD by about 0.04, 0.02 and 0.03psu and the fourth was lower by 1.13psu.

Repeat casts – This is an area of rapidly changing conditions and all casts are shallow so no attempt was made to compare repeat casts.

Historic ranges - Plots were examined with historic ranges superimposed for casts #3, 4, 27 and 28. Temperatures around 200db (just under 10C) are above the historic records, while the salinities are well within them. Three of these casts can be compared with data from cruise 9836 in mid-October which also show temperatures slightly higher than the historic range. Ron Perkin confirms that data from Nanoose Bay indicates temperatures around 200db around 10C through the autumn and early winter of 1998/1999.

- Recalibration

There is no basis for recalibration of this data

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

18. REMOVE

The following channels were removed from all casts: Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

19. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

- Data collected with interval 1.3755s.

- Data collected with interval 1.3755s

3-28. Data collected with interval 0.25s.

# 5.   Secondary temperature and conductivity very noisy and T1 and S1 out of historic range.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	9822

Dates: 	Start:  4 Sept 1998 	  End:	16 Sept 1998

Location: Strait of Georgia

Vessel:		CCGS VECTOR

Chief Scientist:	Mark Trevorrow

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	 0506

Cruise ID#:		9822

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2371 | 8 May 98 | Factory |  |  |
| Conductivity | 2102 | 8 May 98 | Factory |  |  |
| Secondary Temp. | 2449 | 8 may 98 | Factory |  |  |
| Secondary Cond. | 2128 | 5 May 98 | Factory |  |  |
| Transmissometer | 197 | 1 June 98 | IOS |  |  |
| Pressure Sensor | 69698 | 12 Nov 97 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.