## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 30-Mar-2017 | Renamed fluorometer channel name from "Fluorescence:URU" to "Fluorescence:Calibrated" since the data from the fluorometer (on the CTD) has been converted to Chlorophyll-a using linear regression on observed extracted chl-a. Joe Linguanti |
| 4-Feb-2014 | The CTD salinity channel was removed from the Bottle files since the salinity data from the rosette files was not processed or calibrated. For details see section “Rosette Files, April 30, 2003” in file “\\OSD_Data_Archive\Cruise_Data\Documents\CS Wong’s Archive\ Station Papa Data Arching 2002.doc”. J.L. |
| 23-Feb-2000 | Doug Yelland applied calibrations to the Chlorophyll fluorescence channel. Original CTD files were copied to the WORK directory. Joe Linguanti |

PROCESSING NOTES

Cruise: 9836

Agency: IOS

Project: Globec /LaPerouse /COPRA/Georgia Strait/El Nino

Party Chief: Doug Yelland

Platform: J.P.Tully

Date:October 6, 1998 - October 17, 1998

Processed by: Germaine Gatien

Date of Processing : March 26, 1999 -  April 7 , 1999

Number of original CTD casts: 76

Number of casts processed: 75

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0443) was used.

SUMMARY OF QUALITY AND CONCERNS

The descent rate was very noisy causing many shed wakes. Despite this the data quality was generally good and little editing was required.

Irradiance (PAR) and Chl_fluorescence values are nominal and unedited.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained and note was made of problems that occurred during the cruise. Errors in the header for Cast #2 were corrected (GPS position and time readings were misplaced; this is an error noted in other cruises and associated with a surface pressure spike.)

The salinity data was obtained.

The cruise summary sheet was completed.

- Conversion of Raw Data

The raw data was converted using conversion file 9836CTD.con which is identical to the other con files except in the transmissivity section. As was found for the previous cruise 9835, the calibrations for transmissivity do not agree with the records on file nor with cruises 9829 (Sept. 98) and 9839 (Nov. 98) when the same equipment was used. The same conversion information was used for 9835 and 9836. ROS files were created separately to get salinity.

4. ALIGNCTD

The secondary conductivity channel was advanced by 0.060s since this value was found appropriate on the previous cruise using the same equipment. The primary channel was aligned at sea by the deck unit. Comparing cast #107 before and after ALIGNCTD shows that many salinity spikes were removed.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1: Std Dev = 2	Pass 2: Std Dev = 5     Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03 and 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on casts #5 and 107 to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

8. Test Plots

Casts (#5 and 107) were plotted to check for agreement between the pairs of  T and C sensors. The temperature differences are very noisy but average <0.0005C° around 1000db while the conductivity difference varies by ~0.0016units. Salinity values varied by about 0.0018psu. In each case the primary values are higher than the secondary. A quick check against bottle salinity values suggests that the primary salinity is high by about 0.004psu and the secondary low by about 0.013psu. These observations are very similar to those during cruise 9829 and 9835 which immediately preceded 9836. The primary  values will be selected.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

10. Checking Headers

A header summary was produced and a number of errors found in station names (casts #1,13,14,16,18,19,20,22 & 144) and in date/time (cast #2). Those were corrected in the *.ios files.

The surface check was run. The average surface pressure is 2.3db; it appears that the CTD was started below the surface on purpose.  No correction will be made to pressure.

The header check was produced and no errors found.

The cruise track was plotted and looks fine.

11. DELETE

The following DELETE parameters were used:

Surface Record Removal:  Last Press Min

Maximum Surface Pressure (relative):      20.00

Surface Swell Pressure Tolerance:        .50

Pressure filtered over width:   11

Swells deleted. Warning message if pressure difference of      2.00

Drop rates<    .33 m/s (calculated over  11 points) was deleted.

Sample interval =       .04 seconds.

Warnings for the following casts were investigated:

13 - No data was logged from 829db to 995 in the first downcast, so the CTD was raised to 945db and the downcast was resumed. The log notes a gap in data from 897-994db and notes that the OVERFLOW light was on. DELETE skipped a lot of the repeated downcast section so file 98360013.ios was edited to remove the upcast from 994 to 897db. DELETE was then rerun successfully.

18, 98, 101, 103, 107, 110 - There were reversals of the CTD that led to warnings by DELETE; the routine removed the data from the episodes of low or negative descent rate.

132 - Very shallow; delete the cast.

Profiles were plotted at this stage. For casts #16,18,72 and 107 historic ranges were superimposed.

- BOX CAR FILTER

The salinities are reasonable smooth so this step was skipped.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

14. Test Plots

Profiles of all casts were examined (on screen) for large inconsistencies between primary and secondary values and for any evidence of problems with the processing. Two problems were notes:

cast #98 - there was a bad spike in secondary salinity; primary sensors will be used.

cast #101 - there were spikes in all sensors at about 35db but in the primary salinity there appears to be a large offset that took about 30db to disappear. The secondary salinity looks o.k. so should be selected for this cast.

Page plots were produced using T1,S1 for cast #101 and T0,S0 for the rest. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

15. CTDEDIT

CTDEDIT was used to clean noise in S and/or T near the surface in the following casts: 4,11,22, 36,62,86,101,111,147,153 and 155.

CTDEDIT was used for more extensive cleaning in the following casts: 16 and 122

Note was made of the editing details in the relevant files.

- COMPARE

The CTD primary salinities were found to be high by about 0.005psu using the COMPARE routine and including samples from 200m down only. COMPARE was also run for the secondary sensors and the salinity was found to be low by about 0.013psu.

17. Other intercomparisons

Repeat casts - There were no repeat casts.

Previous experience with these sensors -  There were three cruises close in time to this one during which the same sensors (and same calibrations) were used. Two earlier cruises suggest that the CTD was high by 0.003psu (9829 - Aug/Sept 1998) and .0037psu (9836 -Sept/Oct 1998). A later cruise (9839 - Nov. 1998)  found the difference to be 0.0046psu, again with the CTD high.

Historic ranges - The plots of casts #16,18,72 and 107 show the temperature and salinity to be comfortably within the historic ranges.

- RECALIBRATION

Based on the observations of differences between the two salinity channels and the COMPARE results the salinity was recalibrated with an offset of -0.005psu for all casts except #101, and +0.013psu for cast #101. The recalibrations were also applied to the converted rosette files which were then named *.cab. COMPARE was rerun and produced satisfactory results.

19. Final Plots

Page plots were prepared using the edited data.

- REMOVE

The following channels were removed from cast #101: Primary Temperature, Primary Salinity, Scan_Number, Conductivity:Primary, Conductivity:Secondary and Flag.

The following channels were removed from all other casts: Secondary Temperature, Secondary Salinity, Scan_Number, Conductivity:Primary, Conductivity:Secondary and Flag.

20. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

2-Bad header. GPS position and time misplaced. Fixed.

13-The CTD acquisition failed to record data from about 830db down to 995db. The instrument was raised to 890db and then the downcast was resumed. File 98360013.ios was edited to remove the data from the time when the CTD was being raised from 995db to 830db, so that DELETE would select the right data.

18, 98, 101, 103, 107 and 110.- There were low descent rates and patches where the CTD reversed direction during the downcast which led to warnings in DELETE.

65 - There were small instabilities between 30 and 40db; they appear in both sets of sensors. Use of a median filter did not help, so the unfiltered data were used and were left unedited.

101 - There is a large area of offset primary salinity so secondary sensors will be used for this cast only.

132 - This was an acoustics calibration run, very shallow. Cast was deleted.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	9836

Dates: 	Start:	Oct.  6, 1998	    End:	Oct. 17, 1998

Location: Juan de Fuca /  West Coast Vancouver Island

Vessel:		J.P.Tully

Chief Scientist:		Doug Yelland

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | yes | yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443

Cruise ID#:		9836

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 15Aug98 | factory |  |  |
| Conductivity | 1766 | 18Aug98 | factory |  |  |
| Secondary Temp. | 2038 | 1May98 | factory |  |  |
| Secondary Cond. | 1762 | 7May98 | factory |  |  |
| Transmissometer | 182D | 7Aug98 | IOS |  |  |
| PAR | 4495 |  |  |  |  |
| Fluorometer | WS3S-417P |  |  |  |  |
| Pressure Sensor | 63507 | 11Jan96 |  |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.