## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27-Jan-2004 | Added nutrient data obtained from Melanie Quennville to the rosette files. |

PROCESSING NOTES

Cruise: 9909

Agency: IOS/UVIC

Location: Strait of Georgia / Juan de Fuca Strait

Project: Strait of Georgia Monitoring / COPRA

Chief Scientist: Diane Masson

Platform: CCGS VECTOR

Date: April 19, 1999 - April 23, 1999

Processed by: Germaine Gatien

Date of Processing : 4 May 1999 – 13 May 1999

Number of original CTD casts: 70

Number of casts processed: 70

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0443) was used.

SUMMARY OF QUALITY AND CONCERNS

There were large differences between the salinities from the CTD and those from bottles. The secondary sensors, which have consistently produced salinities about 0.01psu low in five previous cruises, behaved much differently during this cruise, with values about 0.06psu low early in the cruise gradually changing  to about 0.04psu low at the end. The primary conductivity cell was new and produced salinities low by about 0.015psu throughout the cruise. The data was of good quality.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained and note was made of problems that occurred during the cruise. The system clock was set to PDT for the first 8 casts. This only affects the *.BL files which are not used during processing so this will not be a problem.

There are header errors in casts #11, 50, 63 and 70 probably caused by pressure spikes. The header files were corrected.

The salinity data was obtained from Bernard Minkley

The cruise summary sheet was completed.

The configuration file was obtained and the calibrations constants were checked. The transmissivity calibrations were changed as those in the file were not the most up-to-date.

- Conversion of Raw Data

The raw data was converted (without salinity) using conversion file 9909ctd.con. ROS files were created separately with salinity.

4. ALIGNCTD

The secondary conductivity channel was advanced by 0.060. The primary channel was aligned at sea by the deck unit. Comparing casts #16 and 53 before and after this step showed that it was very effective at removing spikes from the secondary salinity where the temperature gradient was large.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

8. Test Plots and Channel Check

A sample of casts (#1, 6, 16, 48, 53, 62, 84 and 90) was plotted to check for agreement between the pairs of  T and C sensors. The temperature variations are very noisy but ignoring spikes are <0.001C° below 150db while the conductivity  differences range from 0.0023 to 0.0039units. Salinity values varied by ~0.023psu to 0.041. Primary values were higher in all cases.  The temperature variations are similar to those observed on other cruises using these sensors since the last calibration. The conductivity and salinity variations are much larger than usual and decreasing with time.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. Editing was required to transmissivity at a few points near the end of the upcast in casts 17 and 27 (values of 8.685e-17 were replaced)  before conversion was successful. The header of cast #15 was edited to enter the hemisphere of the latitude correctly. The rosette files were converted and renamed *.bot.

10. COMPARISONS

Because of concerns about the performance of the sensors comparisons were examined earlier in the processing than usual.

Previous experience with three of these sensors - Both temperature sensors and the secondary conductivity sensors were used for 5 recent cruises: 9829 (Aug/Sept), 9835 (Sept/Oct), 9836(Oct), 9839(Nov) and 9905(Jan). In each case the primary sensors were used. The primary salinities were found to be high by 0.003psu, 0.0037psu, 0.005psu, 0.0046psu and 0.003psu respectively. The differences between primary and secondary salinities were fairly consistent during those cruises ranging from 0.010 to 0.018psu with the primary higher than the secondary. Since the primary salinity was high by about 0.004, the secondary salinity must have been low by about 0.01psu.

COMPARE - Hydro bottle salinities were compared with those from the CTD using the IOSSHELL routine COMPARE. The primary sensors were found to be low by about 0.0156psu. The secondary sensors were low by up to 0.065psu.

TIME VARIATION IN SENSORS – Because the difference between the two sets of sensors appeared to decrease with time a study was made of the primary and secondary sensors. The data in the COMPARE files were averaged for each station and each sensor and plots made of the difference between the CTD and bottle salinities versus cast number. This indicated that the differences were fairly consistent for the primary sensors, but were decreasing for the secondary sensors.

11. Checking Headers

A header summary was produced and two errors found in station names were fixed.

The surface check was run. The average surface pressure is 1.9db.

The header check was produced.

The cruise track was plotted and looks reasonable.

- DELETE

Since the CTD was lowered slowly during many of the casts a study was done of the descent rate before running DELETE. A drop rate minimum of 0.1m/s was chosen since it did not appear that it would cause any significant loss of good data whereas using the usual 0.33m/s would do so. The log was examined and no problems found.

The following DELETE parameters were used:

Surface Record Removal:  Last Press Min

Maximum Surface Pressure (relative):      10.00

Surface Swell Pressure Tolerance:        1.0

Pressure filtered over width:   15

Swells deleted. Warning message if pressure difference of  2.00

Drop rates<    .10 m/s (calculated over 25 points) was deleted.

Sample interval =       .04 seconds.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =   1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

14. Test Plots

Profiles of all casts were examined (on screen) for any evidence of problems with the processing. Because the secondary salinity values are so low the primary sensors will be used.

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

15. CTDEDIT

CTDEDIT was used to clean noise in S and/or T near the surface or at the bottom for the following casts: #26,32,37,70,76,78, and 91.

CTDEDIT was used for more extensive cleaning in the following casts: #2,6,8,12,27,31,38,43.

Note was made of the editing details in the relevant files.

16. Intercomparisons

Repeat casts - There were no repeat casts.

Historic ranges - Plots were examined with historic ranges superimposed for several casts; while the data all fell within the ranges, the latter are so broad in this area that little is learned.

- Recalibration

Based on the intercomparisons the edited CTD files, the rosette files and *.sam files were all recalibrated using file 9909rcal.ccf to raise the salinity by 0.0156psu. COMPARE was rerun using the recalibrated salinities and the average differences (using the same data points as were used in the first run of COMPARE) were then found to be 6E-06psu.

19. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

19. REMOVE

The following channels were removed from all casts: Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

20. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

Casts 1-8 – The CTD Log contains a notes that time was recorded in PDT instead of UTC. The log times agree with the header files and they both coincide with the times that the chief scientist recalls for cast 8 and 9. The only headers with errors are the *.bl files which were not used in processing.

Cast 11 - Header problems due to surface pressure spike. Corrected

Cast 15 – Hemisphere omitted from latitude in header. Corrected

Cast 17 – Transmissivity data near end of upcast had to be edited to enable conversion to IOSSHELL.

Cast 27 – Transmissivity data near end of upcast had to be edited to enable conversion to IOSSHELL

Cast 50 - Header problems due to surface pressure spike. Corrected

Cast 63 - Header problems due to surface pressure spike. Corrected

Cast 70 - Header problems due to surface pressure spike. Corrected

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	9909

Dates: 	Start:	April 19, 1999    End:	April 23, 1999

Location: 	Georgia Strait

Vessel:		CCGS VECTOR

Chief Scientist:		Diane Masson

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0437 | Yes | yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0443

Cruise ID#:		9905

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 15Aug 1998 | factory |  |  |
| Conductivity | 2278 | 19Mar 1999 | factory |  |  |
| Secondary Temp. | 2038 | 1May 1998 | factory |  |  |
| Secondary Cond. | 2128 | 5May 1998 | factory |  |  |
| Transmissometer | 182D | 15Mar 1999 | IOS |  |  |
| Pressure Sensor | 63507 | 11Jan 1996 | factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.