## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 30-Mar-2017 | Renamed fluorometer channel name from "Fluorescence:URU" to "Fluorescence:Calibrated" since the data from the fluorometer (on the CTD) has been converted to Chlorophyll-a using linear regression on observed extracted chl-a. Joe Linguanti |
| 23-Feb-2000 | Doug Yelland applied calibrations to the Chlorophyll fluorescence channel. Original CTD files were copied to the WORK directory. Joe Linguanti |

PROCESSING NOTES

Cruise: 9911

Agency: IOS/UBC

Location: West Coast Vancouver Island

Project: WCVI

Chief Scientist: Dave Mackas

Platform: J. P. TULLY

Date: 3 May 1999 – 13 May 1999

Processed by:  Germaine Gatien

Date of Processing : 21 June 1999 –  15 July 1999

Number of original CTD casts: 81

Number of casts processed: 81

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0437) was used. A Wetstar fluorometer (#WSIS097), Seatech Transmissometer (#182D) and Par Sensor were mounted on the CTD.

SUMMARY OF QUALITY AND CONCERNS

The secondary sensors performed very well during this cruise. The primary sensors produced salinities about 0.02psu lower than bottle samples.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The CTD Log Book was obtained. Note is made that there were problems with longitude throughout the cruise; if there is no value or it is obviously wrong use the value in the header notes. Files from cast #12 were named 12110011 in error. These were renamed 99110012. There was a data overflow on cast #53. No GPS was logged for cast #114.

At the time of processing the salinity sheets were missing; copies of analysis done at sea were obtained.

The cruise summary sheet was completed.

The configuration files were obtained. The calibration in the files changed between casts #46 and 47. The earlier one listed an altimeter as part of the configuration. There does not appear to be any data in that channel. The calibration constants were correct for the temperature and conductivity.

The CTD log indicates that transmissometer #182D was used, but the configuration file gives #192D and #192D was recorded as the instrument mounted on CTD #0437 for previous cruises. Con file 0437ctd.con was prepared using the information in 99110047.con except for the substitution of the latest calibrations for transmissometer #192D. Note that it was later discovered that the transmissometer was indeed #182D; recalibration of transmissivity was done late in the processing.

- Conversion of Raw Data

The raw data was converted (without salinity) using conversion 0437ctd.con. No ROS files were created since there were no hydro files with salinity available at the time of processing.

4. ALIGNCTD

The primary conductivity was advanced by –0.013s and the secondary by +0.060s. Comparing cast #142  before and after this step showed that it was effective at removing spikes from the salinity where the temperature gradient was large.  (The choices of parameters was based on results of earlier cruises when +0.060s produced optimal spike reduction. The primary conductivity is aligned at sea by +0.073s. Doing the further alignment of –0.013s brings the net result to +0.060. The result of this further alignment was the reduction of noise in the primary channel in areas of large temperature change.)

- WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

8. Test Plots and Channel Check

A sample of casts (#36, 73,110 and 142) was plotted to check for agreement between the pairs of sensors. The temperature variations are very noisy but ignoring spikes are <0.001C° below 200db while the conductivity  differences are about 0.0013 to 0.0015units. Salinity values varied by ~0.016psu to 0.019. Primary values were higher for temperature and lower for conductivity and salinity. The differences are unusually large for conductivity and salinity.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

10. Checking Headers

A header summary was produced and errors found in station names, positions and times were fixed.

The surface check was run. The average surface pressure is 1.256db.

The header check was produced.

The cruise track was plotted and looks reasonable.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Low Salinity & Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 1.0

Minimum Salinity: 25 psu.

Pressure filtered over width: 15

Swells deleted. Warning message if pressure difference of  2.00

Drop rates<  .3 m/s (calculated over 11 points) was deleted.

Sample interval =       .04 seconds.

The log was examined and all warnings found to pertain to the bottom or upcast.

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =   1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

13. Test Plots

Profiles of all casts were examined (on screen) for any evidence of problems with the processing.

Page plots were produced using T1 and S1 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

14. CTDEDIT

CTDEDIT was used to clean noise in S and/or T near the surface or at the bottom for the following casts: #1, 27, 43, 52, 63, 67, 82, 88, 93, 115 and 117.

CTDEDIT was used for more extensive cleaning in the following casts: #64, 76, 123, 133 and 134.

Note was made of the editing details in the relevant files.

15. Intercomparisons

Previous experience with these sensors: All sensors have been recalibrated since last use. They were used during cruise #9906 that immediately followed 9911 at which time the secondary sensors performed very well with salinities lower than rosette bottle samples by 0.0003psu which is within the resolution of the instrument. The primary sensors were found to produce salinities that were low by 0.021psu.

COMPARE – At the time of processing no salinity data was available in hydro files so COMPARE was not run. Information was obtained from Doug Anderson showing the results of 35 bottle samples analysed at sea. In an EXCEL spreadsheet this information was used to do a comparison of the primary and secondary CTD salinities with those from the rosette samples. After excluding outlyers and points for which the depth was less than 175m, plots were made of the differences vs depth. From the resulting trendlines the primary sensors were found to be low by 0.015psu (at the surface) to 0.025psu (at 1500m). The secondary sensors were 0.0008psu low at all depths.

Repeat casts - Cast #43 was at the same site as #42 but only went to 200m. Comparing metre-averaged temperature and salinity along t -lines the differences are generally less than 0.01C° and 0.001psu which suggests that the CTD was performing in a stable fashion.  Cast #112 was at the same site as #123 and could be compared at 500m. Along t-lines the differences are about <0.01C° and <0.01psu; while the salinity difference is fairly high it is reasonable given that the casts are close to shore and 11 hours apart.

Historic ranges – A few casts were plotted with historic ranges for temperature and salinity superimposed. For casts #37, 67 and 137 all data fell within the historic ranges.

- Recalibration

Based on the intercomparisons no recalibration was considered necessary for salinity. However, due to the error in the transmissivity calibrations it was necessary to reconvert the original data using the correct calibrations; only pressure and transmissivity were converted. The files thus created were then converted to IOS Headers (some editing was required first), put through Delete and metre-averaged. The resulting files were named *.avt.

17. REMOVE

The following channels were removed from all casts (*.edt files): Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary, Scan_Number, Flag and Transmissivity.  MERGE was used to combine the resulting files, *.rem, with the *.avt files prepared as described in section 16. HEADEDIT was used to change the transmissivity serial number in the headers from 192D to 182D. (When HEADEDIT was run error messages noted that the remarks section of the headers was full.)

18. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

19. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

11 – Station name corrected (from C1 to C2).

12 – Originally named 12110011; renamed 99110012

40 – 43 Errors in station names corrected (was LC… should be LB…)

101 – System upload time does not match NMEA Time. Time edited to match NMEA Time, CTD log and Rosette log.

114 –No GPS logged. Latitiude, longitude and time corrected in headers of 99110114.ios

134 – Station name corrected (from LJ9 to LJ8)

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	9911

Dates: 	Start: 3 May 1999 	  End:	13 May 1999

Location: West Coast Vancouver Island

Vessel:		CCGS  TULLY

Chief Scientist:		Dave Mackas

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0437 | Yes | Yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0437

Cruise ID#:		9911

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 11Mar1999 | factory |  |  |
| Conductivity | 1729 | 23Mar1999 | factory |  |  |
| Secondary Temp. | 2142 | 13Mar1999 | factory |  |  |
| Secondary Cond. | 1762 | 16Mar1999 | factory |  |  |
| Transmissometer | 182D | 15Mar1999 | IOS |  |  |
| Pressure Sensor | 63502 | 11Jan1996 | factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.