## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 10-Jan-2006 | Surface loop data was added to the archive. The data was acquired from John Morris at PBS. The data archived in September 2005 was replaced with this new set since chlorophyll was included. The original spreadsheet file from John and more detailed processing notes can be found in the “Cruise_Data\Documents” directory. Any questions regarding this data should be directed to John Morris. J.L. |
| 26-Sep-2005 | Added loop data to the archive. J.L. |

PROCESSING NOTES

Cruise: 9913

Agency: OSAP

Location: Gulf of Alaska

Project: High Seas Salmon/Globec/Nestucca

Chief Scientist: David Welch

Platform: W.E. Ricker

Date: 17 May 1999 – 28 May 1999

Processed by: Germaine Gatien

Date of Processing: 5 August 1999 – 19 August 1999

Number of original CTD casts: 40

Number of casts processed: 40

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (# 506) was used.

SUMMARY OF QUALITY AND CONCERNS

The data was generally of good quality. While no deep bottles were available the surface water was well-mixed so that useful information could be gathered from the surface samples. There appears to be a drift in the primary sensors.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained.

The surface salinity data was obtained.

The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants were checked.

Errors in the stations identification were corrected in the headers for casts #13 and 14.

- Conversion of Raw Data

The raw data was converted (without salinity) using conversion file 99130001.con.

4. ALIGNCTD

A study was done on the choice of parameters for ALIGNCTD. Test runs were made on cast #1, examining the effect of different choices on the spikiness of salinity in an area of large temperature gradient. The smoothest traces were produced by advancing the primary conductivity channel by –0.038s and the secondary conductivity by 0.035s. The primary channel was aligned by +0.073s by the acquisition system, so the net effect is that both channels are advanced by +0.035s.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1 Std Dev = 2	Pass 2 Std Dev = 5	Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels ( = 0.03; 1/ = 9.0.) A check of cast #1 was made comparing the effect on T-S curves of running CELLTM; it was found that the up and down traces were a little closer after running CELLTM.

Note: A closer study was made during cruise 9914 and it was determined that  = 0.025 and 1/ = 7.0 were slightly better choices, but the difference is very slight. In particular it would have no effect at the surface because the temperature gradient was low; hence it will not account for the differences between the loop bottles and CTD values described in section 13.

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

8. Test Plots and Channel Check

A sample of casts (#1, 9 & 38) was plotted to check for agreement between the pairs of T and C sensors. The temperature variations are very noisy but ignoring spikes are <0.001C° below 100db while the conductivity differences range from 0.0003 to 0.0008units. Salinity values varied by ~0.003 to 0.008psu increasing with depth. The primary values were lower in all cases.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

10. Checking Headers

A header summary was produced and an error in the station name for cast #9 was fixed.

The surface check was run. The average surface pressure is 5.726db.

The header check was produced and no errors found.

The cruise track was plotted and looks reasonable.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Low Salt & Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 1.0

Minimum Surface Salinity: 25.0

Pressure filtered over width: 15

Swells deleted. Warning message if pressure difference of 2.00

Drop rates< .3 m/s (calculated over 11 points) was deleted.

Sample interval =  .04 seconds.

The DELETE log was examined and the only warnings that did not correspond to observations at the bottom or during the upcast were ones for cast #31. The descent rate was derived and plotted for this cast and found to be extremely noisy, particularly at 758db where the warnings occurred. It appears that the CTD was stopped and moving up and down for some time. No good data was deleted.

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =   1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

13. Intercomparisons

Previous experience with these sensors – The primary sensors were both new. The secondary conductivity sensor was used with a different temperature sensor during 9839 (Nov. 98) when it produced salinities which were high by about 0.015psu; the secondary temperature sensor was used during 9822 (Sept. 98) but no salinity comparison was possible for that cruise.

COMPARE – The only bottle salinities available are loop samples from the surface. An Excel spreadsheet was prepared comparing these values with the shallowest values from the metre-averaged CTD files. The results indicate that the primary CTD salinity is low by about 0.007psu and the secondary is low by 0.0003psu. A few deep casts were studied and the secondary salinity was found to be higher than the primary by from 0.007 to 0.009psu. This is consistent with the surface observations and suggests that the secondary values are likely to be more accurate.

When the differences between the CTD salinity and the loop bottle salinity were plotted against cast number there was a hint that the primary sensors may have been drifting during the cruise. To further investigate the temporal variation the IOS Shell routine COMPARE was used to compare the primary and secondary salinities versus depth and file pair #. Here the trend is much clearer. The differences are increasing with time. These sensors were used on cruise 9914 during which this trend continued. While this analysis does not indicate which sensor was drifting, the surface bottle comparison suggests that the primary is at fault.

Repeat casts – There are no repeat casts.

Historic ranges - Plots were examined with historic ranges superimposed for casts #11 and 13 and salinities were found to be comfortably within the range. Both temperature traces were slightly above the range for cast #13 around 950db.

14. Test Plots

Profiles of all casts were were produced. Cast #36 showed large differences between the primary and secondary temperature traces.

Based on the intercomparison the secondary sensors were chosen for further processing. Page plots were produced using T1,S1 for all casts. These plots were examined for spikes and instabilities. Cast #36 was found to have very noisy secondary temperature and salinity while the primary values looked fine, so T0,S0 was chosen for this cast only.

15. CTDEDIT

No editing was necessary.

- Recalibration

No recalibration was done to the casts using secondary sensor data. For cast #36 an offset was applied to the primary salinity to bring it into line with the secondary values which were used for all other casts. The choice of +.008psu was based on the differences between the two traces at the bottom of cast #36.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

18. REMOVE

The following channels were removed from all casts except #36: Scan_Number, Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The following channels were removed from cast #36: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

19. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

Cast #31 – Descent rate very noisy; DELETE warnings around 758db. The data from 750-760db is suspect due to the CTD moving up and down repeatedly.

Cast #36 – The secondary temperature was very noisy. Primary values were used and salinity recalibrated to match the secondary values used at all other casts.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	  9913

Dates: 	Start: 17 May 1999 	  End:	28 May 1999

Location:  North Pacific / Gulf of Alaska

Vessel:		W.E. Ricker

Chief Scientist:	 David Welch

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | no | yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		9913

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 24Mar99 | factory |  |  |
| Conductivity | 2280 | 19Mar99 | factory |  |  |
| Secondary Temp. | 2449 | 8May98 | factory |  |  |
| Secondary Cond. | 1763 | 18Aug98 | factory |  |  |
| Transmissometer | 197 | 14Jan99 | IOS |  |  |
| Pressure Sensor | 69698 | 11Dec97 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.