## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 1-Nov-2005 | Added Chlorophyll data to rosette files. J.L. |
| 27-Jan-2004 | Merged nutrients acquired from Linda White to rosette files. |

PROCESSING NOTES

Cruise: 9916

Agency: IOS

Project: Strait of Georgia

Chief Scientist: Diane Masson

Platform: J. P. TULLY

Date: 24 June 1999 –27 June 1999

Processed by: Germaine Gatien

Date of Processing: 28 June 1999 - 30 July 1999

Number of original CTD casts: 70

Number of casts processed: 69

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0437) was used. Two Wetstar fluorometers (#WSIS097 and WS3S-417P), Seatech Transmissometer (#182D) and Par Sensor (GSP-200L4S) were mounted on the CTD.

# SUMMARY OF QUALITY AND CONCERNS

The data was in good order and required only light editing.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The CTD Log Book was obtained.  Salinity sheets were obtained from Bernard Minkley. Hydro files were not available at the time of processing.

The cruise summary sheet was completed.

The configuration files provided had an error in the transmissivity calibrations. This was changed and the resulting file was named 04379916.con.

- Conversion of Raw Data

The raw data were converted (without salinity) using conversion 04379916.con. Errors in the headers of cast #2 (associated with a surface pressure spike) were corrected.

A separate conversion was done to create ROS files with salinity for those casts for which bottle salinities were collected.

Cast 99100063 was renamed 99160001. Cast 99100064 contained no useful data and was deleted.

4. ALIGNCTD

The primary conductivity was advanced by –0.013s and the secondary by +0.060s. (Since the primary channel was advanced by +0.073s at sea, the net effect is to advance both channels by +0.060s.) A check before and after alignment showed that spikes were reduced in areas of large temperature gradient.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2		Pass 2    Std Dev = 5		Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

8. Test Plots and Channel Check

A sample of casts (#8, 45 and 68) was plotted to check for agreement between the pairs of sensors. The temperature variations are very noisy but ignoring spikes are <0.001C° below 150db while the conductivity differences are about 0.001units. Salinity values varied by ~0.01. Primary values were higher in each case. A quick check of CTD salinity versus bottle salinity suggests that the secondary salinity is closest to the bottle values.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. A number of files had to be edited to replace values of transmissivity and/or descent rate of the form 8.685e-17 with 0.000 or 0.00 before conversion would run. (Casts #2,14,22,37,60,62,63)

10. Checking Headers

A header summary was produced and no errors found.

The surface check was run. The average surface pressure is 1.125db.

The header check was produced.

The cruise track was plotted and looks reasonable.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 1.0

Pressure filtered over width: 11

Swells deleted. Warning message if pressure difference of  2.00

Drop rates<  .3 m/s (calculated over 11 points) was deleted.

Sample interval =       .04 seconds.

The log was examined and all warnings found to pertain to the surface or upcast.

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =  .000

Average value will be used

Interpolated values are NOT used for empty bins

13. Test Plots

Profiles of all casts were examined for large differences between sets of sensors. None were found.

Page plots were produced using T1,S1 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

14. CTDEDIT

CTDEDIT was used to clean noise in S and/or T near the surface or at the bottom for the following casts: #2,5,6,15,17,28,30,41,43,45,47,54,55,63,66,68,69,71,72 and 73.

CTDEDIT was used for mid-depth editing of the following casts: #42 and 52.

Note was made of the editing details in the relevant files.

15. Intercomparisons

Previous experience with these sensors – All sensors were used on cruise 9910 but that cruise has not yet been processed; the secondary sensors were used during 9911 and 9906 in May 1999 and the secondary sensors were found to be low by 0.0003psu.

COMPARE – At the time of processing bottle salinities were available on paper, but the files had not been prepared for use in the COMPARE routine. Ten samples were compared with averaged values from the rosette files and the results suggest that the primary sensors were high by about 0.008psu and the secondary low by about 0.0005psu. See file “compare.xls” for this comparison.

Historic ranges - Plots were examined with historic ranges superimposed for several casts; all fell within those bounds.

- Recalibration

Since the intercomparisons of the secondary salinity are consistent with the results for cruise 9911 and 9906 for which the errors were found to be within the resolution of the instrument, no recalibration was deemed necessary.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

18. REMOVE

The following channels were removed from all casts: Scan number, Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

19. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

Cast 99160033 – The UBC fluorometer was used for this cast; the calibration in the con file is the same as that used for the OSAP fluorometer.

Cast 99100063 was rename 99160001.

Cast 99100064 contained no useful data and was deleted.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	9916

Dates: 	Start: 24  June 1999 	  End:	27 June 1999

Location: Strait of Georgia

Vessel:		CCGS  TULLY

Chief Scientist:	Diane Masson

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0437 | Yes | Yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0437

Cruise ID#:		9916

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2062 | 5 May1999 | Factory |  |  |
| Conductivity | 1748 | 5 May1999 | Factory |  |  |
| Secondary Temp. | 2142 | 13 Mar1999 | Factory |  |  |
| Secondary Cond. | 1762 | 16 Mar1999 | Factory |  |  |
| Transmissometer | 182D | 15 Mar1999 | IOS |  |  |
| Fluorometer (OSAP) | WS1S-097 |  |  |  |  |
| Fluorometer (UBC) | WS3S-417P |  |  |  |  |
| PAR Sensor GSP-200L4S | 4495 |  |  |  |  |
| Pressure Sensor | 63502 | 11 Jan1996 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.