# Linda White                                                        July 5, 1999

7036 Bickstan Place

Brentwood Bay, B.C.

V8M 1G6

## Diane Masson

OSAP

Institute of Ocean Sciences

9860 West Saanich Road

Sidney, B.C.

V8L 4B2

Summary Report: 	Nutrient analysis – Georgia Strait Cruise # 9916

Nitrate, Silicate and Phosphate

Summary

Fresh seawater samples were analysed, shipboard, on a Technicon Autoanyzer according to procedures stated in “Institute of Ocean Sciences Nutrient Methods and Analysis” by J.Barwell-Clarke and F. Whitney 1996.

An artificial seawater blank separated each station and a medium check standard was analysed as a sample after the second set of samples.

Standard envelopes of 0, Low, Medium and High nutrient concentrations were analysed at the beginning, middle ( when many samples analysed ) and end of the analytical run.

Duplicate samples were run at the end of each day.

## Observations

The surface layer of seawater was heavily influenced by the Fraser River plume.  The top 3 depths were centrifuged to remove silt particles prior to analysis from the following stations: 65, 62, 56, 46, 42, 39, 27, 2, 3, 6, 9, 12, 14, 16,

Phosphate turbidity was analysed at these stations also.

Friday, June 25, 1999, the ship’s crew  were scrubbing the decks with detergent.  It would appear from the results that there has not been any contamination of the samples.

Duplicate samples have been stored in climate chemistry’s freezer if re-runs are required. Note that duplicate samples were thawed for 48 hours, by accident, when they were offloaded.

Analysis went very well.  The transition from Cruise 9910 to 9916 at Ogden Point went very smoothly.

## Quality Assurance data

## Pooled Standard Deviation  Sp = Square root (sum of diff sqrd/2n)

Samples

Sample Duplicates analysis	Nitrate	Silicate 	Phosphate

0.0244	0.2681	0.1927

n = 33		n = 33		n = 32

## Relative Standard Deviation %

Standards		N03 24.0 µM	Si 60.0 µM		P04 2.0µM

# Calculated

Medium check	23.9 +/-0.27		60.0 +/- 0.17	2.01 +/- 0.03

Analysed		n = 4			n = 4			n = 4

Relstdev %		1.13%		0.29%		1.30%

This report includes data results on an Excel spreadsheet with the following sheets: data and quality assurance results

Filename: 9916nuts.xls