## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27-Jan-2004 | Added nutrient data obtained from Melanie Quennville to the rosette files. |

PROCESSING NOTES

Cruise: 9920

Agency: OSAP

Location: Strait of Georgia

Project: Chief Scientist: Diane Masson

Platform: CCGS VECTOR

Date: 11 August 1999 – 15 August 1999

Processed by: Germaine Gatien

Date of Processing: 28 October 1999 – 2 November 1999

Number of original CTD casts: 70

Number of casts processed: 70

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0530) was used.

SUMMARY OF QUALITY AND CONCERNS

The data was in good order and required little editing.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained.

The salinity data was obtained.

The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants were checked.

- Conversion of Raw Data

The raw data was converted (without salinity) using conversion file 99200001.con. ROS files were created separately with salinity.

- ALIGNCTD

The primary conductivity was aligned by the acquisition system by 0.073s. A study was made of cast #5 to determine the best setting for removing spikes in the secondary salinity. Choices of 0.050s and 0.060s each showed very notable improvement over no alignment and slight improvement over 0.073s. Though it makes little difference to the primary channel it was decided to use the same choices for both channels. So for all casts, the primary conductivity channel was advanced by –0.013s (for a net advancement of 0.060s) and the secondary conductivity was advanced by 0.060s. This is the same setting used for the same CTD setup during cruise 9931.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in pressure, temperature and conductivity.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

8. Test Plots and Channel Check

A sample of casts (#5,27,55,77) was plotted to check for agreement between the pairs of T and C sensors. The temperature, conductivity and salinity variations were of the order +0.002C°, -0.0007units and         -0.0085psu, respectively. (+/- indicate that the secondary/primary values were higher.)

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. Errors in the Sea-Bird headers had to be corrected for casts #1 and 2 to enable conversion.

10. Checking Headers

A header summary was produced and a few errors found in station names were fixed.

The surface check was run. The average surface pressure is 1.87db.

The header check was produced and no problems found.

The cruise track was plotted and looks reasonable.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 1.0

Pressure filtered over width: 15

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted. See note below for one exception.

Sample interval =       .04 seconds.

NOTE: All warnings pertained to upcast or near surface data except for cast #29. During that cast the CTD moved very slowly from 176db to 183db and all data was lost between these two depths when the Low Drop Rate of 0.3m/s was used. DELETE was rerun using 0.1m/s and this left useful data while removing sections during which the CTD moved upwards. However, this data may be less reliable than the other parts of the cast.

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

13. Intercomparisons

Previous experience with these sensors – All sensors were used during cruise 9931 which immediately preceded this cruise. At that time drift was noticed in salinity for both channels: the primary salinity started low by ~0.01psu and ended high by ~0.005psu; the secondary salinity started low by ~0.009psu and ended low by ~0.005psu. The secondary salinity was recalibrated by adding 0.008psu. The primary conductivity sensor was also used for cruises 9829,9835,9836,9839 and 9905 when the salinity was found to be high by 0.003 to 0.005psu; however, the sensor has been recalibrated since that time.

COMPARE - Hydro bottle salinities were compared with those from the CTD using the IOSSHELL routine COMPARE. The primary and secondary salinities were found to be high by ~0.005psu and low by  ~0.0035psu, respectively. There is a lot of scatter in the data and there appears to be a pressure dependence with the CTD salinity increasing with depth relative to the bottle salinity. Looking at the variations with time there is no evidence of a drift like that observed during cruise 9931, which immediately preceded 9920. The primary salinity shows a slight variation with time, while the secondary seems quite flat.

Historic ranges - Plots were examined with historic ranges superimposed for several casts #1-6,10-12,20-23, and 41-42. All traces were within those ranges.

Repeat Casts – The only repeat casts were 2 at station CPF1; plots show differences of <0.03psu and <0.1C along a t-surface near 300m. This is an area of large variations so these results are reasonable.

14. Test Plots

Profiles of all casts were examined for any evidence of problems with the processing. None were found.

Page plots were produced using T1,S1 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

15. CTDEDIT

CTDEDIT was used to clean noise in S and/or T in the top 20m for the following casts: #1,18,59.

CTDEDIT was used for more extensive cleaning in the following casts: #3,12,16 and 69.

Note was made of the editing details in the relevant files.

- Recalibration

Based on the intercomparisons the edited CTD files, the rosette files and *.sam files were all recalibrated using file 9920rcal.ccf to raise the salinity by 0.0035psu. COMPARE was rerun using the recalibrated salinities and the average differences (using the same data points as were used in the first run of COMPARE) were then found to be -0.0005psu.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

18. REMOVE

The following channels were removed from all casts: Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary, Scan_Number and Flag.

19. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

Cast #1 – Bad header associated with pressure spikes; header corrected before conversion to IOS Headers.

Cast #2 – Error in longitude – corrected before conversion to IOS Headers.

Cast #29 – CTD Log Note: Ship drifting at 3knots.

– In DELETE the Low Drop Rate was changed from 0.3m/sec to 0.1m/sec to avoid losing useful data since the CTD moved very slowly from 176 – 183db.

Cast #39 – CTD Log note: hit jellyfish.

Cast #69 – CTD Log note: hit bottom.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:  9920

Dates: 	Start: 11 August 1999      End: 15 August 1999

Location:  Strait of Georgia

Vessel:	CCGS Vector

Chief Scientist:	  Diane Masson

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0530 | Yes | Yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial #:	SEABIRD/911+/	0530

Cruise ID#:		9920

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2030 | 16Jun99 | Factory |  |  |
| Conductivity | 1766 | 6May99 | “ |  |  |
| Secondary Temp. | 2374 | 5Jun99 | “ |  |  |
| Secondary Cond. | 2173 | 29May99 | “ |  |  |
| Transmissometer | 192D | 20Jul99 | IOS lab |  |  |
| Pressure Sensor | 75636 | 4Jun99 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.