## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 9929

Agency: Stock Assessment/OSAP

Location: West Coast Vancouver Island

Project: COPRA/La Perouse

Chief Scientist: Dan Ware

Platform: CCGS W.E. RICKER

Date: 9 August 1999 – 16 August 1999

Processed by: Germaine Gatien

Date of Processing: 28 January 2000 – 31 January 2000

Number of original CTD casts: 14

Number of casts processed: 12

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0506) was used.

SUMMARY OF QUALITY AND CONCERNS

Only the primary conductivity data is useful and there is doubt about the temperature calibrations since the two sensors gave values different by more than 1C. There were no salinity calibration samples. There were pressure spikes in 3 casts.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained. The primary conductivity sensor is identified as #2270 in the log, but #2280 in the con file. The con file is presumed to be correct since there is no sensor #2270.

No salinity calibration data was available.

The cruise summary sheet was completed.

The configuration files used at sea did not have the most up-to-date calibration constants for the secondary sensors. A new configuration file 9929ctd.con was prepared with the latest calibrations.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted (without salinity) using conversion file 9929ctd.con. There were no rosette casts. The CTD log notes problems with a number of casts. A preliminary look shows that the secondary conductivity did not operate properly. There also appears to be a large offset between the two temperature sensors. It was decided to proceed with processing the primary data only.

Errors in the headers caused by pressure spikes were fixed in casts # 5, 9 and 13 before processing further.

4. ALIGNCTD

The primary conductivity channel was advanced by –0.018s. Since it was advanced at time of acquisition by +0.073s the net effect was to advance by +0.055s.

Comparing casts # 9 and 15 before and after this step showed that it was effective in removing spikes from the salinity where the temperature gradient was large.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in pressure, temperature and conductivity.  Parameters used were:

Pass 1    Std Dev = 2			Pass 2    Std Dev = 5			Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for the primary only (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on casts #13 and 32 to calculate the differences between primary and secondary channels for temperature and the descent rate.

8. Test Plots and Channel Check

As noted earlier there are large differences between primary and secondary temperatures (averaging about 1.15C at depth) and the secondary conductivity data is bad. The descent rate was about 1m/s.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

10. Checking Headers

A header summary was produced and no errors found.

The surface check was run. The average surface pressure is 5.3db. Casts # 5, 9 and 13 have negative surface pressures. These were casts with pressure spiking.

The header check was produced.

The cruise track was plotted and looks reasonable.

11. Test Plots

Profiles with historic range plots superimposed were produced for all casts to compare up and down casts, check for processing problems and especially to check the temperature calibration. The secondary temperature is below the historic range below 150db, but the primary temperatures look reasonable.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 1.0

Pressure filtered over width: 15

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =       .04 seconds.

COMMENTS ON WARNINGS: All warnings refer to surface and bottom depths only.

- CTDEDIT

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

CTDEDIT was used to clean noise in S and T in the top 10db for cast #9,11,13,15 and 30. Note was made of the editing details in the relevant files.

14. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

15. Intercomparisons

Previous experience with these sensors – The primary conductivity sensor was used during cruises 9913, 9933 and 9938; there were no rosette samples, but comparisons with a few deep bottles during 9933 showed salinity low by 0.013psu and loop samples during 9913 and 9938 showed values low by 0.007 and 0.021psu.

COMPARE – There are no bottles available.

Historic ranges - Plots were examined with historic ranges superimposed for all casts. As noted earlier the secondary temperature is below the historic range for depths greater than 150m. The primary temperature is within the historic range. Only a small section of the salinity trace for cast #34 (around 40db) falls below the historic range. Only 2 casts are deeper than 200db.

- Recalibration

While there is little evidence to justify recalibration, it does seem likely that the salinity is low. The values found during 9933 will be used to recalibrate the salinity, applying an offset of +0.013psu.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

- REMOVE

The following channels were removed from all casts: Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

19. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

1.Cast aborted. Delete.

4.Pumps not operating. Delete

5. Secondary conductivity data bad.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 9929

Dates: 	Start:  9 August 1999		  End:	16 August 1999

Location:   West Coast Vancouver Island

Vessel:   CCGS W. E. RICKER

Chief Scientist:	Dan Ware

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		9929

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 24MAR1999 | factory |  |  |
| Conductivity | 2280 | 19MAR1999 | factory |  |  |
| Secondary Temp. | 2449 | 28JUL1999 | factory |  |  |
| Secondary Cond. | 1763 | 28JUL1999 | factory |  |  |
| Transmissometer | 197 | 1JAN1999 | IOS |  |  |
| Pressure Sensor | 69698 | 12NOV1997 | factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.