## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 9931

Agency: OSAP

Location: BC Inlets

Project: BC Inlets

Chief Scientist: Dario Stucchi

Platform: CCGS VECTOR

Date: 26 July 1999 – 11 August 1999

Processed by: Germaine Gatien

Date of Processing: 9 September 1999 – 21 September 1999

Number of original CTD casts: 159

Number of casts processed: 158

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTDs  (#0550) were used with transmissometer (#192D) and Wetlab fluorometer (WSIS–097).

SUMMARY OF QUALITY AND CONCERNS

The data contained many spikes in pressure and conductivity. The calibration of the primary conductivity sensors drifted significantly with time and with depth.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book was obtained.

The salinity data was obtained.

The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants were checked.

- Conversion of Raw Data

The raw data was converted (without salinity) using conversion files 99310001.con. ROS files were created separately with salinity. A preliminary check showed all channels present, reasonable agreement between pairs of sensors and between up and down-casts. There were a lot of pressure and conductivity spikes in the data.

- ALIGNCTD

Based on a study of two casts it was decided to advance the primary conductivity channel by –0.013s and the secondary conductivity by 0.060s. The primary channel was aligned by +0.073 by the acquisition system, so the net effect is that both channels are advanced by +0.060s.

5. WILDEDIT

This data set had many spikes. Program WILDEDIT was used to remove spikes in Pressure, Temperature and Conductivity.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

A check of a few casts shows that the spikes were removed.

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

8. Test Plots and Channel Check

A sample of casts (#6,22,44,109 & 147) was plotted to check for agreement between the pairs of T and C sensors. The following table shows typical differences between pairs of channels. Note that while the temperature channels are very consistent there is a steady change in conductivity and salinity. During the casts there was a drift with depth towards more negative values of C and S.

| Cast # | Depth | Tsec – Tpri | Csec – Cpri | Ssec – Spri |
| --- | --- | --- | --- | --- |
| 6 | 350 | +0.002 | +0.0003 | +0.0012 |
|  | 500 | +0.002 | +0.00025 | +0.0008 |
| 22 | 350 | +0.002 | +0.0002 | +0.0002 |
|  | 500 | +0.002 | +0.00015 | -0.0005 |
| 44 | 350 | +0.002 | -0.00002 | -0.0018 |
|  | 580 | +0.002 | -0.00005 | -0.002 |
| 109 | 350 | +0.002 | -0.0005 | -0.0075 |
| 147 | 350 | +0.002 | -0.00065 | -0.0085 |

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers (including the rosette files.) Some files required editing of the headers before conversion was possible.

10. Checking Headers?

A header summary was produced and errors found in station names and times were fixed.

The surface check was run. The average surface pressure is 1.683db.

The header check was produced.

The cruise track was plotted and looks reasonable.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 1.0

Pressure filtered over width: 15

Swells deleted. Warning message if pressure difference of 2.00

Drop rates< .3 m/s (calculated over 15 points) was deleted.

Sample interval =       .04 seconds.

There were many warnings for cast #132; examination of the file shows that the only useful data for this cruise is at the surface and around 16db.

There were warnings that no records were removed from casts #152, 153 and 159 because they were less than 20db deep.

All other warnings pertained to the surface, bottom or upcast.

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

- Intercomparisons

Previous experience with these sensors – These sensors have not been used on other cruises since the last calibration.

COMPARE - Hydro bottle salinities were compared with those from the CTD using the IOSSHELL routine COMPARE. The results suggest that the calibration of the primary conductivity was changing significantly with time and depth. Near the beginning the primary sensors gave salinity low by about       –0.01psu but by the end is high by about +0.005. The secondary sensors show a slight change with time,  from –0.009 to –0.005psu. There is a lot of scatter in the compare results. A linear regression gives a value of –0.0078psu at 500db and averaging all points from 200db or below gives the same result.

Historic Ranges – Range plots were not available for most of the casts but the one for cast #81 showed that the data was within the historic range for T and S.

14. Test Plots

Profiles of all casts were examined (on screen) for any evidence of problems with the processing.

Page plots were produced using T1,S1 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

15. CTDEDIT

CTDEDIT was used to clean noise in S and/or T near the surface or at the bottom for the following casts: #36,76,99,100,105,118,121,125,128,130,147,150 and 152.

CTDEDIT was used for more extensive cleaning in cast #131.

Note was made of the editing details in the relevant files.

- Recalibration

Based on the intercomparisons the edited data files, the rosette files and *.sam files were all recalibrated using file 9931rcal.ccf to raise the salintiy by 0.008psu. COMPARE was rerun using the recalibrated salinities and the average differences (using the same data points as were used in the first run of COMPARE) were then found to be +0.0002psu at 500db.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

18. REMOVE

The following channels were removed from all casts: Scan_Number, Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The fluorometer channel was removed from the following casts: 7-10 and 17-19.

19. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

99310007 – Error in station name corrected.

99310036 – Error in station name corrected.

99310058 – Errors in headers (data file and rosette file) associated with surface pressure spike corrected.

99310078 – Times in CTD log are local. Headers are in UTC. Header errors were fixed to enable conversion to IOS Headers.

99310094 – Error in station name corrected.

99310130 - Error in station name corrected.

99310132 – Bad header due tosurface pressure spike. Fixed header file and dat file. When converted the pressures were found to be full of spikes, with almost no useful data. Delete.

99310155 – In the CTD log there are tow casts named #155. The first at station A5, was overwritten by the second, which is for station A4.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 9931

Dates: 	Start:     26 July 1999           	  End:	11 August 1999

Location: BC Inlets

Vessel:	CCGS VECTOR

Chief Scientist:	  Dario Stucchi

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | yes | yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0550

Cruise ID#:		9931

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 06-JUN-99 | factory |  |  |
| Conductivity | 1766 | 06-MAY-99 | “ |  |  |
| Secondary Temp. | 2374 | 05-JUN-99 | “ |  |  |
| Secondary Cond. | 2173 | 29-MAY-99 | “ |  |  |
| Fluorometer | WSIS-097 | 30-JUL-97 | ? |  |  |
| Transmissometer | 192D | 20-JUL-99 | IOS |  |  |
| Pressure Sensor | 75636 | 04-JUN-99 | factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.