## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 9934

Agency: OSAP

Location: Haro Strait

Project: Haro Strait 1999

Chief Scientist: Svein Vagle

Platform: CCGS VECTOR

Date: 6 October 1999 –14 October 1999

Processed by: Germaine Gatien

Date of Processing: 22 November 1999 – 26 November 1999

Number of original CTD casts: 76

Number of casts processed: 75

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#550) was used.

SUMMARY OF QUALITY AND CONCERNS

There were large spikes in the data as has occurred in many recent cruises. There have been problems with these sensors on previous cruises and there continue to be larger differences between pairs of sensors than are usual. There were no bottles for comparison.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.dat.

2. Preliminary Steps

The Log Book and salinity data were obtained. The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. The con file contains information about a transmissometer and fluorometer, but none were actually installed for this cruise.

- Conversion of Raw Data

- All files were converted using 99340001.con as configuration file.

- ALIGNCTD

Based on experience with these sensors during cruise 9920 the primary conductivity channel was advanced by –0.013s and the secondary conductivity advanced by 0.060s. The primary channel was aligned by +0.073 by the acquisition system, so the net effect is that both channels are advanced by +0.060s. A check of a few casts indicates that there were few salinity spikes that were not associated with temperature spikes so that the effect of running ALIGNCTD is not obvious.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure, Temperature and Conductivity.  Parameters used were:

Pass 1    Std Dev = 2		Pass 2    Std Dev = 5		Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity.

8. Test Plots and Channel Check

Casts #1,27,49 and 68 were plotted to check for agreement between the pairs of T and C sensors. No casts were deeper than 310db. The temperature variations are ~0.0025C°, while the conductivity differences are ~0.0007 units and salinity differences ~0.009psu. The descent rates were low (0.5 to 0.8m/s) but were very steady with no reversals noted during the down-casts. The temperature and salinity variations are larger than usually observed, but similar results have been found for these sensors on other cruises. The temperature differences are two or three times higher than those usually seen with SeaBird 911+ sensors.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. Header errors (associated with pressure spikes) had to be fixed for casts #21 and 35 to enable conversion.

10. Checking Headers

A header summary was produced and errors which were found in most headers were fixed.

The surface check was run. The average surface pressure is 0.878db.

The header check was produced and no errors found.

The cruise track was plotted and looks fine.

DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 1.0

Pressure filtered over width: 15

Swells deleted. Warning message if pressure difference of 2.00

Drop rates< 0.1m/s (calculated over 15 points) was deleted.

Sample interval =       .04 seconds.

All warnings were examined and most were found to pertain to the bottom of casts. Warnings for casts #20, 29, 35, 36, 50, 52 and 63 arose because these were towyo casts. The resulting *.del files are pieced together from sections of several downcasts. In order to obtain the single deepest downcast, the converted files (*.ios) were edited by removing data from shallow downcasts that occur before the deepest ones. Details are noted in the Particulars Section.  On rerunning these casts through DELETE only cast #50 produced warnings in the downcast section; those warnings refer to small sections in which the CTD moved upwards. No useful data was removed.

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure		Averaging interval = 1.000

Minimum bin value =   .000	Average value will be used

Interpolated values are NOT used for empty bins

13. Intercomparisons

Previous experience with these sensors – Since they were last calibrated these sensors were used during cruises 9931 and 9920 in August 1999 and during 9940 in late October 1999. There was drift in the sensors during 9931; the secondary sensor was the more stable and was recalibrated by adding 0.008psu. During 9920 the primary sensors were high by 0.005psu and the secondary sensors were low by 0.0035; there was a slight temporal variation in the primary calibration, but not in the secondary. During 9940 the primary sensors were found to be high by about 0.001psu while the secondary was low by 0.013psu. No recalibration was done during 9940 since there were no deep comparisons and very few samples. Primary values showed slight drift but were chosen since the differences were much smaller.

COMPARE – There were no bottle salinities available. The compare routine was used to study the differences between the pairs of sensors using metre-averaged, thinned files. The results fit the observations of section 8 above and show no evidence of variation with time.

Historic ranges – Five casts were plotted with historic ranges superimposed; all fell within those ranges.

Conclusion – The primary sensors will be used since they were closest to bottles during 9940.

14. Test Plots

Profiles of all casts were examined for any evidence of problems with the processing; none were found.

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

15. CTDEDIT

CTDEDIT was used to clean noise in S and/or T near the surface for the following casts: #1,8,9,11,12,14,15,16,31,32,33,34,51,58,59,60,61,62,64,66,67 and 69.

Note was made of the editing details in the relevant files.

- Recalibration

No recalibration was done.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

18. REMOVE

The following channels were removed from all casts: Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag and Scan_Number.

19. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

20 - Towyo cast. First 56263 records removed to select deepest downcast section.

29 - Towyo cast. First 37378 records removed to select deepest downcast section.

35 - Towyo cast. First 136351 records removed to select deepest downcast section.

36 - Towyo cast. First 6034 records removed to select deepest downcast section.

48 - Cast aborted. Only surface records in file. Deleted.

50 - Towyo cast. First 11708 records removed to select deepest downcast section.

52 - Towyo cast. First 17660 records removed to select deepest downcast section.

63 - Towyo cast. First 99577 records removed to select deepest downcast section.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	9934

Dates: 	Start: October 6, 1999  	  End:	October 14 1999

Location: Haro Strait

Vessel:	CCGS VECTOR

Chief Scientist:	 Svein Vagle

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | No | yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0550

Cruise ID#:		9934

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 16-Jun-1999 | Factory |  |  |
| Conductivity | 1766 | 6-May-1999 | “ |  |  |
| Secondary Temp. | 2374 | 5-Jun-1999 | “ |  |  |
| Secondary Cond. | 2173 | 29-May-1999 | “ |  |  |
| Pressure Sensor | 75636 | 04-Jun-1999 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.