## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| ### 23 Nov 2021 | ### Corrected Salinity:Bottle precision. S.H. |
| ### 12 January 2021 | ### Added HPLC Data. S.H. |
| 10 August 2001 | Reprocessed SeaPoint Fl. Channel applying median filter and adding fluorescence offset correction to calibration file. Details in §18. |
| 21 January 2002 | PAR recalibrated to correct calibration constant. Details in §18. |

PROCESSING NOTES

Cruise: 2000-03

Agency: OSAP/UBC

Location:  West Coast Vancouver Island

Project: LaPerouse/Line P

Party Chief: Angelica Peña

Platform: CCGS J.P.Tully

Date: 15 March 2000 –19 March 2000

Processed by: Germaine Gatien

Date of Processing: 10 November 2000 – 16 November 2000

Number of original CTD casts: 21

Number of casts processed: 21

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0550) mounted with a transmissometer, SeaPoint fluorometer, PAR sensor and SeaPoint fluorometer was used.

SUMMARY OF QUALITY AND CONCERNS

The descent rate was extremely noisy for all casts with many reversals of the CTD; while steps were taken to remove obviously bad data, some lowering of data quality is unavoidable.

The fluorometer was unpumped and the data is corrupted by features that reflect variations in the descent rate.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants checked. A PH sensor is mentioned in the calibration file but there was no channel recorded for it.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using conversion file 20030001.con. Rosette files were converted at the same time. An initial examination of the data shows reasonable correspondence between up and down casts and between the two sets of sensors, but there were many spikes in all channels.

4. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

5. ALIGNCTD

No record was kept of the deck unit serial number, so cast #10 was studied to determine if the secondary channel was aligned or not. Applying a variety of alignments shows that the best choice is to apply the same advancement to both channels which suggests that the deck unit advanced both the primary and secondary conductivity by 0.073s in the acquisition process. Fine-tuning the advancement makes very little difference, so the ALIGNCTD step was skipped.

6. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

7. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

8. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and the descent rate. These were placed in a test directory and will not be archived.

9. Test Plots and Channel Check

The three deepest casts (#6,8,13) were plotted to check for agreement between the pairs of T and C sensors and to study descent rate. The descent rates were extremely noisy with many reversals in direction of the CTD. The differences were very noisy and the following values represent average values near the bottom of each cast. These differences are lower than those noted in the previous cruise that used the same sensors, but the casts are deeper so this is to be expected. When the shallow cast #14 was examined it gave results closer to those noted in 2000-02.

| Cast # | Maximum depth | T1-T0 | C1-C0 | S1-S0 |
| --- | --- | --- | --- | --- |
| 6 | 1000 | Noisy Avg -0.0008 | -0.0012 | -0.014 |
| 8 | 900 | Noisy Avg -0.0004 | -0.00098 | -0.011 |
| 13 | 1350 | Noisy Avg -0.0002 | -0.00093 | -0.010 |
| 14 | 300 | Extremely Noisy | -0.0015 | -0.018 |

The chl_fluorescence channel appears to be heavily corrupted by the descent rate noise. The PAR channel trace is much smoother but also shows some evidence of features associated with CTD reversal.

The descent rate was examined for all casts and found to be extremely noisy with many reversals.

10.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

11. Initial comparisons

Previous experience with these sensors – These sensors were used during cruise 2000-02 when the primary pair gave salinities low by 0.0035units and the secondary low by 0.0273units. The secondary conductivity sensor was recalibrated after mission 2000-03.

COMPARE - Comparison with the rosette bottle salinities gives primary salinities low by 0.007units and secondary low by 0.021units. Only bottles below 200m were included in the comparison, but the top 50m was well-mixed for most casts and that data could have been included; it is consistent with the deeper bottles.

Based on these results it was decided to choose the primary channels for archiving.

12. Checking Headers

A header summary was produced and no errors found.

The surface check was run. The average surface pressure is 2.19db.

The header check was produced and no errors were found.

The cruise track was plotted and looks reasonable.

13. Test Plots

Profiles of all casts were produced and no evidence was found of problems with the processing.

- DELETE

REVERSE was run on cast #6 so that the upcast could be processed.

Delete was run on *.IOS files except for cast #6 which was run on *.REV.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS

Most warnings refer to surface and upcasts only; others were associated with CTD reversals.

The data from the SeaPoint fluorometer is still very spiky after DELETE.

All DEL files were copied to EDT files.

15. DETAILED EDITING

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

While DELETE removed many bad data records associated with low descent rate (including many reversals of the CTD) a lot of data corrupted by shed wakes remains. CTDEDIT was used to remove many such bad records and to clean salinity spikes where appropriate. All casts required extensive editing. Note was made of the editing details in the relevant files. The edited files were copied to EDT files.

A study was made of cast #13 to see whether extensive editing below 200m had a significant effect on the data. Comparing metre-averaged data with and without editing showed significant differences all the way to the bottom (1400m). This cast was well-mixed at the surface and the differences are quite small there. As expected the unedited data has higher temperatures and lower salinities for downcasts since the shed wakes generally bring water from a higher depth; however, there were complete reversals of the CTD which complicates things. The differences between edited and unedited metre-averaged data below 1000m are up to 0.03Cº and 0.004units for T and S respectively.

Plots of temperature and salinity vs depth for the BOT files were examined to check for any bad values in these files. Problems were found in casts #13 and 18 and one record removed from each using a text editor; note was made in the headers.

16. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

17. Other Intercomparisons

Historic ranges – Plots were made of the metre-averaged data for casts #6,8 and 13 with historic ranges superimposed. All data fell within the ranges.

Repeat casts - There were many repeat casts at station P4. When plotted on a density surface the range of temperatures and salinities at around 300m was 0.06Cº and 0.01units which seems reasonable since the casts were over a 2-day period and in an area of fairly large gradients.

- Recalibration

The metre-averaged files (from EDT and BOT files) were recalibrated using calibration control file 2003rcal.ccf which raised the primary salinity by 0.007units and the secondary salinity by 0.021units. The recalibrated rosette files were renamed with extensions RAC.

The SAM files were recalibrated and COMPARE was rerun. The average differences from bottles for the primary salinity was +0.00003units and for the secondary, -0.0003units.

On Aug. 10, 2001 the files were reprocessed, applying a median filter (size 11) to the SeaPoint fluorometer channel in the EDT files. The results were recalibrated adding an adjustment to the SeaPoint fluorometer channel to the salinity recalibrations. An offset of 0.045g/l was subtracted. An updated version of 2003rcal.ccf was created. The resulting files were bin-averaged and reordered. The bottle files were also recalibrated.

On January 21, 2002 file 2003PAR.ccf was used to recalibrate the PAR channels in the CHE, RAC and CTD files; the original PAR calibration constant was too low by a factor of 10 producing values too high by a factor of 10.

19. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

20. REMOVE

The following channels were removed from all casts: Scan_Number; Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The Irradiance (PAR) channel was removed from casts #6, 8,10 and 13 only.

21. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

6- Pump not started until half way down (~500m). Use upcast data.

8,13 – PAR channel recorded but all values the same. Channel deleted from final files.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2000-09

Dates: Start: 15 March 2000 End: 19 March 2000

Location: West Coast Vancouver Island

Vessel: CCGS J.P.Tully

Chief Scientist: Angelica Peña

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0550

Deck Unit/Serial #:   unknown

Cruise ID#:		2000-03

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 7Dec1999 | Factory |  |  |
| Conductivity | 1764 | 28Dec1999 | “ |  |  |
| Secondary Temp. | 2371 | 16June1999 | “ |  |  |
| Secondary Cond. | 2102 | 24June1999 | “ |  |  |
| Transmissometer | 192D | 20Jul1999 | IOS |  |  |
| Fluorometer | 2229 | ? | ? |  |  |
| PAR sensor | 4565 | 9Sept1999 |  |  |  |
| Pressure Sensor | 69698 | 12Nov1997 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.