## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 2000-05

Agency: Stock Assessment

Location: NE Pacific

Project: Offshore Hake Survey

Party Chief: Mark Saunders

Platform: CCGS W. E. Ricker

Date: 8 August 2000- 14 August 2000

Processed by: Germaine Gatien

Date of Processing: 8 May 2001 – 1 June 2001

Number of original CTD casts: 12

Number of casts processed: 12

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0506) was mounted with Transmissometer 197. The deck unit was a SeaBird 11+ S/N 11P18377-0471.

SUMMARY OF QUALITY AND CONCERNS

The sampling interval is 1.375s so the data quality is much lower than usual for this instrument.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

No CTD Log was available. A spreadsheet with positions and activities for CTD casts was obtained.

Salinity calibration data was obtained, but without a log the depth of the samples is uncertain. It will be assumed that the deep samples were taken when the CTD was at the bottom of the cast so the bottle salinity is likely 1db above the maximum pressure as this has been the practice on previous PBS cruises.

The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants were checked. There is an error in the pressure calibration. The same error was found during cruise 2000-16. At that time it was noted that the surface pressures seemed too high and correcting the error reduced the pressures. The calibration file 506ctd.con was created with the corrected pressure calibration.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using conversion file 506ctd.con.

An initial examination of the data shows all channels present. The pairs of sensors appear to be in reasonable agreement as do the down and up-casts. However, there is very little data. The sampling interval is 1.375s although the CON file was set at 24samples/second.

4. ALIGNCTD

This step was skipped as there is insufficient data for it to be useful.

5. WILDEDIT

This step was skipped as there is insufficient data for it to be useful. Any bad points will be identified in the editing process.

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run to calculate primary and secondary salinity. A second run of DERIVE was used to calculate differences between the pairs of sensors and to examine the descent rate. This data will not be archived.

8. Test Plots and Channel Check

Four deep casts (# 4,5,7,8) were plotted to check for agreement between the pairs of T and C sensors. All channels were present. The descent rate was fairly noisy, but high enough that it should not cause major problems. The differences in temperature, conductivity and salinity were found to be about 0.0008C, 0.0007units and 0.008units, respectively. These are a little higher than for previous uses of this equipment.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. The GPS headers needed editing for about half the casts. (Latitude, longitude and time entries were in the wrong places.)

10. Checking Headers

A header summary was produced. Times for the first two casts differ from the spreadsheet times by one hour. However they appear to fit the Net Tow Log times if it is assumed that the tows followed the CTD casts, as was the case for most of the casts.

The surface check was run. The average surface pressure is 2.1db, which is reasonable. However, for 5 casts the surface pressure was off-scale. It appears that the pump was not turned on when logging started. Sea-Bird CLEAN plus editing should remove bad data caused by this.

The header check was produced and no errors were found.

The cruise track was plotted and looks reasonable.

11. Test Plots

Profiles of all casts were examined for any evidence of problems with the processing and none were found. Given the sampling rate the up and down-casts are reasonably similar as are the pairs of sensors.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 20.00

Minimum Surface Salinity: 10.0

Surface Swell Pressure Tolerance:  0.5

Pressure NOT filtered

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted.

Sample interval =  1.375 seconds.

The secondary sensors were chosen for further processing since a preliminary check suggests they are closer to the bottles and they were used for cruise 2000-34.

IOS SHELL routine Sea-Bird CLEAN was used to remove records with negative pressures.

All CLN files were copied to EDT files.

- DETAILED EDITING

Page plots were produced using T1,S1 for all casts. These plots were examined for spikes and instabilities. Since the sampling interval was high the plots are quite smooth. CTDEDIT was used only to remove 4 surface records from cast #12. Note was made of the editing in the relevant file. The edited file was copied to an EDT file.

14. BIN AVERAGE

This step was skipped since the sampling interval was 1.375s.

15. Intercomparisons

Previous experience with these sensors –

| Mission | Start Date | Sensors 2278/2710 | Sensors 2280/2095 |
| --- | --- | --- | --- |
| 2004 | 1/4/2000 | 3 bottles only –0.014 | 3 bottles only –0.004 |
| 2009 | 1/5/2000 | Drift from –0.003 to +0.002 at end | +0.0005 |
| 2015 | 31/5/2000 | No reliable info | No reliable info |
| 2016 | 26/6/2000 | Surface only - +0.001 | Surface only –0.003 |
| 2034 | 03/10/2000 | 8 bottles +0.0015 | 8 bottles -0.002 |

COMPARE – There was no rosette. Deep salinity samples were gathered for 2 casts during 2000-05 and 3 during 2000-29. The comparison is shown in file 2005comp.xls. While pressures were recorded for 3 of the 5 bottles, they appear to be wrong. For cast #6 the spreadsheet indicates that the bottle was tripped at 660m whereas the maximum pressure is 501db. It was assumed that the pressures were 1db less than the maximum depth recorded. This is the difference used for previous Stock Assessment cruises. The primary salinity was found to be low by an average of 0.012units and the secondary low by 0.003units. The result for the secondary salinity is reasonably close to those of 2000-16 and 2000-34. The difference between the two salinities is close to the difference found between the primary and secondary salinity in step 8 (0.008units). However, the primary salinity error is larger than for all other cruises since last calibration with the exception of 2000-04 for which there were only 3 samples available. While the pressures may well be wrong, choosing salinities from a shallower depth would only increase the error.

Historic ranges – Historic ranges were found for all casts and all data fell within the ranges.

- Recalibration

No recalibration was done since the secondary CTD salinities are reasonably close to the bottles and the comparison is based on very few bottles; similarly the results of previous cruises using this equipment are based on little information.

- HEADEDIT AND REORDER

A note was added to the comment section of all the headers to advise users of the low sampling rate.

The following channels were removed from all casts: Scan_Number, Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The data format was corrected as needed. The output files had extensions CTD.

18. Final Plots and Cross-Reference Listing

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the files with extensions CTD.

A cross-reference listing was produced.

### Particulars

20050002 – UTC time not updating

20050007 – No upcast due to equipment problems

20050010 – Problems after blackout

20050011 - Spikes on Transmissometer

20050012 – Position in file does not match the bridge log, but does match the Tucker Trawl location. File probably correct.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2000-05

Dates: Start: 8 August 2000 End: 14 August 2000

Location: NE Pacific

Vessel: CCGS W. E. RICKER

Chief Scientist: Mark Saunders

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		2000-05

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 14Dec1999 | Factory |  |  |
| Conductivity | 2278 | 16Dec1999 | “ |  |  |
| Secondary Temp. | 2095 | 16June1999 | “ |  |  |
| Secondary Cond. | 2280 | 16Dec1999 | “ |  |  |
| Transmissometer | 197 | 12Jan2000 | IOS |  |  |
| Pressure Sensor | 69698 | 12Nov1997 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.