## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |
| 8 November, 2001 | Reprocessed SeaPoint FL channel applying median filter and adding offset correction to calibration. RAC files recalibrated. Formats corrected. |

PROCESSING NOTES

Cruise: 2000-06

Agency: OSAP

Location: WCVI and Juan de Fuca

Project: LaPerouse/Effingham/JuandeFuca/Hot Vents/Georgia Strait/CRD

Party Chief: Tom Juhasz

Platform: CCGS John P. Tully

Date: 12 May 2000 –26 May 2000

Processed by: Germaine Gatien

Date of Processing: 17 November 2000 – 1 December 2000

Number of original CTD casts: 111

Number of casts processed: 112 (including 1 upcast)

INSTRUMENT SUMMARY

Two SeaBird Model SBE 911+ CTDs  (#0443 and 0585) were used.

SUMMARY OF QUALITY AND CONCERNS

The descent rate is very noisy for some casts with data lost due to shed wakes.

There are problems with the fluorometer for cast #36 onwards with zero values in the upcasts and frequent zero values in the downcasts.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants were checked. The pressure calibrations for CTD 0585 were not the same as those on file. Investigation proved that the file values were wrong.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using conversion files ctd0443.con for cast #1 and ctd0585.con for all other casts. Many casts had errors in the headers (misplacement of start time, latitude and longitude entries) and these were fixed in both the CNV and ROS files.

An initial examination of the data suggests that the two sets of sensors tracked well and the up and downcasts are reasonably similar. However, the secondary conductivity is very noisy.

4. ALIGNCTD

Examination of cast #27 using different alignments shows that the deck unit must have aligned both channels. Tests were made of the effect of applying further adjustments to both channels and –0.018s improved the smoothness of the salinity in areas of large temperature gradient. So both conductivity channels were advanced  –0.018s and have a net advancement of +0.055s.

Cast #1 used a different CTD but the same alignment was applied.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts (19,61,69,80,107) to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Max Depth | T (C) | C (units) | S (units) | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 19 | 230 | +0.0001 | +0.0004 | +0.004 | Noisy but fairly high |
| 61 | 300 | +0.0004 | +0.0004 | +0.004 | Noisy but fairly high |
| 69 | 850 | ~0 | +0.0004 | +0.004 | Noisy, some reversals |
| 80 | 1600 | ~0 | +0.0004 | +0.006 | Noisy, some reversals |
| 107 | 1750 | ~0 | +0.0004 | +0.005 | Noisy, some reversals |

These are similar to the results during 2000-07 except that temperature differences are lower, which is expected since these casts are deeper than those during 2000-07.

The fluorometer appears to have been unpumped and descent rate variations are reflected in the fluorescence trace.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. Header errors in some casts had to be corrected to enable conversion.

The ROS files were converted to IOS files (after fixing some headers), and the extensions were changed to BOT.

10. Checking Headers

A header summary was produced and a few errors in headers were found and corrected.

The surface check was run. The average surface pressure is 2.2db. A few upcasts were examined to check the pressure calibration and the actual surface as judged by conductivity values of –9.999 reads between +0.3 and -0.3db. No offset of pressure is deemed necessary.

The header check was produced and no errors were found.

The cruise track was plotted and looks reasonable.

11. Test Plots

Profiles of all casts were examined for any evidence of problems with the processing. Cast #68 has bad secondary temperature and conductivity for most of the downcast. Cast #92 has bad primary salinity on the upcast from the bottom to about 500m. The fluorescence channel looks very strange from cast #36 on. Most of the upcasts have zero FL values and many zero values occur in the downcasts as well. Casts #38 and 81 are examples of very bad FL data.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS

Most warnings refer to surface and upcasts only. Mid-depth downcast warnings were investigated for casts #1,61,82,91-95 and 104. Most were related to major slowdowns or reversals of the CTD during descent. There were problems with cast #1 with two mini-casts to about 20m preceding the full cast. There appears to have been a problem with the pumps which were corrected before the full cast. A text editor was used to remove the first 5500 records so that DELETE would select the full cast. DELETE was rerun on the edited cast.  Similarly, for cast #92 the first 2247 records were removed to prevent DELETE from selecting questionable data from several shallow initial lowerings of the CTD.

All DEL files were copied to EDT files.

It was decided to use primary sensors for further processing since the secondary salinity has a lot of fine noise in it and since the secondary signals were bad for cast #68.

- DETAILED EDITING

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

Many casts had unstable features in the top 5 or 10m which are not seen in the upcasts and heavy editing was done to remove such features.

Casts from the Juan de Fuca section contained many deeper unstable features that are not associated with descent rate noise; these appear to be real features and were left unedited.

The casts from the west coast of Vancouver Island contain many instabilities that are associated with shed wakes.

CTDEDIT was used to clean noise in S and T near the surface and bottom for casts: 1,6,9,10,12, 14,16,17,20,24,26,27,29,33,34,44,46,56,62,72,73.

CTDEDIT was used for more extensive cleaning in the following casts: 2-4,7,13,15,19,23,30,32, 35-41,50,51,53-55,57,60,61,63,66-69,75-85,90,92-94,98,101-103,105,107,110,111,9981 (upcast of 81).

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

Plots of temperature and salinity vs depth for the BOT files were examined to check for any bad values in these files. A few points were removed from the primary salinity for cast #91.

REVISION: (Nov. 8,2001) The fluorometer channel was put through a median filter, size 11.

14. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

15. Intercomparisons

Previous experience with these sensors – The sensors for CTD #0443 were used during cruises 2000-01 and 2000-02 with primary sensors found low by 0.0041 and 0.0039units respectively. The secondary sensors were found to be low by 0.006 and 0.0047units. There were deep bottles available from 2000-02 but not from 2000-01 so the figure of –0.0047 is probably more accurate.

The sensors for CTD #0585 were used during cruise 2007 when a comparison using only 7 bottles found the primary sensors to give salinity low by 0.0083 and the secondary, low by 0.0036. The secondary sensors were found to contain a lot of fine-scale noise.

COMPARE – Comparison is only possible for CTD #0585 since there was only one shallow cast using CTD#0443 (cast #1). The rosette bottle salinities show that the primary sensors gave values low by about 0.015units and the secondary, low by about 0.01units. Note that while cast #1 was included in the file list, none of the data was used in the comparison.

Historic ranges - Plots were examined with historic ranges superimposed for casts #69,80,81,82,107 and all data fell within the ranges.

- Recalibration

File 2010rcal.ccf was used:

a) to raise the primary salinity by 0.004units and the secondary by 0.005units for cast #1 only (CTD #0443).

b) to raise the primary salinity by 0.015units and the secondary by 0.010units for all other casts (CTD #0585).

COMPARE was rerun after calibration and the average differences were –0.00002units for the primary salinity and +0.00025units for the secondary.

REVISION: (Nov. 8, 2001) 2006rcal.ccf was edited adding an offset to the fluorescence channel (decreasing the fluorescence by 0.04g/l, the “dark value”.)

FL was recalibrated in the CHE files using 2006fcal.ccf.

- REORDER

REORDER was run to assign the number of decimal points in all variables to reflect their accuracy and to remove channels as follows:

From all casts - Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

From casts #70-111 - The PAR channel.

REVISION: (Nov. 8,2001) REMOVE was used to remove the channels mentioned above and HEAD EDIT was used to fix the formats.

18. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

1. Because of pump problems with the CTD, there were two down and upcasts between 0 and 20m before a complete cast was done. A text editor was used to remove the first two lowerings so that DELETE would choose the best data.

29. The time in the CTD log is different from that in the headers. A note in the log suggests that the log time is incorrect.

68. Bad T1, C1 for most of downcast.

92. Bad S0 signal for most of upcast to 500m.

92. A text editor was used to remove several initial shallow lowerings of the CTD so that DELETE would choose the best data.

36-111. Problems with chl_fluorescence signal especially in upcasts. Many zero values in downcasts too.

70-111 – no PAR sensor

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2000-06

Dates: Start: 12 May 2000 End: 26 May 2000

Location: LaPerouse/Effingham/JdeF/Hot Vents/Georgia Strait/CRD

Vessel: CCGS John P. Tully

Chief Scientist: Tom Juhasz

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | yes | yes |
| 2 | SEABIRD | 911+ | 0550 | Yes | yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0443

Cruise ID#:		2000-06

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 07May1999 | Factory |  |  |
| Conductivity | 2128 | 24June1999 | “ |  |  |
| Secondary Temp. | 2449 | 27July1999 | “ |  |  |
| Secondary Cond. | 1763 | 28July1999 | “ |  |  |
| Transmissometer | 192D | 20July1999 | IOS |  |  |
| Fluorometer | 2229 |  | ? |  |  |
| PAR Sensor | 4565 |  | ? |  |  |
| Pressure Sensor | 63507 | 11Jan1996 | Factory |  |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2000-06

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2663 | 07Mar2000 | Factory |  |  |
| Conductivity | 2399 | 03Mar2000 | “ |  |  |
| Secondary Temp. | 2668 | 07Mar2000 | “ |  |  |
| Secondary Cond. | 2424 | 03Mar2000 | “ |  |  |
| Transmissometer | 192D | 20July1999 | IOS |  |  |
| Fluorometer | 2229 |  | ? |  |  |
| PAR Sensor | 4565 |  | ? |  |  |
| Pressure Sensor | 77511 | 30SEP1999 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.