## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |
| 19-Jan-2004 | When moving files to new archive a couple of things were noted. 1) The CTD files did not contain the Dec 5, 2000 change. I’m not sure why but when Germaine redid the fluorometer data we must of copied files from the previous version. 2) The CHE flies from were missing from my M: drive (probably copied them there so to hide them) so I copied the files from Bernard’s disk and reprocessed them since Bernard’s CHE files had the phosphate data to one decimal place only. |
| 16 November 2001 | Reprocessed SeaPoint FL channel by applying median filter and adding offset correction to calibration. RAC files recalibrated. Formats corrected. G.Gatien |
| 5-Dec-2000 | In processing 2000-06 Germaine discovered an error in the pressure calibrations entered in the Seabird calibration file on the network before doing the conversion so there is no problem with that data. However, the wrong pressure calibration was used for 2000-07. She calculated that the effect was that the pressures in the archive are too high by about 1db. Therefore I subtracted 1 db from both the *.CTD and *.CHE files. Joe Linguanti |

PROCESSING NOTES

Cruise: 2000-07

Agency: IOS/OSAP

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG / JdeF

Party Chief: Diane Masson

Platform: CCGS John P. Tully

Date: 25 April 2000 –29 April 2000

Processed by: Germaine Gatien

Date of Processing: 20 June 2000 –14 June 2000

Number of original CTD casts: 74

Number of casts processed: 73

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (S/N 0585) was used.

SUMMARY OF QUALITY AND CONCERNS

The salinity is considered to be within +/- 0.01units. This cruise was subject to heavy weather particularly when the deepest bottle sampling was done. Given the great variability in this region the effect of shed wakes is expected to be large and thus the quality of the bottle comparison is reduced. When data from other cruises using these sensors has been analyzed it may be wise to revisit the calibration issue.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants were checked. An error was found in the pressure calibration in the configuration file. A configuration file with proper calibrations was prepared (2007ctd.con).

- Conversion of Raw Data

The raw data was converted using conversion file 2007ctd.con.

An initial examination of the data suggests that the secondary conductivity is very noisy.

The salinity channel was stripped from the CNV files. (Otherwise when DERIVE is run a 2nd set of salinity channels will be created.)

4. ALIGNCTD

Examination of cast #14 using different alignments shows that the best choice is to advance the conductivity by 0.055s. The primary conductivity was advanced by 0.073s by the acquisition system, so ALIGNCTD was set to advance the primary by –0.018s and the secondary by +0.055s, in order that both channels have a net advancement of +0.055s.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure, Temperature, Conductivity, Transmissivity and Fluorometry.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate..

8. Test Plots and Channel Check

A sample of casts (#4,10,13,52, 65 and 74) was plotted to check for agreement between the pairs of T and C sensors. The deepest of these casts was 416m and it is an area of rapid change. The descent rate was very noisy for some of these casts. All of the differences were extremely noisy with average values of about 0.002, 0.0005 and 0.005 for temperature, conductivity and salinity, respectively. The primary channel gave higher values for temperature and lower values for conductivity and salinity.

The fluorometry channel was investigated. Values for Saanich Inlet look reasonable. Where there are sharp minima in the descent rate these are reflected in positive spikes in the fluorometer channel.

- Conversion to IOS Headers

The headers for 55 casts (and 14 corresponding rosette files) had to be edited to enable conversion. The entries for NMEA positions and times were scrambled as has been noted previously in casts in which there are pressure spikes. A few casts had no positions in the headers at all so these were entered manually in the IOS files using the positions given in the CTD log. For one cast there were words entered in the lines where the NMEA positions should have been. These lines had to be removed from the CNV files and after conversion to IOS headers the position had to be entered manually using the CTD log positions.

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

10. Checking Headers

A header summary was produced and one small header error was found and corrected.

The surface check was run. The average surface pressure is 3.4db.

The header check was produced and no errors were found.

The cruise track was plotted and looks reasonable.

11. Test Plots

Profiles of all casts were examined and no evidence of problems with the processing was found.

- DELETE

- The following DELETE parameters were used:

- Surface Record Removal: Last Press Min

- Maximum Surface Pressure (relative): 20.00

- Surface Swell Pressure Tolerance: 1.0

- Swells deleted. Warning message if pressure difference of 2.00

- Drop rates < 0.3m/s (calculated over 15 points) was deleted.

- Sample interval =  .04 seconds.

- COMMENTS ON WARNINGS

- Most warnings refer to surface and upcasts only. Mid-depth downcast warning was investigated for casts #1,10,13,33,51,52,53,54,56. All are associated with reversals of the CTD and no useful data appears to have been lost.

- All DEL files were copied to EDT files.

- DETAILED EDITING

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Some instabilities were left since such areas as Haro Strait do have such features.

For casts 9-14, 52-54 and 56 the descent rate was very noisy causing many shed wakes. While DELETE removes much of the bad data, there is a lot left that must be removed using CTDEDIT. Data records were removed where corruption by shed wakes was obvious.

CTDEDIT was used to clean noise in S and T near the surface and bottom for cast #20, 23-28, 36, 37, 39, 40, 42-45, 49 and 50.

CTDEDIT was used for more extensive cleaning in the following casts: 1, 3, 4, 8, 9-14, 32, 33, 35, 48, 52-57, 60, 62-64, 66-69 and 71-73.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

Plots of temperature and salinity vs depth for the BOT files were examined to check for any bad values in these files. Because of the large variability of temperature and salinity in this area and because the descent rate was very noisy for many casts, there is a lot of noise in the rosette files. Data from shed wakes was not removed since it is assumed that the bottle would also have sampled the shed wake. No obviously spurious data was found in the BOT files.

REVISION: (Nov.16,2001) The fluorometer channel was put through a median filter, size 11.

14. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

15. Intercomparisons

Previous experience with these sensors – These are new sensors and there is no information available on previous use.

COMPARE – Comparison with the rosette bottle show a lot of scatter. There are no comparisons below 400db. Averaging the 7 deepest values indicates that the primary salinity is low by 0.0083units and the secondary is low by 0.0036units.

Historic ranges - Plots were examined with historic ranges superimposed for casts #12, 13, 31 and 45. All data was within the ranges, but they are very broad at these locations.

- Recalibration

File 2007rcal.ccf was used to raise the primary salinity by 0.0083units and the secondary by 0.0036units. It was applied to the averaged files (both rosette and CTD files).

REVISION: (Nov.16,2001) 2007rcal.ccf was edited adding an offset to the fluorescence channel (decreasing the fluroescence by 0.07μg/l, the “dark value”.) This was applied to the metre-averaged CTD files.   2007fcal.ccf was created to recalibrate fluorescence only in the RAC files.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

- REMOVE

The following channels were removed from all casts: Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

Cast 1 – Downcast only

Cast 2 – Upcast only

Cast 6 – Didn’t start recording until 18db.

Casts 9-14 – Heavily edited to remove data corrupted by shed wakes.

Cast 14 – First record at 14db.

Casts 52-54 - Heavily edited to remove data corrupted by shed wakes.

Cast 56 - Heavily edited to remove data corrupted by shed wakes.

Cast 69 – Bad section around 125db with large reversal of CTD.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2000-07

Dates: Start: 25 April 2000 End: 29 April 2000

Location: Strait of Georgia, Juan de Fuca Strait

Vessel: CCGS VECTOR

Chief Scientist: Diane Masson

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | yes | yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2000-07

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2663 | 07May2000 | Factory |  |  |
| Conductivity | 2399 | 03May2000 | “ |  |  |
| Secondary Temp. | 2668 | 07May2000 | “ |  |  |
| Secondary Cond. | 2424 | 03May2000 | “ |  |  |
| Transmissometer | 333DR | 06Mar2000 | IOS |  |  |
| Fluorometer | 2229 | ? | ? |  |  |
| Pressure Sensor | 77511 | 30Sep2000 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.