## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 26-Feb-2004 | Original nutrient files obtained from Bernard Minkley’s data disk. |

PROCESSING NOTES

Cruise: 2000-09

Agency: OSAP

Location: B.C. Inlets

Project: Inlets

Party Chief: Dario Stucchi

Platform: CCGS VECTOR

Date: 1 May 2000 –13 May 2000

Processed by: Germaine Gatien

Date of Processing: 2 November 2000 – 10 November 2000

Number of original CTD casts: 110

Number of casts processed: 109

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0506) mounted with a transmissometer and SeaPoint fluorometer and deck unit (11P 14262-0425) were used.

SUMMARY OF QUALITY AND CONCERNS

The data was in good order.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants checked. An error in the pressure calibration (same as error found for 2000-04) was corrected and the new configuration file was named 2009ctd.con.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using conversion file 2009ctd.con. Rosette files were converted at the same time. An initial examination of the data shows good correspondence between up and down casts and between the two sets of sensors. There were misplaced entries in the NMEA lat/long/time section of the headers for casts #9-61, 86-97; these were corrected in the ROS and CNV files. These errors are of a type noted in many previous data sets and are associated with pressure spikes.

4. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

5. ALIGNCTD

Examination of cast #40 using different alignments shows that the best choice is to advance the conductivity by 0.050s (0.060s worked almost as well at removing salinity spikes in areas of large temperature gradient). For this model deck unit the primary conductivity is advanced by 0.073s in the acquisition process, but the secondary is not aligned. So ALIGNCTD was set to advance the primary by  –0.023s and the secondary by +0.050s, so that both channels have a net advancement of +0.065s.

6. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

7. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

8. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on a selection of casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and the descent rate. These were placed in a test directory and will not be archived.

9. Test Plots and Channel Check

A sample of casts (#1,11,47 and 112) was plotted to check for agreement between the pairs of T and C sensors and to study descent rate. The descent rates were generally high enough and very steady.

| Cast # | Maximum depth | T1-T0 | C1-C0 | S1-S0 |
| --- | --- | --- | --- | --- |
| 1 | 216 | <0.0005 | +0.0005 | +0.005 |
| 11 | 600 | -0.0008 | +0.0002 | +0.0025 |
| 47 | 600 | -0.0004 | +0.0001 | +0.0015 |
| 112 | 360 | -0.0002 | -0.0001 | -0.001 |

The differences in conductivity appear to be decreasing through the cruise as do the salinity differences.

10.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

11. Initial comparisons

Previous experience with these sensors – These sensors were used during cruise 2000-04 (April), 2000-15 (early June) and 2000-16 (June/July). There was no reliable salinity comparison from 2000-15 and there was a significant change in the calibration of the primary sensors between 2000-04 and 2000-16. For 2000-04 the primary salinity was low by 0.014units and the secondary low by 0.004units. For 2000-16 the primary was high by 0.001units and the secondary low by 0.003units.

COMPARE - Comparison with the rosette bottle salinities shows that the primary sensors gave values low by 0.001units and the secondary high by 0.0005units. There was a slight pressure dependence for the primary and virtually none for the secondary. Because there was a notable difference between the comparisons done for the primary sensors for cruises 2000-04 and 2000-16, COMPARE was used to analyze the differences between bottles and CTD versus the file pair number, to see if there was any time dependence for the calibrations. The secondary sensors produce quite flat results, but there is variation with time for the primary sensors with differences on the order of –0.003units at the beginning of the cruise and on the order of +0.002units at the end.

Based on these results it was decided to choose the secondary channels for archiving.

12. Checking Headers

A header summary was produced and two small header errors were found and corrected.

The surface check was run. The average surface pressure is 1.76db.

The header check was produced and no errors were found.

The cruise track was plotted and looks reasonable.

13. Test Plots

Profiles of all casts were examined on-screen and no evidence was found of problems with the processing.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS

All warnings refer to surface and upcasts only.

All DEL files were copied to EDT files.

15. DETAILED EDITING

Page plots were produced using T1,S1 for all casts. (Cast #12 would not plot until the first line was removed from the IOS file and DELETE rerun.) These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

For many casts there were unstable features in the top 5 or 10m. After metre-averaging most of these features will disappear and some were removed using CTDEDIT since they appeared to be the result of propeller wash or shed wakes. Others were left since many of these stations were in inlets where there are real instabilities.

CTDEDIT was used to clean noise in S and/or T near the surface and bottom for casts: 1-3,5-7,9,13,15-17, 22,38,46-49,56,57,59-61,63,69,78,79,87,94,100,101,104,107,108,112,113,115,116.

CTDEDIT was used for more extensive cleaning in the following casts: 10-12,18,20,21,23-34,36,37,39, 40,43,51,83.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

Plots of temperature and salinity vs depth for the BOT files were examined to check for any bad values in these files. Problems were found in casts # 42, 70 and 109. CTDEDIT was used to remove bad records from these casts and note was made in the headers of the editing. The edited files were renamed BOT.

16. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

17. Other Intercomparisons

Historic ranges – None were available for this data set.

- Recalibration

No recalibration was done since the secondary salinity was within 0.0005units of the bottles.

19. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

20. REMOVE

The following channels were removed from all casts: Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

21. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

8. This was an upcast only so was not processed.

12. First data line removed from the *.IOS file to enable plotting.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2000-09

Dates: Start: 1 May 2000 End: 13 May 2000

Location: B.C. Inlets

Vessel: CCGS VECTOR

Chief Scientist: Dario Stucchi

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Deck Unit/Serial #:   11P 14262 –0425

Cruise ID#:		2000-09

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 14Dec1999 | Factory |  |  |
| Conductivity | 2278 | 16Dec1999 | “ |  |  |
| Secondary Temp. | 2095 | 16June1999 | “ |  |  |
| Secondary Cond. | 2280 | 16Dec1999 | “ |  |  |
| Transmissometer | CST333DR | 6Mar2000 | IOS |  |  |
| Fluorometer | 2229 | ? | ? |  |  |
| Pressure Sensor | 69698 | 12Nov1997 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.

A second CTD (#0550) was on board but was not used.