PROCESSING NOTES

Cruise: 2000-11

Agency: OSAP / UBC

Project: Zooplankton Inlet Monitoring, COPRA

Chief Scientist: Steve Romaine

Platform: CCGS VECTOR

Date:  May 20,2000 – May 26, 2000

Processed by: Germaine Gatien

Date of Processing: 30 November 2001 – 3 December 2001

Number of original CTD casts: 28

Number of casts processed: 23

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD  (S/N#1294) was used.

SUMMARY OF QUALITY AND CONCERNS

There are no salinity calibration samples.

The descent rate was mostly high and steady and the data quality is generally good.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

The Log sheets were obtained; there was no mention of any problems with the CTD.

There were 2 con files available: 2127NEW.con and COPRA.con. The former contains the most up-to-date calibration coefficients and was used to process 2001-27. There are no dates in the COPRA.con file and no information on file that corresponds to the calibrations

- Conversion of Raw Data

- The raw data was converted using conversion file 2127NEW.con.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature. The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

5. ALIGNCTD

Temperature was advanced relative to pressure by 0.9s using ALIGNCTD. This value was chosen as it produced the best overall results in spike reduction on a test cast although a range of values from 0.7s to 1.2s produce the best results on particular spikes. In previous use of this instrument 0.7s was used.

6. DERIVE

Program DERIVE was run to calculate salinity. DERIVE was run a second time to calculate the descent rate; these files were written to a test directory and will not be archived. The descent rate was checked and is generally very steady with 1 sharp minima only in each of casts #48 and 67.

7.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

8. Checking Headers

There were many errors in the headers; the hemisphere was missing from positions and all the times are 00:00 UTC and all dates are May 22, 2020. The times look ok in the SeaBird headers. These errors were corrected by editing the *.IOS files. The start time in the CTD log for cast #72 is clearly wrong; it is later than the end time and the start time of the cast that follows. It is assumed that it should be 16:59 rather than 17:59.

After fixing the times the header check was run and the speeds look reasonable.

A header summary was produced and no further errors found.

The surface check was run. The average surface pressure is –0.24db. No adjustment was made.

9. Test Plots

Profiles were plotted for all casts and up and down traces compared. There were no unreasonable discrepancies noted.

10. DELETE

The following DELETE parameters were used:

Surface Record Removal:  Low Salt & Last Press Min

Minimum Salinity: 15

Maximum Surface Pressure (relative):      10.00

Surface Swell Pressure Tolerance:        1.0

Pressure filtered over width:   5

Swells deleted. Warning message if pressure difference of      2.00

Drop Rates <0.2m/s (calculated over 5 points) was deleted.

11. Test Plots

Page plots were produced and were examined for spikes and instabilities to guide the use of CTDEDIT.

12. CTDEDIT

CTDEDIT was used to remove surface records and/or clean salinity for casts #:18-27,43-58,67-73. Generally the editing of salinity was light; heavier editing was required for cast #73. The details are entered as comments in the headers of the files and are noted on the test page plots at the end of the processing binder.

13. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

- Intercomparisons

COMPARE - There were no bottles available for comparison.

Previous Use of CTD – There were no useful bottles available for cruises 1999-19, 1998-08 and 1997-48 which used this equipment. Cruise 2001-27 in October 2001 found that the CTD salinity was low by about 0.001units, within the noise level of the instrument.

Historic ranges – Profiles of temperature and salinity were plotted with the historic ranges superimposed for six casts. The temperatures were high near the surface for casts #70, 71 and 72. The salinity was low near the surface for casts #70, 71 and 72 and at depth for casts #72 and 73. Dario Stucchi checked other profiles close to stations #72 and 73 and found deep salinities comparable to the ones found during this cruise.

Other comparisons – Cast #54 included two lowerings of the CTD. A comparison was made between the two casts to check for repeatability. In the region of 35 to 50db the differences were on the order of 0.01Cº and 0.01units of salinity.

15. Final Plots

THIN and DERIVED QUANTITIES routines were used to produce information for tables. Page plots were prepared using the edited data and tables.

16. REMOVE and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity and Flag.  The pressure format was corrected.

17. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

Casts #1,4,6,11,13 – Data missing

Cast #72,73 – Deep salinity lower than historic ranges.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2000-11

Dates: 	Start:	20 May 2000	    End:	26 May 2000

Location:	Strait of Georgia and surrounding inlets

Vessel:	CCGS VECTOR

Chief Scientist:		Steve Romaine

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19 SEACAT | 1294 | No | Yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/19 SEACAT / 1294

Cruise ID#:		2000-11

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 1294 | 17 June 1999 | Factory |  |  |
| Conductivity | 1294 | 17 June 1999 | Factory |  |  |
| Pressure Sensor | 163223 | 17 June 1999 | factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.