## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 10-Jan-2006 | Surface loop data was added to the archive. The data was acquired from John Morris at PBS. The data archived in August 2005 was replaced with this new set since chlorophyll was included. The original spreadsheet file from John and more detailed processing notes can be found in the “Cruise_Data\Documents” directory. Any questions regarding this data should be directed to John Morris. J.L. |
| Aug 3, 2005 | Added loop data to the archive. J.L. |

PROCESSING NOTES

Cruise: 2000-16

Agency: OSAP

Location: North East Pacific

Project: High Seas Salmon

Party Chief: David Welch

Platform: CCGS W. E. Ricker

Date: 26 June 2000 –7 July 2000

Processed by: Germaine Gatien

Date of Processing: 17 October 2000 – 26 October 2000

Number of original CTD casts: 44

Number of casts processed: 43

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0506) was used.

SUMMARY OF QUALITY AND CONCERNS

The data was averaged over 10 scans at the time of logging. This will reduce the quality of the data.

Cast #25 is from upcast data.

The only salinity calibrations are from the surface.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained. Only surface samples were taken.

Five different versions of CON files were used at sea. The differences were examined and a new file (2016ctd.con) was prepared that corrects a pressure calibration constant error, and otherwise resembles the con file used for most of the casts in acquisition.

The cruise summary sheet was completed.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using conversion file 2016ctd.con for all casts except cast #11.

The deck unit averaged 10 scans for all casts except  #1.

Cast #11 could not be converted using 2016ctd.con. On advice from Sea-Bird the configuration was changed to set “LAT/LON added by NMEA” to NO. Cast #11 was converted using 2016ctd1.con.

An initial examination of the data shows that:

• cast #1 contains very large spikes and has bad secondary conductivity and the up and down traces are quite different;

• cast #25 upcast and downcast are very different from about 40 to 65m;

• all other casts show good correspondence between sensors and between the up and down traces.

4. ALIGNCTD

For cruise 2000-04 the best alignment was achieved by advancing both sensors by 0.060s. Examination of cast #22 confirms that 0.060s is a good choice for this data set too. The primary conductivity was advanced by 0.073s by the acquisition system so ALIGNCTD was set to advance the primary by –0.013s and the secondary by +0.060s, so that both channels have a net advancement of +0.060s.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure, T0, T1, C0,C1 and Tr.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of casts (#3,24,27,42) was plotted to check for agreement between the pairs of T and C sensors. The differences were very noisy but most values were close to –0.001Cº for temperature, -0.0003units for conductivity and –0.003units for salinity. These values are very different from those found for 2000-04 which used the same sensors.

The descent rate was examined and was found to be very noisy. However, since the rate was fairly high there should not be major problems with shed wakes. Casts with the most serious descent rate problems were 1,3,6,7,12,13,23-27, 29,38,40-43 with 42 and 43 being the worst. Since the deck unit averaged over 10 scans the noise problem may be worse than it appears.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. Header errors in casts 1-4 were fixed (there were numbers where the hemisphere should be in Lat/Long).

10. Checking Headers

A header summary was produced and errors in time were corrected for casts #1-4 and the station names were corrected in casts #1,11,12,27 and 28. There are many differences between positions given in the CTD log and those recorded in the headers. A note in the CTD log indicates that the positions in the headers should be used since there were problems with the system being used for the log entries.

The surface check was run. The average surface pressure is 2.7db. It is noted in the CTD log that the pressures at the surface seemed high by 1-2m. The error found in the pressure calibration accounts for about 0.7db. For the upcasts the minimum pressures were typically 1.5db suggesting that any error is less than 1db since the instrument logging is stopped while the CTD is in the water..

The header check was produced and no errors were found.

The cruise track was plotted and looks reasonable.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0

Swells deleted. Warning message if pressure difference of 2.00

Pressure filtered over 15 points for cast #1 and 5 points for casts #2-44 since the deck unit averaged over 10 scans.

The Drop rates feature was not used since the deck unit averaged over 10scans.

COMMENTS ON WARNINGS

Most warnings refer to surface and upcasts only. Mid-depth downcast warning were investigated for a few casts and were found to correspond to reversals of the CTD.

All DEL files were copied to EDT files.

12. Initial Intercomparisons

Previous experience with these sensors – These sensors were used during cruise 2000-04 in April 2000 during which cruise 3 deep bottle samples were taken. The primary and secondary salinity were found to be low by 0.014units and 0.004units respectively.

COMPARE – Only surface bottles were available. 7 casts with well-mixed surface water were selected and S0 and S1 from the CTD were recorded at the shallowest points available from the downcast (rejecting points for which the CTD does not seem equilibrated). The results show that S0 is high by about 0.001units and S1 low by about 0.003units. Since the CTD data points are from 1.7 to 2.8db and it is to be expected that the CTD salinities are both a little lower nearer the surface, the error in S1 is likely to be greater than 0.003. It would seem wise to choose the primary salinity channel.

- DETAILED EDITING

Since the primary values of salinity appear to be closer to the bottles the primary sensors will be selected for further processing.

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

CTDEDIT was used to clean salinity lightly in casts #2-11, 17, 20, 24, 29 and 32. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

Cast #1 contained many large instabilities and the up and down casts are quite different even at the bottom of the cast. The secondary channels contain bad data and there were many pressure spikes. It is probable that this data is not good, so the cast was deleted.

Cast #25 has an area of unstable features in both salinity channels in the downcasts from about 40 to 65m. While both conductivity channels are noisy they are quite different from each other. The temperatures are quite smooth in this area. The upcast looks very different, no sign of this large unstable feature. The upcast will be used for this cast.

14. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

15. Further Intercomparisons

Historic ranges – Cast #25 was plotted with historic ranges superimposed and all data fell within those limits. For other deep casts only temperature ranges were available and all data fell within the ranges for those.

- Recalibration

There is insufficient evidence to warrant recalibration. The results of cruise 2000-04 are not useful since the differences between the pairs of sensors has changed notably betweeen the two cruises.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

- REMOVE

The following channels were removed from all casts: Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

- Cast deleted because data does not look reliable.

11. This cast was logged using WINDOWS software, contains many pressure spikes and could only be converted by turning the LAT/LON added by NMEA off in the CON file.

25.  Upcast used since downcast has an area of unstable features in salinity.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2000-16

Dates: Start: 26 June 2000 End: 7 July 2000

Location: High Seas Salmon

Vessel: CCGS W. E. RICKER

Chief Scientist: David Welch

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Deck Unit/ Serial #:   11P18377-0471

Cruise ID#:		2000-16

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 14Dec1999 | Factory |  |  |
| Conductivity | 2278 | 16Dec2000 | “ |  |  |
| Secondary Temp. | 2095 | 16June1999 | “ |  |  |
| Secondary Cond. | 2280 | 16Dec1999 | “ |  |  |
| Transmissometer | 197 | 12Jan2000 | IOS |  |  |
| Pressure Sensor | 69698 | 12Nov1997 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.