## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |
| 12-Feb-2006 | Changed mission number from 2000-06 to 2000-28 in CTD files. |
| 24 March 2002 | Fl channel was filtered and recalibrated in the CTD files. FL channel was recalibrated in the RAC & CHE files. For details see §15 & §18. |

PROCESSING NOTES

Cruise: 2000-28

Agency: OSAP

Location: WCVI and Juan de Fuca

Project: La Perouse etc.

Party Chief: Tom Juhasz

Platform: CCGS John P. Tully

Date: 5 October 2000 – 15 October 2000

Processed by: Germaine Gatien

Date of Processing: 1 December 2000 – 7 December 2000

Number of original CTD casts: 79

Number of casts processed: 79 (plus 31 upcasts prepared for Rick Thomson)

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTDs  (#0585) was used. The deck unit was SBE model 11plus, S/N 0508. A transmissometer and fluorometer were attached to the CTD.

SUMMARY OF QUALITY AND CONCERNS

The data was found to be in good order, with none of the pressure spiking and header problems that have been troublesome over the past 2 seasons.

The fluorometer data for casts #1-6 is considered of doubtful quality and should be used with caution.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained.

The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants were checked.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using conversion file 20280001.con.

An initial examination of the data suggests that the two sets of sensors tracked well and the up and downcasts are reasonably similar. However, the secondary conductivity has some fine-scale noise.

4. ALIGNCTD

The deck unit is one of the newer versions that advances both primary and secondary conductivity by 0.073s. Tests were made of the effect of further advancement and using –0.018s improved the data, using -0.023s was almost as good, but –0.028s and –0.013s did not work as well. Thus –0.018s was applied to all casts.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

8. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to check the descent rate. These were placed in a test directory and will not be archived.

9. Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Max Depth | T (C) | C (units) | S (units) | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 10 | 779 | -.0002 noisy | +0.0003 | +0.0045 | Noisy but high |
| 14 | 363 | -.0002 noisy | +0.0003 | +0.0035 | Mostly good, few bad spots |
| 66 | 259 | -.0002 noisy | +0.0003 | +0.0045 | Noisy but fairly high |
| 74 | 352 | Very noisy | +0.0003 | +0.005 | Noisy |

These are similar to the results during 2000-06, 2000-07 and 2000-24 for C and S. Since the secondary temperature sensor was changed it is not surprising that T is also different.

The fluorometer data is extremely noisy for casts 1,2,4 and 5. There is a note in the log that the flow rate to the sensor was adjusted for cast #7 and the trace looks better thereafter though still very noisy. A note has been made in the headers of the suspect casts.

10.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files and the extensions were changed to BOT.

11. Checking Headers

A header summary was produced and no errors found.

The surface check was run. The average surface pressure is 1.94db.

The header check was produced and no errors were found.

The cruise track was plotted and looks reasonable.

12. Test Plots

Profiles of all casts were examined on-screen for any evidence of problems with the processing.

- REVERSE

Casts #42 – 65 and #75-81 were put through REVERSE and stored in a different directory for Rick Thomson. They were put through DELETE and metre-averaged and recalibrated. After calculation of theta they were transferred to Rick Thomson.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0

Pressure was filtered over width: 25

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS

Most warnings refer to surface and upcasts only. Mid-depth downcast warnings were investigated for casts #1,17,38 and 67 and were found to be related to major slowdowns or reversals of the CTD during descent.

All DEL files were copied to EDT files.

It was decided to use primary sensors for further processing since the secondary conductivity contains a lot of fine-scale noise.

- DETAILED EDITING

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

CTDEDIT was used to clean noise in S and T near the surface and bottom for casts: 23,25-27, 29,32,33,43,45,47-50,53,54,57-59,61,63,81.

CTDEDIT was used for more extensive cleaning in the following casts: 1,4,7,8,10,12-15,17-19, 21,22,24,28,30,31,34-41,66-68,70-74,77,79,80.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

Plots of temperature and salinity vs depth for the BOT files were examined to check for any bad values in these files. A few salinity spikes were cleaned in casts #22, 27 and 37 and note was made in the headers of the editing. It was noted that there is a lot of variability in the salinity near the surface for many casts, particularly casts #13-15 and 21-35. The salinity range was as high as 3units in surface samples, but there was no obvious way to edit the data.

REVISION – March 2002 – The fluorescence data were put through a median filter, size 11.

16. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

17. Intercomparisons

Previous experience with these sensors –

The primary sensors were used during cruise 2000-07 when a comparison using only 7 bottles found the primary sensors to give salinity low by 0.0083 and during 2000-06 when they were found to be low by 0.015units. There were lots of deep bottles available for the 2000-06 comparison.

The secondary conductivity sensor was used during 2000-07 and 2000-06 but with a different temperature sensor and the secondary salinity was found to be low by 0.0036units and 0.01units, respectively. The secondary salinity had a lot of fine-scale noise during those cruises.

COMPARE – Comparison with rosette bottle salinities indicates that the primary salinities are low by 0.021units and the secondary, by 0.017units.

Historic ranges - Plots were examined with historic ranges superimposed for casts #10 and 74 and all data fell within the ranges.

- Recalibration

File 2028rcal.ccf was used to raise the primary salinity by 0.021units the secondary by 0.017units.

COMPARE was rerun after calibration and the average differences were 0.003units for the primary salinity and 0.002units for the secondary.

REVISION – March 2002 – File 2028fl.ccf was used to lower the fluorescence by 0.06μg/l in the  CTD, RAC and CHE files.

- REMOVE and REORDER

REMOVE and REORDER were run to assign the number of decimal points in all variables to reflect their precision and to remove the following channels from all CTD casts: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag

20. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

- Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

1-6 – Some doubt about flow to fluorometer.

39 – Pumps stopped before last bottle fired.

77 – Logging not started until 45m. Upcast was substituted for downcast.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2000-26

Dates: Start: 5 October 2000 End: 15 October 2000

Location: LaPerouse/Effingham/JdeF/Hot Vents/Georgia Strait/CRD

Vessel: CCGS John P. Tully

Chief Scientist: Tom Juhasz

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | yes | yes |
| 2 | SEABIRD | 911+ | 0550 | No** | no |

** The second CTD was not used.

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2000-28

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2663 | 07Mar2000 | Factory |  |  |
| Conductivity | 2399 | 03Mar2000 | “ |  |  |
| Secondary Temp. | 2371 | 16June1999 | “ |  |  |
| Secondary Cond. | 2424 | 03Mar2000 | “ |  |  |
| Transmissometer | 192D | 20July1999 | IOS |  |  |
| Fluorometer | 2229 |  | ? |  |  |
| Pressure Sensor | 77511 | 30SEP1999 | Factory |  |  |