## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 2000-29

Agency: Stock Assessment

Location: Georgia Stait

Project: August Strait of Georgia Hake Survey

Party Chief: Mark Saunders

Platform: CCGS W. E. Ricker

Date: 15 August 2000- 20 August 2000

Processed by: Germaine Gatien

Date of Processing: 8 May 2001 – 1 June 2001

Number of original CTD casts: 11

Number of casts processed: 11

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0506) was mounted with Transmissometer 197. The deck unit was a SeaBird 11+ S/N 11P18377-0471.

SUMMARY OF QUALITY AND CONCERNS

Since the sampling interval was only 1.375s the quality of the data is lower than usual for this instrument.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

No CTD Log was available. A spreadsheet with positions and activities for CTD casts was obtained.

Salinity calibration data was obtained, but without a log the depth of the samples is uncertain. It will be assumed that the deep samples were taken when the CTD was at the bottom of the cast so the bottle salinity is likely 1db above the maximum pressure as this has been the practice on previous PBS cruises. The cruise summary sheet was completed.

This cruise was a continuation of cruise 2000-05 so the same configuration file was used: 506ctd.con.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using conversion file 506ctd.con.

An initial examination of the data shows all channels present. The pairs of sensors appear to be in reasonable agreement as do the down and up-casts. However, there is very little data. The sampling interval is 1.375s although the CON file was set at 24samples/second.

4. ALIGNCTD

This step was skipped as there is insufficient data for it to be useful.

5. WILDEDIT

This step was skipped as there is insufficient data for it to be useful. Any bad points will be identified in the editing process.

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run to calculate primary and secondary salinity. A second run of DERIVE was used to calculate differences between the pairs of sensors and to examine the descent rate. This data will not be archived.

8. Test Plots and Channel Check

Two of the deeper casts (# 4,6) were plotted to check for agreement between the pairs of T and C sensors. All channels were present. The descent rate was quite high and steady. The differences in temperature, conductivity and salinity were found to be about 0.0004C, 0.0007units and 0.008units, respectively. These values are a little higher than those found 2000-34 in October 2000. They are similar to the results for 2000-05 except that the temperature differences are lower, probably because the descent rate was much steadier.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. The GPS headers needed editing for 3 of the casts. (Latitude, longitude and time entries were in the wrong places.)

10. Checking Headers

A header summary was produced. Two errors in station names were corrected.

The surface check was run. The average surface pressure is 2.3db, which is reasonable. However, for 3 casts the surface pressure was off-scale. It appears that the pump was not turned on when logging started. Sea-Bird CLEAN plus editing should remove bad data caused by this.

The header check was produced and no errors were found.

The cruise track was plotted and looks reasonable.

11. Test Plots

Profiles of all casts were examined for any evidence of problems with the processing and none were found. Given the sampling rate the up and down-casts are reasonably similar, as are the pairs of sensors.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 20.00

Minimum Surface Salinity: 10.0

Surface Swell Pressure Tolerance:  0.5

Pressure NOT filtered

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 5 points) was deleted.

Sample interval =  1.375 seconds.

The secondary sensors were chosen for further processing since a preliminary check suggests they are closer to the bottles and they were used for cruise 2000-34.

IOS SHELL routine Sea-Bird CLEAN was used to remove records with negative pressures.

All CLN files were copied to EDT files.

- DETAILED EDITING

Page plots were produced using T1,S1 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Since the sampling interval was large there is little need for editing. CTDEDIT was used to remove 3 records at the surface for cast #11. Note was made of the editing details in the relevant file and it was copied to an EDT file.

14. BIN AVERAGE

This step was skipped since the sampling interval was 1.375s.

15. Intercomparisons

Previous experience with these sensors –

| Mission | Start Date | Sensors 2278/2710 | Sensors 2280/2095 |
| --- | --- | --- | --- |
| 2004 | 1/4/2000 | 3 bottles only –0.014 | 3 bottles only –0.004 |
| 2009 | 1/5/2000 | Drift from –0.003 to +0.002 at end | +0.0005 |
| 2015 | 31/5/2000 | No reliable info | No reliable info |
| 2016 | 26/6/2000 | Surface only - +0.001 | Surface only –0.003 |
| 2034 | 03/10/2000 | 8 bottles +0.0015 | 8 bottles -0.002 |

COMPARE – There was no rosette. Deep salinity samples were gathered for 2 casts during 2000-05 and 3 during 2000-29. The comparison is shown in file 2005comp.xls. While pressures were recorded for 3 of the 5 bottles, they appear to be wrong. For cast #6 the spreadsheet indicates that the bottle was tripped at 660m whereas the maximum pressure is 501db. It was assumed that the pressures were 1db less than the maximum depth recorded. This is the difference used for previous Stock Assessment cruises. The primary salinity was found to be low by an average of 0.012units and the secondary low by 0.003units. The result for the secondary salinity is reasonably close to those of 2000-16 and 2000-34. The difference between the two salinities is close to the difference found between the primary and secondary salinity in step 8 (0.008units). However, the primary salinity error is larger than for all other cruises since last calibration with the exception of 2000-04 for which there were only 3 samples available. While the pressures may well be wrong, choosing salinities from a shallower depth would only increase the error.

Historic ranges – Historic ranges were found for four casts and all data fell within the ranges.

- Recalibration

No recalibration was done since the secondary CTD salinities are reasonably close to the bottles and the comparison is based on very few bottles; similarly the results of previous cruises using this equipment are based on little information.

- HEADEDIT AND REORDER

A note was added to the comment section of all the headers to advise users of the low sampling rate.

The following channels were removed from all casts: Scan_Number, Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The data format was corrected as needed.

The output files had extensions CTD.

18. Final Plots and Cross-Reference Listing

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the files with extensions CTD.

A cross-reference listing was produced.

### Particulars

20290001 – Pump turned on after CTD. Surface values not good.

20290010 – Pump turned off during upcast from 125m to 113m approx.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2000-29

Dates: Start: 15 August 2000 End: 20 August 2000

Location: NE Pacific

Vessel: CCGS W. E. RICKER

Chief Scientist: Mark Saunders

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		2000-29

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 14Dec1999 | Factory |  |  |
| Conductivity | 2278 | 16Dec1999 | “ |  |  |
| Secondary Temp. | 2095 | 16June1999 | “ |  |  |
| Secondary Cond. | 2280 | 16Dec1999 | “ |  |  |
| Transmissometer | 197 | 12Jan2000 | IOS |  |  |
| Pressure Sensor | 69698 | 12Nov1997 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.