## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 10-Jan-2006 | Surface loop data was added to the archive. The data was acquired from John Morris at PBS. The data archived in September 2005 was replaced with this new set since chlorophyll was included. The original spreadsheet file from John and more detailed processing notes can be found in the “Cruise_Data\Documents” directory. Any questions regarding this data should be directed to John Morris. J.L. |
| 26-Sep-2005 | Added loop data to the archive. J.L. |

PROCESSING NOTES

Cruise: 2000-34

Agency: OSAP

Location: NE Pacific

Project: High Seas Salmon

Party Chief: John Morris

Platform: CCGS W. E. Ricker

Date: 3 October 2000- 21 October 2000

Processed by: Germaine Gatien

Date of Processing: 22 February 2001 – 23 April 2001

Number of original CTD casts: 139

Number of casts processed: 136

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0506  ) was mounted with Transmissometer 197. The deck unit was a SeaBird 11+ S/N 11P18377-0471.

SUMMARY OF QUALITY AND CONCERNS

The quality of the data was high despite the fact that the descent rate of the CTD was very noisy for many casts. This is probably due to the fact that the descent rate was kept high.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The CTD Log was obtained. It was noted that in the log book times are in PDT for the first 59 casts and UTC thereafter. The times are correct in the data files.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants were checked. There is an error in the pressure calibration. The same error was found during cruise 2000-16. At that time it was noted that the surface pressures seemed too high and correcting the error reduced the pressures. The calibration file 0506ctd.con was fixed.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using conversion file 0506ctd.con.

An initial examination of the data shows all channels present; the secondary conductivity looks noisy.

4. ALIGNCTD

The deck unit used for this mission advances only the primary conductivity channel by 0.073s. Tests on casts #1 and 47 showed that advancing the secondary by anything between 0.035s and 0.060s improved the data. However, the data is quite noisy in general so it is hard to tell which of these settings is best. For previous use of this equipment in 2000-04 and 2000-16 an alignment of 0.060s was used. The primary conductivity was advanced by –0.013s (for a net advancement of +0.060s) which improved the primary salinity data significantly.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2		Pass 2    Std Dev = 5		Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of deep casts (# 25,70 and 106) was plotted to check for agreement between the pairs of T and C sensors. All channels were present. The descent rate was noisy. The differences in temperature, conductivity and salinity were found to be very noisy but averaged about 0.001C, 0.0004units and 0.004units, respectively below 900m. These values are similar to those found during 2000-16.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. There was no rosette used, so there are no rosette files to convert.

10. Checking Headers

A header summary was produced and a few errors in station names and one error in longitude were found and corrected.

The surface check was run. The average surface pressure is 2.4db, which is reasonable.

The header check was produced and no errors were found.

The cruise track was plotted and looks reasonable.

11. Test Plots

Profiles of all casts were examined for any evidence of problems with the processing. The downcasts and upcasts are reasonably close as are the two temperature and salinity traces except for the following casts:

#2 - the sensors were quite different near 20m. T1 S1 look wrong.

#50 - the primary temperature and salinity appear to be bad; probably something was stuck in the pump.

#81 – the primary salinity looks wrong in the upcast

#97 – the upcast looks quite different from the downcast, but the range is small and the pairs of sensors track well.

There were a many spikes in salinity in both channels but the primary is perhaps a little worse than the secondary.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance:  1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS

Most warnings refer to surface and upcasts only. Mid-depth downcast warnings were investigated for 12 casts and there does not appear to have been a significant loss of data. For casts #26,27,80,81 and 82 there was a loss of what seems to be good data. For these cases the record containing bad pressure data  was replaced with a copy of the previous record. DELETE was then rerun for these casts.

The secondary sensors were chosen for further processing except for cast #2.

All DEL files were copied to EDT files.

- DETAILED EDITING

Page plots were produced using T1,S1 for all casts (T0,S0 for cast #2). These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

CTDEDIT was used to clean noise in S and T near the surface and bottom only for casts #3,5, 19,91,113,114,129,131.

CTDEDIT was used for more extensive cleaning in the following casts: #1,4,6-18,20-43,46,48-51,53,54,56-61,63-72,74,75,81,86,90,98,99,103,105,107-108,111,122,125,126,128,134-137.

Cast #97 was left unedited. It is a 275m deep cast with only 0.7C range of temperature. There are many small spikes and blocks of data appear offset. The upcast looks very different from the downcast, but the two sets of sensors behaved alike. Nearby stations look quite different but the closest is shallow and the only deep one in the area is not very close and is further from open ocean. The spikes mostly line up with sigma-t lines.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

14. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

15. Intercomparisons

Previous experience with these sensors –

| Mission | Start Date | Sensors 2278/2710 | Sensors 2280/2095 |
| --- | --- | --- | --- |
| 2004 | 1/4/2000 | 3 bottles only –0.014 | 3 bottles only –0.004 |
| 2009 | 1/5/2000 | Drift from –0.003 to +0.002 at end | +0.0005 |
| 2015 | 31/5/2000 | No reliable info | No reliable info |
| 2016 | 26/6/2000 | Surface only - +0.001 | Surface only –0.003 |

COMPARE – There was no rosette. Deep salinity samples were gathered for 8 casts using a Niskin bottle. The comparison is shown in file 2034comp.xls. Since the depths are approximate there is some uncertainty in the comparison. Data from files with extensions IOS were used for the comparison and 20 values were averaged from the estimated depth of the bottle. The primary salinity was found to be 0.0015units high and the secondary 0.0020units low. This is reasonably close to the 2000-16 results, and is consistent with the difference found between the primary and secondary salinity in step 8 (0.004units).

Historic ranges – No historic ranges were found for these casts.

- Recalibration

No recalibration was done since the CTD salinities are reasonably close to the bottles and the comparison is based on very few bottles; similarly the results of previous cruises using this equipment are based on little information.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

- REMOVE

The following channels were removed from all casts except #2: Scan_Number, Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The following channels were removed from cast #2: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The data format was corrected as needed.

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

2034000t – aborted

20340002 – Problem with secondary sensor around 20m; primary used for this cast.

20340007 – station name wrong in header; changed to EP01.

20340011 – aborted – file exists check later

20340035 – aborted – file exists check later

20340050 – station name wrong in header; changed to X004. The up and down casts look very different.

20340097 – both sets of sensors recorded many spikes and offset blocks of data. The upcast looks very different from the downcast. Nearby casts are quite different. The temperature range is small and the spikes mostly align with sigma-t lines.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2000-34

Dates: Start: 3 October 2000 End: 31 October 2000

Location: NE Pacific

Vessel: CCGS John P. Tully

Chief Scientist: John Morris

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		2000-34

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 14Dec1999 | Factory |  |  |
| Conductivity | 2278 | 16Dec1999 | “ |  |  |
| Secondary Temp. | 2095 | 16June1999 | “ |  |  |
| Secondary Cond. | 2280 | 16Dec1999 | “ |  |  |
| Transmissometer | 197 | 12Jan2000 | IOS |  |  |
| Pressure Sensor | 69698 | 12Nov1997 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.