# PROCESSING NOTES

Cruise: 2000-36

Agency: PBS

Project: Beam

Geographic Area: Strait of Georgia

Chief Scientist: Beamish R.

Platform: W. E. Ricker

Date of Cruise: 28 March 2000- 1 April 2000

Processed by: Germaine Gatien

Date of Processing: 18 May 2005 – 23 May 2005

Number of original casts: 28

Number of casts processed: 10

# INSTRUMENT SUMMARY

One CTD was used during this cruise.

Instrument Type:  Guildline 8715 CTD

Instrument Serial Number: 53501

# SUMMARY OF QUALITY AND CONCERNS

There was no calibration sampling available. In previous use (in early 1998) the salinity was found to be within ±0.005psu for most samples. This instrument is generally quite stable, so an estimate of ±0.01psu is reasonable for this cruise. Only 3 decimal places will be archived for salinity.

# PROCESSING SUMMARY

- PRELIMINARY STEPS

There was no Log Book available, but there is a spreadsheet that contains some data, positions, station names and a few comments.

There was one comment that suggests that bottle sampling was done, but no analysis data were available.

The cruise summary sheet was completed.

A cruise track plot was made and it looks reasonable.

- CONVERSION OF RAW DATA

- The files were converted using the IOSSHELL routine to convert GUILDLINE files.

- CALIBRATION

- All data were calibrated using file 2000-36cal.ccf which was based on file 9818cal.ccf.

- The header summary and header check were examined. There are no station names in the headers, but there is a spreadsheet file with such names for some casts and some names in comments in the headers. The spreadsheet entries were compared with the header summary and the maximum pressure check to see that it was appropriate for that station; the station names were entered in the headers.

- There were separate downcast and upcast files. Some files contained no useful data, either there were just a few records from the surface, or the pressure channel was bad. Upcast files will not be processed further.

- The files that will be processed further are: 5, 8, 12, 15, 19, 20, 22, 24, 26, 28, 31.

- The surface check gave an average surface pressure of -1.1db. Individual files were examined to see if there is a problem with a pressure offset, and the low values appear to be from the period when the pressure sensor equilibrates. By the time the temperature and conductivity show that the CTD was being lowered, the pressures look reasonable. For cast #5 it is very clear that the surface value is very close to 0, as conductivity goes from near-zero to “in-water” values at pressures very close to zero. Checking a few upcasts also confirm that the pressures are good.

- DESPIKE

- CLEAN was used to remove pad values by interpolation.

- The following ADVANCED DESPIKE parameters were used:

- CHANNEL        FIT   OVER-  MIN       MAX       MIN         MAX         SPIKE

- WIDTH LAP    VALUE   VALUE  STDDEV  STDDEV     TOL.

- -------------------- ----- ----- ------- ------- ------- ------- -----

- Pres 		55     5        -2.00       4000.00	.10000        999.000    3.00

- Temp   	55     5        -2.00          15.00	.00500        999.000    3.00

- Cond    	55     5           .00            1.50	.00010        999.000    3.00

- A few plots were examined to ensure that salinity spikes were removed successfully, and they were.

- TIME COMPENSATION

- Temp. probe Dist (m):   .0700  Sample period (sec):   .0400

- DELETE

- The following DELETE parameters were used:

- Surface Record Removal: Low Sal & Last Press Min

- Maximum Surface Pressure (relative):      10.00

- Minimum Surface Salinity:     5.000

- Surface Swell Pressure Tolerance:       1.00

- Pressure filtered over width:   11

- Swells deleted. Warning message if pressure difference of      2.00

- Drop rates<    .30m/s (calculated over  11 points) will be deleted

- Drop rate applies in the range    10db to 10db less than the maximum pressure

- Sample interval =       .04 seconds.

- There were no warnings.

- FILTER

- A median filter with filter width 15 was applied to the salinity channel.

- CTDEDIT

CTDEDIT was used to edit temperature and salinity. Some records near the top and bottom of the cast were removed. Salinity was smoothed.

Details were added as comments to the headers.

All files required some editing. The salinitydata for cast #19 look unbelievable.

9.   BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure    Averaging interval =      1.000

Minimum bin value =    .000     Average value will be used

Interpolated values are NOT used for empty bins

10.Final Plots

Plots were made of cast #5 and 19 together since they were from the same site and this confirms that #19 has bad salinity. The temperature data for #19 are probably good, but since there is already a good cast from this site (#5) this cast will not be archived.

The bin-averaged files were thinned to standard depths and then put through DERIVED QUANTITIES to calculate theta, gamma, depth, SVA and speed of sound.

Page plots were produced.

11. Quality Control Tests

There were no bottle data available.

Profile plots, with historic ranges superimposed, were examined on-screen. All data were found to fall within the local climatology and no further editing was found necessary.

Only one processed cruise could be found during which this instrument was used; that was 9818. There was some salinity calibration sampling and the salinity was found to be within ±0.005psu for most samples. Generally Guildline CTDs are quite stable, so we might expect that the salinity is within ±0.01psu for this cruise.

12.	REMOVE CHANNELS

The following channel was removed from all casts: Conductivity_Ratio.

13.	PRODUCE FINAL FILES

a.) HEADEDIT was used to fix channel names and formats, fix the ship name and add the following comments:

There was no salinity calibration sampling.

The last calibration on file was from 1998. This is probably the

right one for this data set, but it is not certain.

Calibration sampling in 1998 showed that most samples were within +/-

0.005psu, so a reasonable estimate for this cruise is +/-0.01psu.

b.) The final files were renamed *.ctd. The Standard Check routine and HEADEDIT were rerun until all problems were resolved. Note that the salinity format is non-standard to indicate the limitations in the quality of that data.

c.) The start and stop times had the year 1980. These were changed to 2000 using a text editor.

A cross-reference listing was produced.

PARTICULARS

6, 9,13,16,18,21,23,25,27,29,32 – upcast data only

7 – empty file

10 – only 11db of data; repeat of cast #8 which went to 301db.

11, 30 – bad pressure data.

14, 17 – only a little data at the surface.

19 – according to the spreadsheet the values were very low for this cast. This was confirmed by comparison with cast #5.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2000-36

Dates: 	Start:	28 March 2000	    End:	1 April 2000

Location:	Strait of Georgia / Juan de Fuca Strait

Vessel:		W. E. Ricker

Scientific Party Chief:    Beamish R.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | Guildline | 8715 | 53501 | no | yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	Guildline /8715/53501

Cruise ID#:	2000-36

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure |  |  |  |  |  |
| Temperature |  | 7 Aug 98 | I.O.S. |  |  |
| Conductivity |  | 7 Aug 98 | I.O.S. |  |  |
| Transmissometer | 572 | 31 July 98 | I.O.S. /Doug Anderson |  |  |

| Calibration Coefficients |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Channel | Formula | Coefficients |  |  |  |  |  |
| Name | # | C1 | C2 | C3 | C4 | C5 | C6 |
| Temperature | 10 | .0027318 | .999875 |  |  |  |  |
| Conductivity | 10 | .0000503156 | .999729 |  |  |  |  |