# PROCESSING NOTES

Cruise: 2000-37

Agency: PBS

Project: Beam

Geographic Area: Strait of Georgia

Chief Scientist: Beamish R.

Platform: W. E. Ricker

Date of Cruise: 15 April 2000- 28 April 2000

Processed by: Germaine Gatien

Date of Processing: 20 May 2005 – 24 May 2005

Number of original casts: 38

Number of casts processed: 18

# INSTRUMENT SUMMARY

One CTD was used during this cruise.

Instrument Type:  Guildline 8715 CTD

Instrument Serial Number: 53501

# SUMMARY OF QUALITY AND CONCERNS

There was no calibration sampling available. In previous use (in early 1998) the salinity was found to be within ±0.005psu for most samples. This instrument is generally quite stable, so an estimate of ±0.01psu is reasonable for this cruise. Only 3 decimal places will be archived for salinity.

The descent rate of the CTD was very noisy with frequent reversals in direction especially for the casts off the west coast of Vancouver Island. DELETE and a graphical editor were used to remove data that were clearly corrupted by shed wakes, but it is expected that some of the data that remains is affected by such corruption

# PROCESSING SUMMARY

- PRELIMINARY STEPS

There was no Log Book available, but there is a spreadsheet that contains some data, positions, station names and a few comments.

The cruise summary sheet was completed.

A cruise track plot was made and it looks reasonable.

- CONVERSION OF RAW DATA

- The files were converted using the IOSSHELL routine to convert GUILDLINE files.

- CALIBRATION

All data were calibrated using file 2000-37cal.ccf which was based on file 9818cal.ccf.

The header summary and header check were examined. There are no station names in the headers. There were separate downcast and upcast files. Some files contained no useful data, either there were just a few records from the surface, or the pressure channel was bad. Upcast files will not be processed further.

The casts that will be processed further are: 2, 3, 6, 7, 10, 13, 15, 18, 21, 23, 25, 27, 29, 31, 33, 35, 37 and 39.

The surface check for the 18 downcast files gave an average surface pressure of -0.7db. Individual files were examined to see if there is a problem with a pressure offset, and the low values appear to be from the period when the pressure sensor equilibrates. By the time the temperature and conductivity show that the CTD was being lowered, the pressures look reasonable. Checking a few upcasts suggests that the pressures may be slightly low, but within 1db.

- DESPIKE

- CLEAN was used to remove pad values by interpolation.

- The following ADVANCED DESPIKE parameters were used:

- CHANNEL        FIT   OVER-  MIN       MAX       MIN         MAX         SPIKE

- WIDTH LAP    VALUE   VALUE  STDDEV  STDDEV     TOL.

- -------------------- ----- ----- ------- ------- ------- ------- -----

- Pres 		55     5        -2.00       4000.00	.10000        999.000    3.00

- Temp   	55     5        -2.00          15.00	.00500        999.000    3.00

- Cond    	55     5           .00            1.50	.00010        999.000    3.00

- A few plots were examined to ensure that salinity spikes were removed successfully, and they were.

- TIME COMPENSATION

- Temp. probe Dist (m):   .0700  Sample period (sec):   .0400

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Low Sal & Last Press Min

Maximum Surface Pressure (relative):      10.00

Minimum Surface Salinity:     5.000

Surface Swell Pressure Tolerance:       1.00

Pressure filtered over width:   11

Swells deleted. Warning message if pressure difference of      2.00

Drop rates<    .30m/s (calculated over  11 points) will be deleted

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval =       .04 seconds.

There were no warnings.

- FILTER

- A median filter with filter width 15 was applied to the salinity channel.

- CTDEDIT

The descent rate of the CTD was very noisy for some casts, especially from casts #21 through 31 plut #39; there is evidence of corruption by shed wakes.

CTDEDIT was used to edit temperature and salinity. Some records near the top and bottom of the cast were removed as were mid-cast records that appeared to be corrupted. Salinity was smoothed.

Cast #6 required no editing. Details were added as comments to the headers.

7.   BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval =      1.000

Minimum bin value =       .000

Average value will be used

Interpolated values are NOT used for empty bins

- Final Plots

Profile plots, with historic ranges superimposed, were examined on-screen. All data were found to fall within the local climatology and no further editing was found necessary.

The bin-averaged files were thinned to standard depths and then put through DERIVED QUANTITIES to calculate theta, gamma, depth, SVA and speed of sound.

Page plots were produced.

- Quality Control Tests

There were no bottle data available.

This CTD was used for 4 cruises during 2000, but there was no salinity sampling available for any of those cruises. Only one other processed cruise could be found during which this instrument was used; that was 9818. There was some salinity calibration sampling and the salinity was found to be within ±0.005psu for most samples. Generally Guildline CTDs are quite stable, so we might expect that the salinity is within ±0.01psu for this cruise.

13.	REMOVE CHANNELS

The following channel was removed from all casts: Conductivity_Ratio. The year was wrong in the start and top dates; this was changed using a text editor.

14.	PRODUCE FINAL FILES

a.) HEADEDIT was used to fix channel names and formats, fix the ship name and add the following comments:

There was no salinity calibration sampling.

The last calibration on file was from 1998. This is probably the

right one for this data set, but it is not certain.

Calibration sampling in 1998 showed that most samples were within +/-

0.005psu, so a reasonable estimate for this cruise is +/-0.01psu.

b.) The final files were renamed *.ctd. The Standard Check routine and HEADEDIT were rerun until all problems were resolved. Note that the salinity format is non-standard to indicate the limitations in the quality of that data.

c.) A cross-reference listing was produced.

15. PARTICULARS

4, 8, 11, 14, 16, 19, 22, 24, 26, 28, 30, 32, 34, 36, 38, 40 – upcast data only

5, 9 – only a little surface data

6 – record does not start until 15.9db

12 – empty file

20 – bad pressure data

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2000-37

Dates: 	Start:	15 April 2000	    End:	28 April 2000

Location:	Strait of Georgia / Juan de Fuca Strait

Vessel:		W. E. Ricker

Scientific Party Chief:    Beamish R.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | Guildline | 8715 | 53501 | no | yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	Guildline /8715/53501

Cruise ID#:	2000-37

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Pressure |  |  |  |  |  |
| Temperature |  | 7 Aug 98 | I.O.S. |  |  |
| Conductivity |  | 7 Aug 98 | I.O.S. |  |  |

| Calibration Coefficients |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Channel | Formula | Coefficients |  |  |  |  |  |
| Name | # | C1 | C2 | C3 | C4 | C5 | C6 |
| Temperature | 10 | .0027318 | .999875 |  |  |  |  |
| Conductivity | 10 | .0000503156 | .999729 |  |  |  |  |