## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 2001-02

Agency: Stock Assessment

Location: Eastern Pacific

Project:  Hake Survey

Party Chief: Mark Saunders

Platform: CCGS W. E. Ricker

Date: 12 January 2001- 4 February 2001

Processed by: Germaine Gatien

Date of Processing: 4 June 2001 – 14 June 2001

Number of original CTD casts: 45

Number of casts processed: 45

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0506) was mounted with Transmissometer 197.

SUMMARY OF QUALITY AND CONCERNS

The quality of this data is lower than usual probably because of a missing bridle on the CTD leading to flow-rate problems. The bottle calibration suggests errors of up to ±0.01units. No recalibration was attempted because of problems in the calibration.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

Sheets from the CTD Log was obtained. There is no list of the equipment used.

Salinity calibration data was obtained. The cruise summary sheet was completed.

It is believed that the same equipment was used for cruise 2001-09 during which it was discovered that the configuration file had errors in it.  The new one created after that discovery contained an error in the pressure calibration. The same error was found during cruise 2000-16 (and later Ricker cruises). At that time it was noted that the surface pressures seemed too high and correcting the error reduced the pressures. The calibration file 506Final.con was created with the corrected pressure calibration.

The sensor history was found.

Notice was received from Gail Jewsbury that the bridle for the CTD was found to be missing in April 2001. This means that the CTD could hang at an angle during descent. It is presumed that it was missing for all missions from January 2001 to late April 2001. The data collected during this time may be affected.

- Conversion of Raw Data

The raw data was converted using conversion file 506ctd.con.

An initial examination of the data shows all channels present. When this equipment was used for cruise  2001-09 there were significant differences between the primary and secondary conductivity particularly in the top 30m of many casts. This was believed to be because of a missing bridle. This does not appear to be the case for this cruise possibly because there is a deep mixed layer.

4. ALIGNCTD

It is assumed that the same deck unit was used for this mission as was used for cruise 2001-09. That unit advances only the primary conductivity channel by 0.073s. Tests were made on casts #20 & #40 by       trying a variety of advancements. The results were inconclusive with different advancements having the best results for different parts of the casts. This may be a result of the CTD orientation varying with time due to the absence of the bridle. It was decided to use +0.055s since that was used for 2001-09 and it did improve the traces overall. So the primary conductivity was advanced by –0.018s (for a net advancement of +0.055s) and the secondary was advanced by +0.055s.

The fact that bi-polar spikes remain after this step is considered diagnostic of flow-rate problems. This is probably due to the lack of a bridle on the CTD.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2		Pass 2    Std Dev = 5		Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of deep casts (# 13,20 and 40) was plotted to check for agreement between the pairs of T and C sensors. The descent rate was noisy but average values were high. The differences were all very noisy averaging at about 0.0008C for temperature, 0.0001units for conductivity and 0.0005units for salinity at about 650db with the secondary sensors giving the higher values for all three variables. These differences  are similar to those found during 2000-09 for temperature, but much lower for conductivity and salinity. The depths of the comparison are greater for this cruise which would partly explain the lower values. Problems with orientation of the sensors would also lead to great variability.

All casts were examined for problems. For most casts the pairs of sensors produced similar results and the upcasts resemble the downcasts. Exceptions were casts #35 where the two salinity channels produced very different results between 60 and 85db and neither signal resembles the upcast; the temperature channels are similar and quite smooth through that region. The CTD log notes that the pump was not turned on until 50db. The upcast will be used for this cast. For cast #38 there is a sudden increase in T and S at 50db which does not appear in the upcast and which does not correspond to a shed wake. This may be a real feature or may be a case of the pump coming on late. This will be investigated more later.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. Several casts had rotated NMEA headers which had to be corrected to enable conversion.

10. Checking Headers

The cruise track was plotted and one error in position found. Examining the CNV file it was found that the NMEA longitude had several odd characters in it. After fixing this error the file was reconverted and the plot looks reasonable.

A header summary was produced. Errors in station names were found and corrected. A descrepancy in position and time between the CTD log and the NMEA headers was found for cast #21. The NMEA reading is in good agreement with the positions and times given for the bottom and the end of the cast, but disagrees with the beginning position and time. The NMEA reading appear to be correct rather than the log so no change was made.

The surface check was run. The average surface pressure is 2.7db.

The header check was produced and no errors were found.

- DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance:  1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS – No problems found.

All DEL files were copied to EDT files.

File 21020001.CLN was put through REVERSE and DELETE since there is only upcast data.

Files 21020035.CLN and 21020038.CLN were put through REVERSE, renamed 21029935.REV and 21029938.REV and then put through DELETE to create UPCAST versions of the two casts for comparison with what appear to be problem downcasts.

12. Test Plots

Page plots for the upcasts and downcasts for cast #35 and 38 were compared.

The upcast is clearly better for cast #35 and so 21029935.DEL will be copied to 21020035.DEL.

For cast #38 there is an odd feature in T and S but no evidence that it is not real. While the upcast does not have as large a change in T and S as the downcast the general variation is about the same. There is no evidence of shed wakes, the feature is stable and both sets of sensors “see” the same feature. Quite a lot of data was removed by WILDEDIT in this region and if it is real there does not seem to be any justification for such removal. So the processing of this cast was rerun with only Pressure put through WILDEDIT. The downcast will be used for #38.

13. Intercomparisons

Previous experience with these sensors – These sensors were used for cruise 2001-09 in March 2001 when the bridle was also missing. There were also problems with a pump and the calibration drfited during the cruise. The secondary salinity was raised by 0.005units. There was insufficient information to recalibrate the primary salinity.

##### Bottle comparisons

- DEEP BOTTLES - There was no rosette. Deep salinity samples were gathered for 11 casts. The depth of the samples is approximate. Estimated errors due to uncertainty in the depth of the Niskin bottle is ±0.002units. The average differences found between the CTD and the bottles were +0.009 and +0.010units for the primary and secondary sensors respectively. The errors appear to vary with pressure but there is a lot fo scatter and 4 out of 7 bottles at 700m have values close to zero. (See file 2102comd.xls)

- SURFACE BOTTLES - There were Niskin bottle samples at 5m for 14 casts. Again, it is not known if the Niskin bottles were at 5m or if the pressure from the CTD read 5db. The average differences were –0.006 and –0/005units for the primary and secondary sensors. (See file 2102coms.xls.)

Thus the CTD appears to read low at the surface and high at depth. Such strong depth dependence is unusual for this instrument. The results at the surface do not seem likely to be caused by uncertainties in the depth of the bottle since the surface waters are well mixed for these casts. The surface results also are consistent with the results for 2001-09, but there were few bottles for that cruise and they were not very deep.

Historic ranges – Historic ranges were only available for a few casts. All data fell with the ranges.

14. DETAILED EDITING

Page plots were produced using T1,S1 for all casts since there appeared to be less noise in the secondary salinity. These plots were used to guide the use of CTDEDIT.

CTDEDIT was used to clean noise in S and T near the surface and bottom only for casts #1,20,33 & 34.

CTDEDIT was used for more extensive cleaning in the following casts: #4,7,10,13-15,18,22,24,27,29-32,35-43 & 45.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

15. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

- Recalibration

Given the uncertainties in the depth of the samples, the unusual depth dependence of the calibrations and the differences from the calibrations done during 2001-09 it seems unwise to recalibrate the data.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

- REMOVE

The following channels were removed from all casts: Scan_Number, Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The data format was corrected as needed.

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

1 - Calibration cast, upcast only.

31 – Pump not turned on for downcast. Use upcast.

35 – Pump not turned on until aobut 50db- use upcast.

38 – Odd feature in T and S around 50db. Appear to be real features. Use Downcast.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001-02

Dates: Start: 10 January 2001 End: 4 February 2001

Location: Eastern Pacific

Vessel: CCGS W. E. Ricker

Chief Scientist: Mark Saunders

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		2001-02

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 30Nov2000 | Factory |  |  |
| Conductivity | 2278 | 6Dec2000 | “ |  |  |
| Secondary Temp. | 2710 | 30Nov2000 | “ |  |  |
| Secondary Cond. | 2280 | 6Dec2000 | “ |  |  |
| Transmissometer | 197 | 14Dec2000 | IOS |  |  |
| Pressure Sensor | 69698 | 12Nov1997 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.