## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

PROCESSING NOTES

Cruise: 2001-04

Agency: West Vancouver Lab and IOS/OSAP

Location: Broughton Archipelago

Project: Aquaculture

Chief Scientist: Dario Stucchi

Platform: CCGS VECTOR

Date: February 9, 2001 – February 17, 2001

Processed by: Germaine Gatien

Date of Processing: 22 October 2001 – 25 October 2001

Number of original CTD casts: 27

Number of casts processed: 26

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with Transmissometer 192D and Seapoint Fluorometer #2229. There is no record of the deck unit serial number.

SUMMARY OF QUALITY AND CONCERNS

The data were in good order.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The CTD Log was obtained.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked; the only error was that the transmissivity calibration was an old one. A new file 2104CTD.con was created with the new values.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using configuration file 2104CTD.con

A preliminary check shows all expected channels present.

There is no note in the CTD log about what cable was used. Dark values were checked for the fluorometer; the deepest cast is 470db where the minimum concentration is about 0.08g/l. These values are consistent with the 10X cable. The observed maxima seem reasonable.

The temperature and salinity traces track reasonably well. The secondary salinity shows fine-scale noise (on the order of 0.005units) as was noted in other cruises using these sensors. This is due to noise in the secondary conductivity channel. Cast #21 has data spikes.

4. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels. The bottle number channel was stripped from the CNV files.

5. ALIGNCTD

The deck unit serial number is unknown. A test was made on cast #18 and it was determined that the secondary conductivity had not been advanced by the deck unit, which means it is one of the older units which only advance the primary conductivity channel by 0.073s.  A net advancement of +0.055s was used for both sensors during 2000-28, 2000-30 and 2001-32. These values look reasonable for this data as well, so all casts were advanced by –0.018s for the primary conductivity and +0.055s for the secondary.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels except scan number.  Parameters used were:   Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

Tests were run on cast #18 to determine the optimal parameter choice. The results with (α, 1/β) = (0.02,9.0) or (0.03,9.0) were similar with the former giving slightly better results.  The conductivity cell thermal mass correction was done for both channels using α = 0.02; 1/ β = 9.0 as was used for 2001-32.

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present.

The differences in conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 7 | 230 | -0.0002 | +0.00025 | +0.0025 | High, steady |
| 8 | 470 | -0.0002 | +0.0002 | +0.0025 | High, steady |
| 18 | 325 | -0.0003 | +0.0002 | +0.002 | High, steady |

All casts were plotted on-screen to check for problems with processing. Reasonable agreement was found between pairs of sensors. The up and down traces are as similar as might be expected in such a rapidly changing environment. The spikes noted in cast #21 have been removed.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

10. Checking Headers

A header summary and header check were produced. No errors were found.

The average surface pressure is 2.4db.

The cruise track was plotted and looks reasonable.

The cruise summary and cruise report were produced and no errors found.

- DELETE

- CLEAN was run to replace pad values in pressure with interpolated values.

- The following DELETE parameters were used:

- Surface Record Removal: Last Press Min

- Maximum Surface Pressure (relative): 10.00   Pressure Tolerance: 1.0

- Pressure filtered over 15 points

- Swells deleted. Warning message if pressure difference of 2.00

- Drop rates < 0.3m/s (calculated over 15 points) was deleted.

- Sample interval =  .04 seconds.

- COMMENTS ON WARNINGS: There were no warnings.

- A comparison of maximum depths before and after DELETE indicated no problems.

- All DEL files were copied to EDT files.

- 12. COMPARE

- The BOT files were examined for errors and a few bad points deleted from bottle #3, cast #7.

- COMPARE was run. There are only 8 bottle values. The results are in excellent agreement with those found when these sensors have been used previously. The primary sensors have salinity values low by 0.020units and the secondary, low by 0.017units

- DETAILED EDITING

The primary sensors were chosen for further processing because the secondary conductivity is noisy.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

The following casts were edited lightly: 2,4,17,19,21,23-27.

The descent rate was generally steady and moderately high.

Note was made of the editing details in the relevant files.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

14. SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness. Cast #18 was examined before and after this step and showed that the filter was effective. BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. The ¼m-averaged files were stored in a separate directory for the use of Angelica Peña.

15. BIN AVERAGE

The following Bin Average values were applied to the BOX files for general use:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen to check whether further editing was needed.

15. Other comparisons

Previous experience with these sensors – Both pairs of sensors were used during 2000-07,2000-06,2000-30 and 2000-28 in April, May, September and October of 2000 respectively and for 2001-32 in September 2001. The primary sensors were selected for archiving in all cases since the secondary conductivity was full of fine-scale noise. The CTD salinity values were lower than bottles with differences varying from –0.009 to –0.021units for the primary and –0.004 to –0.017units for the secondary. The largest differences occurred for the most recent cruises.

Repeat Casts – There were 6 casts at station PE-8. However this is a shallow site with rapidly changing conditions so does not provide a check on repeatability of the instrument.

Historic ranges – There were no historic range data available for this data set.

- Recalibration

File 2104rcal.ccf was used to recalibrate the AVG files (including those in the IOS directory, the HYDRO directory and the FLUOR directory) as follows:

- primary salinity –  offset: +0.020units

- secondary salinity –  offset: +0.017units

- Chl_fluorescence -  offset: -0.08g/l

The SAM files were recalibrated and named SAC and COMPARE was rerun on the recalibrated data. The results show that the differences are +0.00008 and -0.0004units for the primary and secondary salinity respectively.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data

18. REMOVE and REORDER

The following channels were removed from all CTD casts: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The channels were reordered and data format corrected as needed.

The channels for the rosette files were also reordered but only Conductivity:Primary, Conductivity:Secondary and Flag channels were removed.

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

- Cast not mentioned in CTD log and very shallow. DELETED.

21.  Spikes in data. Removed by WILDEDIT.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001- 04

Dates: Start: 9 February 2001 End: 17 February 2001

Location: Broughton Archipelago

Vessel: CCGS VECTOR

Chief Scientist: Dario Stucchi Party Chief: Terri Sutherland

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2001-04

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2663 | 02/05/00 | Factory |  |  |
| Conductivity | 2399 | 28/09/00 | “ |  |  |
| Secondary Temp. | 2371 | 07/03/00 | “ |  |  |
| Secondary Cond. | 2424 | 28/09/00 | “ |  |  |
| Fluorometer -pumped | 2229 | 18/06/01 | IOS |  |  |
| Transmissometer | 192D | 03/11/00 | IOS |  |  |
| Pressure Sensor | 77511 | 13/05/00 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.