## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| August 16, 2005 | Added DIC, Alkalinity, pH, PC and PN data from the Haida project to rosette files. Some files already contained DIC and Alkalinity data so that data was replaced with the new quality controlled version. J.L. |
| 29-Nov-2004 | CTD cast 5 – It looks like it had been bin averaged to every 0.3 dbars using the Sea-bird software so I meter averaged it using IOS Shell. J.L. |
| April 17, 2003 | Move files to new archive. |
| 16-jul-2001 | Reprocessed fluorometer data using a median filter (size 11) before bin averaging and recalibrating to remove offset. Bottle files recalibrated only. |
| 22-may-2001 | Casts 21,22 – removed transmissivity channel. J.L. |

PROCESSING NOTES

Cruise: 2001-06

Agency: OSAP

Location: North-East Pacific

Project: Line P/Haida Eddy

Party Chief: Marie Robert

Platform: CCGS John P. Tully

Date: February 5, 2001 – February 22, 2001

Processed by: Germaine Gatien

Date of Processing: April 23, 2001 – May 4, 2001

Number of original CTD casts: 48

Number of casts processed: 47

INSTRUMENT SUMMARY

Two SeaBird Model SBE 911+ CTDs (#0550 & #0443) were mounted with Transmissometer 333D and Seapoint Fluorometer # 2229. One or two PH sensors were mounted for some casts. The deck unit was a Seabird 11+, serial number 0508.

SUMMARY OF QUALITY AND CONCERNS

There were pressure spikes in 12 casts with associated errors in headers.

The descent rate of the CTD was very noisy for many casts resulting in corruption of data by shed wakes.

DELETE and CTDEDIT were used to remove many of the bad records. The lowest quality data is for casts #21-23 and #46-48 when the seas were very rough and the rosette was off.

No calibration or editing was done in processing the transmissivity, fluorometer and PH channels.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained. Salinity calibration data was obtained. The cruise summary sheet was completed. The configuration files were obtained and the calibration constants were checked.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using four different conversion files:

21060001.con was used for casts 1-3. (CTD#0550 without PH)

21060004.con was used for casts 4-19,24-45 and 49. (CTD#0550 with PH)

21060021.con was used for casts 21-23 and 46-48. (CTD#0443 without PH)

21060052.con was used for casts 52-67. (CTD#0443 with PH)

An initial examination of the data suggest some problems with the fluorometer when it was mounted on CTD #0443. The upcast data looks bad in the top 100-200m. The downcast data looks ok but should be investigated before use.

For temperature and salinity the pairs of sensors track well and the upcasts and downcasts look reasonably similar. Pressure spikes were noted in a few casts.

4. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

5. ALIGNCTD

The deck unit used for this mission advances both conductivity channels by 0.073s. Separated tests were done for the two CTDs.

- For CTD #0550 cast #26 was studied at two depths using advancements of +0.050s, +0.060s, +0.073s and +0.083s; the setting used by the deck units (+0.073s) was found to produce the best results for both sets of sensors.

- For CTD #0443 casts #21 was studied using advancements of +0.045s, +0.050s, +0.055s, +0.060s, +0.073s and +0.083s; the best results for both channels were with +0.050s. A second cast was checked to confirm that +0.050s was better than +0.073s.

All casts using CTD #0443 were realigned using –0.023s for a net advancement of +0.050s for both conductivity channels.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2		Pass 2    Std Dev = 5		Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of casts (#26 and 34 for CTD0550 and #21, 46,48 and 52 for CTD0443) was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. The descent rate was noisy for CTD #0550 and extremely noisy for CTD #0443. The casts for which no rosette was mounted have the worst descent rates; this may be because they were run in the worst weather or the rosette may stabilize the package. The average descent rate is fairly high.

- CTD 0443: The differences in temperature, conductivity and salinity were found to be about +0.0005C, +0.0005units and +0.005units, respectively below 4000m; the positive values indicate that the secondary sensors gave higher values. There was considerable variability in these values with cast 21 having lower differences than the others. The temperature differences were extremely noisy but the average is close to zero.

- CTD 0550: The differences in temperature, conductivity and salinity were found to be about              -0.0008C, -0.0005units and -0.005units, respectively below 4000m; the negative values indicate that the primary sensors gave higher values. The temperature differences were very noisy.

For both sets of sensors the salinity differences are fairly constant with pressure.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. Errors in headers of 12 files had to be corrected before this step was successful.

The ROS files were converted to IOS files, and the extensions were changed to BOT. Again, headers had to be corrected first in 12 files.

10. Checking Headers

A header summary and header check were produced and one error in position and one in station name were found and corrected. For cast #5 the PH channels which had been identified as USERPOLY0 and USERPOLY1 are now called Fluorometer:0 and Fluorometer:1 and for cast #49 USERPOLY1 has become Fluorometer:1. That will be corrected later.

The surface check was run. The average surface pressure is 2.6db.

The cruise track was plotted and looks reasonable.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 0.5

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS

Most warnings refer to surface and upcasts only. Mid-depth downcast warnings were investigated for a few casts and found to correspond to pad values in the pressure channel or reversals of the CTD. To avoid loss of good data the pad values in pressure were replaced with the pressure of the previous record. Some smaller pressure spikes were also found which had not been replaced by pad values, probably because they occurred in small groups. Where the T and S values make it obvious that the pressure did not really change as much as is indicated the pressure was also replaced with copies of the last good record. After these editing steps DELETE was rerun on all affected casts.

All DEL files were copied to EDT files.

12. COMPARE

- CTD #0443 – With only 11 deep bottles available for comparison and a lot of scatter an estimate was made of pressure-independent errors. The primary sensors were found to be low by about 0.004units and the secondary, high by 0.002units. There is no evidence of change during the cruise.

- CTD #0550 – There were many deep bottles available for comparison. When plotted against file pari number it became obvious that cast #34 was very different from the other casts. When cast #34 was excluded from the comparison the primary sensors were found to be high by about 0.0011units and the secondary, low by 0.0047units. On the Salinity Analysis sheets for the samples taken during cast #34 there is a note that the salinometer was unstable and that there was a draft in the lab. So it is reasonable to exclude that cast form the comparison.

These results are in reasoanble agreement with the differences found between the pairs of sensors as described in section 8. (See plots in Processing Section of Report.)

The details of the comparison were checked for indications of errors in the hydro files.

- Cast #34, as stated above has salinities that are low by about 0.005units.

- Cast #53, sample #188 appears to have the pressure wrong. There is confusion in the Rosette log but this sample appears to be from the pressure that is crossed out (1000db).

- Cast #62, sample #231 has a salinity that looks more like 175db than 200db. There is no obvious error, but the value was dropped from the comparison. The difference is similar for both sensors.

- DETAILED EDITING

While the secondary sensors are closer to the bottles for CTD #0443, the primary sensors have less noise in the signals. For CTD #0550 the primary sensors give salinity closer to the bottle values and there is little difference in noise between the two pairs. So primary salinity was chosen for further processing for both systems.

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

In general the descent rate of the CTDs was very noisy (except for the Rivers Inlet section) and there are many data records that appear to be corrupted by shed wakes. The DELETE routine has removed a lot of bad data, but much remains just above and below the blank areas left by DELETE. Such bad data was removed using CTDEDIT where the effect on the final data set appeared likely to be significant (a shallow feature with large temperature or salinity differences relative to the local gradient, or a deep feature with smaller differences).

A study was made of cast #37 to determine the effect of shed wakes to data at depth. Metre-averaged data was compared with and without editing to remove corrupted records. Averaging over the whole water column the effect of not editing is small (0.0005C and 0.0002units for T and S). However, individual bins may have unedited temperatures lower than edited temperatures by as much as 0.01C even below 1000db. Almost all the errors have the same sign reflecting the average gradient. (See plot in Processing Section of Report.)

CTDEDIT was used to clean S and T in all casts except  #61 and 62. The editing was very slight for casts 60 and 63-67.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

Plots of temperature and salinity vs record # for the BOT files were examined to check for any bad values in these files. Cast #56 was edited to clean 3 points in salinity at 200m depth. (See plot in Processing Section of Report.)

14. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

15. Other comparisons

Previous experience with these sensors – No data has been processed using these sensors since their last calibration.

Historic ranges – Four casts were plotted with the historic ranges superimposed. All data fell within those ranges. (See plots in Processing Section of Report.)

Comparison of nearby casts – T-S plots were made for the deep parts of casts 10-15, 15-23, 29-34 and 43-48 and adjacent casts are in good agreement. (See plots in Processing Section of Report.)

Comparison of  CTDs – The T-S plots of cast #10-13 and #43-48 include casts from each of the two CTDs. The casts are in reasonable agreement but the two CTDs are clearly grouped with a difference of about 0.025 salinity units for the P12-P18 line (at 1400db) and 0.005 for the ED3-ED7 line (at 2000db). The latter difference is close to the difference between the two comparisons done using bottles. A larger difference is expected for the P12-P18 line since it is 3 times as long and the casts were shallower than for the ED3-ED7 line. (See plots in Processing Section of Report.)

Repeat Casts – Casts #34 and 37 were at the same site. The two casts agree extremely well at 1500db. (See plot in Processing Section of Report.)

- Recalibration

File 2106rcal.ccf was used to recalibrate the data (including the rosette files) as follows:

- CTD 0443 - primary salinity – increase by 0.004units

- CTD 0443 - secondary salinity – decrease by 0.002units

- CTD 0550 - primary salinity – decrease by 0.0011units

- CTD 0550 - secondary salinity – increase by 0.0047units

After recalibration the deep T-S plot for stations ED3-ED7 was repeated and the casts are much closer.

The SAM files were recalibrated and named SAC; these files were used to rerun COMPARE. The results (see 443recal.xls and 550recal.xls) give average differences of points in the fit after recalibration as        –0.0015 and –0.0014units of salinity for the primary and secondary sensors for CTD #0443 and               –0.00003units for both sensors for CTD #0550. (See plots in Processing Section of Report.)

17. REMOVE and REORDER

The following channels were removed from all casts: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The channels were reordered and data format corrected as needed. The various channel names that were used for PH sensors (Userpoly0, Userpoly1, Fluorometer:0, Fluorometer:1) were renamed to PH.

The channels for the rosette files were also reordered but only Conductivity:Primary, Conductivity:Secondary and Flag channels were removed.

18. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

21060002 – Descent rate of CTD low and noisy with many shed wakes corrupting the data.

21060005 – Two PH sensors mounted on CTD. Both data sets look suspicious.

21060012 – Pressure spikes noted in raw data.

21060015 – 21060019 – The FL signal is odd near the surface, but up and downcasts agree.

21060021–21060022 – Transmissivity full of spikes but the top 100m of downcast looks ok.

21060023 – Pressure spikes noted in raw data.

21060024–21060044 – Fluorometer data looks odd in upcast with many spikes in the top 200m.

21060037 – PH sensor mounted but cap was left on probe so no signal recorded.

21060049 – Fluorometer data looks odd in upcast with many spikes in the top 200m.

21060049 – end of cruise – PH sensor mounted with CTD; data recorded and looks ok.

21060053 – Upcast corresponding to downcast #52.  Use rosette file but delete CTD cast.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001- 06

Dates: Start: 5 February 2001 End: 22 February 2001

Location: North-East Pacific

Vessel: CCGS John P. Tully

Party Chief: Marie Robert

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |
| 2 | SEABIRD | 911+ | 0550 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0550

Cruise ID#:		2001-06

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 02/05/00 | Factory |  |  |
| Conductivity | 1764 | 28/09/00 | “ |  |  |
| Secondary Temp. | 2668 | 07/03/00 | “ |  |  |
| Secondary Cond. | 2102 | 28/09/00 | “ |  |  |
| Fluorometer | 2229 | ? |  |  |  |
| Transmissometer | 333D | 14/07/00 | IOS |  |  |
| PH - Userpoly0 | 180339 | ? |  |  |  |
| PH -Userpoly1 | 180293 | 08/02/01 |  |  |  |
| Pressure Sensor | 75636 | 04/03/99 | Factory |  |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0443

Cruise ID#:		2001-06

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 02/11/00 | Factory |  |  |
| Conductivity | 1729 | 28/12/99 | “ |  |  |
| Secondary Temp. | 2023 | 01/11/00 | “ |  |  |
| Secondary Cond. | 2128 | 31/10/00 | “ |  |  |
| Fluorometer | 2229 | ? |  |  |  |
| Transmissometer | 333D | 14/07/00 | IOS |  |  |
| PH | 180293 | 08/02/01 |  |  |  |
| Pressure Sensor | 63507 | 31/11/96 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.

# Moving files to new archive. April 17, 2003 – Joe Linguanti

- Any of the rosette files that contained only one record were deleted since these data were collected for the loop file.

- Program HEADER EDIT was used to change non-standard channel names and units and add additional metadata to the CTD and Rosette files.

- Program Clean was used to;

- reset number of records, minimum and maximum values

- delete empty channels where applicable

- set channel name initial letters to uppercase

- Files .CLN were copied to new archive and renamed to the new file convention “YYYY-CC-NNNN.* “.

- The loop file was obtained from Frank Whitney. I added two new columns to Frank’s spreadsheet, Sample_Method and Pressure. The latitude and longitude values were computed to 3 decimal places. The spreadsheet file was then converted to IOS Header format with meta-data added to it. The original .XLS file is in the DOC directory. Bernard’s original loop file is in the HISTORY directory.

- The thermosalinograph files were copied from the TRACKOBS directory in the OSAP Data Archive.