## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| June 14, 2006 | CTD cast 183 – recalculated oxygen in mL/L as the original calculation was in error. J.L. |
| July, 29, 2005 | Added DIC, Alkalinity and pH data from the Haida project to rosette files. Added PC, PN and TOC data, from GLO-FLO bottles, later in August after Lisa Miller had supplied me with dates and locations. The GLO-FLO files have cast numbers > 9000. Some files already contained DIC and Alkalinity data so that data was replaced with the new quality controlled version. J.L. |
| Oct 22, 2001 J. Linguanti | Bottle cast 220 – Salinity value at 1250db was removed since value was low due to questionable analysis. |
| March 18, 2002 G.Gatien | SBE oxygen data was calibrated in CTD,RAC,CHE files using 2108oxy.ccf. Formats were changed and channels added as needed. See oxygen_calib.doc A note was put in headers that oxygen is to be considered ±5%. |
| April 17, 2003 | Moved files to new archive. |

PROCESSING NOTES

Cruise: 2001-08

Agency: OSAP

Location: NE Pacific

Project: Haida Eddy/Line P

Party Chief: Frank Whitney

Platform: CCGS John P. Tully

Date: May 28, 2001 – June 29, 2001

Processed by: Germaine Gatien

Date of Processing: July 6, 2001 – August 24, 2001

Number of original CTD casts: 100

Number of casts processed: 99

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with Transmissometer 333D, Seapoint Fluorometer #2228 and a PAR sensor. The deck unit was a Seabird 11+ (serial number 0424).

SUMMARY OF QUALITY AND CONCERNS

The data was generally in good order. The descent rate was very noisy for many casts and the secondary conductivity was spiky.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. As found during 2001-10 there is a small difference between the calibrations for pressure in the con files and on the web site, but the con file appears to be correct.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using configuration file 0443fluo.con.

A preliminary check shows all expected channels present. As noted in the log the oxygen sensor was not on for cast #1 and the PAR sensor was used only for casts #1-27,61,85,106,132,148,182,192,203,211, and 219. Although there was no sensor present some casts had a very noisy signal for PAR.

Dark values were checked for the fluorometer and are generally about 0.06mg/l at 2000db. These values are consistent with the use of the 10X cable which is believed to be the one used for this cruise. The fluorometer channel has an odd look at casts 214 and 215 and possibly others; the minimum value is around 200-300db, with higher values at depth. This type of profile does not occur at all casts.

There was a problem converting cast #179. There are a lot of spikes, but no useful data. This data looked ok at sea, so repeated attempts were made to convert it changing a number of things in the configuration file with no effect. The routine FIXDAT.EXE was then used and this produced good data. The repaired DAT file was renamed 21089179.DAT, but in subsequent processing steps the name 21080179 was used for this cast.

Cast #185 has an empty DAT file. This cast was deleted.

There was a problem converting cast #189. It was noted in the CTD Log that the cast was full of spikes. Neither changes to the CON file nor using FIXDAT produced usable data.

The temperature and salinity traces track reasonably well and up and downcasts are similar, but there are many spikes in the secondary channels. The differences in salinity appears to be about 0.002units.

4. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

5. ALIGNCTD

The deck unit did not advance the secondary conductivity channel.

Tests on cast #27 show that an alignment of conductivity channels by +0.060s minimizes salinity spikes. This matches the results of cruise 2001-13.

All casts were realigned using –0.013s for the primary and +0.060s for the secondary so that the net alignment is +0.060s for both channels.

A study was made of the alignment of the SBEOXY channel. This is the first data set processed that includes this instrument. The manual suggests advancing the oxygen channel for a pumped instrument by +2s to +5s. Casts #116 and 121 were studied in detail by plotting the up and downcasts together displaying just oxygen and primary temperature. Getting the two traces to overlie each other is the method suggested in the SeaBird manual. In shallow water this might be reasonable although local effects might cause real changes and in deep water the effect of internal waves would cause the traces to not overlie each other. So deep casts were chosen and the temperature was compared and taken to be the standard so that the internal wave problem is allowed for. Casts without rosette sampling were selected to simplify the comparison. For cast #116 an advancement of +8s led to optimal results, whereas an advancement of +9s or  +10s was needed for cast #121.

Casts #31 and 228 were chosen for testing to see if there is any obvious temporal variation.The former was best with an advancement of +8s and the latter around +6 or +7. It was noted that cast #228 was had a much noisier signal than the earlier casts. On investigation all casts from #188 to the end were noisy.

It has been noted by those studying the calibration, that there was a notable change after the 4233m-cast (#186); the change in noise level occurs at the same time.

A shallow cast (#106) was studied to see if the results were similar. An advancement of +7 or +8s seemed optimal.

So there appears to be considerable variation, but an advancement of +8s seems a reasonable compromise and was applied to all casts.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels except scan number.  Parameters used were:

Pass 1    Std Dev = 2		Pass 2    Std Dev = 5		Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present.

The descent rate was often very noisy but the average rate was kept reasonably high for most casts.

The differences in temperature, conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 1 | 225 | +.001 | -0.00012 | -0.0023 | ~.9m/s,ok |
| 32 | 3200 | ~0 | -0.0002 | -0.0025 | Low and extremely noisy |
| 107 | 3200 | ~0 | -0.0002 | -0.0025 | Fairly noisy, ~.7m/s |
| 186 | 4200 | +0.0002 | -0.0002 | -0.0025 | Extremely noisy avg ~1m/s |
| 204 | 3200 | ~0 | -0.00018 | -0.0023 | Fairly noisy, high |
| 209 | 3200 | ~0 | -0.00018 | -0.0023 | Noisy, fairly high |
| 220 | 1200 | ~0.001 | -0.0001 | -0.002 | High |

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

10. Checking Headers

A header summary and header check were produced and errors in the headers for casts #1 and 37 were corrected. The surface check was run. The average surface pressure is 3.1db. The cruise track was plotted and looks reasonable.

- DELETE

- CLEAN was run to replace pad values in pressure with interpolated values.

- The following DELETE parameters were used:

- Surface Record Removal: Last Press Min

- Maximum Surface Pressure (relative): 20.00   Pressure Tolerance: 1.0

- Pressure filtered over 15 points

- Swells deleted. Warning message if pressure difference of 2.00

- Drop rates < 0.3m/s (calculated over 15 points) was deleted.  Sample interval =  .04 seconds.

- COMMENTS ON WARNINGS: The only warning was for cast #179 which is full of spikes and appears to have no useful data.

- Maximum pressure check: No problems noted other than for cast #179.

- All DEL files were copied to EDT files.

- 12. COMPARE

- The BOT files were examined for errors and CTDEDIT was used to fix a few bad points in casts #: 107, 115,155,186,196 and 203. Note was made in the headers of editing done. The edited files were renamed BOT.

- COMPARE was run using HYD and BOT files. The details of the comparison were checked for indications of errors in the hydro files. The bottle at 1250db for cast #220 and all bottles from cast #198 had anomalously low salinity. These points and a few other outliers were removed from the comparison as were all values from above 200db.

- The primary salinity was found to be high by 0.0023units and the secondary high by 0.00008units. The trendlines were quite flat when outliers were excluded. Both pairs of sensors showed a slight drift with time, but there is so much scatter around the trendlines that this is not really clear. What is clear is that  the large drift noted in the secondary sensors during 2001-10 and 2001-13 did not occur during 2001-08.

- DETAILED EDITING

All casts were examined for differences between the pairs of sensors and differences between up and down casts for evidence of problems with the processing. The secondary salinity is unusually spiky in many casts. The secondary salinity looks wrong in casts #86 and 214 and the primary seems wrong for cast #183. Nearby casts were compared and the secondary salinity for cast #183 looks close to adjacent casts while the primary seems too low.

The primary sensors were chosen for further processing. While the results of COMPARE show the secondary sensors having values closest to the bottles, there was more noise in those channels and the history of drift during 2001-10 and 2001-13 raises some doubts about their performance. The secondary sensors will be chosen for cast #183 only.

Page plots were produced using T0,S0 for all casts except #183 for which T1,S1 was plotted. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

All casts except #106 required some editing.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

14. SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness.

[Due to a problem with some IOS SHELL routines the Head Editor had to be used to change the channel name for the oxygen sensor to OXYGEN(SBE). This was done to the BOX files and the BOT files. Changes should be made to the parameter file to avoid this problem in future.]

BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. These files were then stored in a separate directory for Angelica Peña.

15. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen to check whether further editing was needed.

15. Other comparisons

Previous experience with these sensors –

Since its latest calibration the primary conductivity sensor has been used for 2001-10 in April and 2001-13 in May 2001. The primary salinity was found to be low by about 0.003 during 2001-10 and high by 0.0033units during 2001-13. There were no deep bottles during the earlier cruise and many during the latter.

The secondary sensors were used for 2001-06 in February, 2001-10 in April and 2001-13 in May. For 2001-06 the secondary salinity was high by 0.002units (there were few bottles and a lot of scatter for that comparison). For 2001-10 there was a drift from CTD being low by 0.03 at the beginning to low by 0.015 at the end. For 2001-13 there was a drift from low by 0.01units at the beginning to values very close to the bottle values at the end.

Historic ranges – Three deep casts were plotted with the historic ranges superimposed. All salinity fell within those ranges as did shallow-water temperatures. At about 2000db the temperatures are a little above the range maxima. (See plots in Processing Section of Report.)

- Recalibration

File 2108rcal.ccf was used to recalibrate the AVG files (including those in the IOS directory, the HYDRO directory and the FLUOR directory) as follows:

- primary salinity –  offset: -0.0023units

- secondary salinity –  no change

- Chl_fluorescence -  offset: -0.055mg/l

The SAM files were recalibrated and named SAC; these files were used to rerun COMPARE. The results (2108com2.xls) indicate that the deep salinity values after recalibration were within 0.00003units of the bottles. (See plots in Processing Section of Report.)

17. REMOVE and REORDER

The following channels were removed from all casts except #183: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The following channels were removed from cast#183 only: Scan_Number, Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The oxygen channel was stripped from cast #1.

The PAR channel was removed for casts #: 31-58,62-82,86-102,107-130,133-147, 155-180,183-191,193-202,204-209,212-218,220-225.

The channels were reordered and data format corrected as needed.

The channels for the rosette files were also reordered but only Conductivity:Primary, Conductivity:Secondary and Flag channels were removed from all casts. The Oxygen channel was removed from cast #1 in the bottle files.

18. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

- no oxygen sensor installed, no GPS data

- File wrongly entered as 21080002, renamed 21080003; oxygen sensor installed

- The primary and secondary conductivity look quite different in the top 10m suggesting a pump problem.

- Log notes GPS problems

- Pump turned on late.

86   secondary salinity is very odd.

179  spikes noted at sea.

181 file wrongly entered as 21080181, renamed 21080182.

183 salinity sensors very different. Primary appears to be wrong.

185 empty DAT file, no other files with this #. DELETE

189 spiky, no good data obvious. CHECK

202 problem with short between rosette and deck unit. Data to 429m on downcast and 1483-1307 on upcast.

204 spooling problem in upcast so returned to 3350 and went up again.

214 secondary salinity looks odd during downcast.

220 spooling problem during upcast stopped at 496 and went back down to 527 then up again.

PAR sensor on for casts 1-27, 61,85,106,132,148,182,192,203,211,219.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001- 08

Dates: Start: 28 May 2001 End: 29 June 2001

Location: NE Pacific

Vessel: CCGS John P. Tully

Party Chief: Frank Whitney

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0443

Cruise ID#:		2001-08

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 02/11/00 | Factory |  |  |
| Conductivity | 1763 | 31/10/00 | “ |  |  |
| Secondary Temp. | 2023 | 01/11/00 | “ |  |  |
| Secondary Cond. | 2128 | 31/10/00 | “ |  |  |
| Fluorometer -pumped | 2228 | ? |  |  |  |
| SBE43 Oxygen sensor | 0047 | 16/05/01 | ? |  |  |
| Transmissometer | 333DR | 26/05/01 | IOS |  |  |
| PAR | 4615 | 08/02/01 |  |  |  |
| Pressure Sensor | 63507 | 31/11/96 | Factory |  |  |

A second SeaBird 911+ (S/N 0585) was taken on the cruise but not used.

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.

# Moving files to new archive. April 17, 2003 – Joe Linguanti

- Any of the rosette files that contained only one record were deleted since these data were collected for the loop file.

- Program HEADER EDIT was used to change non-standard channel names and units and add additional metadata to the CTD and Rosette files.

- Program Clean was used to;

- reset number of records, minimum and maximum values

- delete empty channels where applicable

- set channel name initial letters to uppercase

- Files .CLN were copied to new archive and renamed to the new file convention “YYYY-CC-NNNN.* “.

- The loop file was obtained from Frank Whitney. I added two new columns to Frank’s spreadsheet, Sample_Method and Pressure. The latitude and longitude values were computed to 3 decimal places. The spreadsheet file was then converted to IOS Header format with meta-data added to it. The original .XLS file is in the DOC directory. Bernard’s original loop file is in the HISTORY directory.

- The thermosalinograph files were copied from the TRACKOBS directory in the OSAP Data Archive.

- Program HEADER EDIT was  used to change non-standard channel names and units and add additional metadata to the files.

- IOS Header Date and Time channels were added.

- The FLAG channel was removed

- Channels were re-ordered

- Program Clean was used to set the start and end times and the geographical area.